## Slide 1: TASC Basic Analysis
Kenji YOSHIDA

## Slide 2: Pedestal in TASC

## Slide 3: Pedestal measurements in TASC 2nd – 12th layer PWOs
An example pedestal measurement (TASC 8th layer, lower PWO with APD H.G.)

## Slide 4: Time Variation of Pedestal Peaks
TASC 3rd layer, middle PWO with APD H.G.

## Slide 5: Rejected pedestal data
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
*
Fitting Line
Error×3
Error×3
Rejected
An example of rejected pedestal
ADC count [ch]
counts
As a general rule, muon, e, and p run data are corrected by using the last pedestal run data.
(*)Pedestal data separated > 3 sigma are rejected.

## Slide 6: Muon calibration

## Slide 7: 1MIP determination with muon beams
Si Strip
TASC
PWO | Ttop 1 | Ttop2 | Ttop3
Distance from Si Strip center [cm] | -2.2~-1.2 | -0.5~0.5 | 1.3~2.0
beam
＊IMC is omitted in this figure
12 layers
Electrode gap for Si strip：0.183mm
Read out by collecting 4 electrodes
-> Si strip interval: 0.732mm
-> Track determination with0.7mm
Ttop2
Ttop1
Ttop3
Si
Event Selection
The upper PWO has lower events than the lower and middle PWOs.

## Slide 8: 1MIP determination by using muon beams
Convolved function of Gauss to Landau distribution
1MIP = the Most Provable value of the Landau component
Pedestal data are subtracted

## Slide 9: Temperature effect for 1MIP determination

## Slide 10: Temp. Correction
PWO : Temperature dependent amount of Luminescence
APD : Temperature dependent properties
Temperature correction for 1MIP
=> 1MIP Temperature Correction for each PWO APD H.G. signal
Temp. measurements with thermocouples
6Layer
100
TASC BOX
PWO
350
[mm]
460
120
318
130
beam
An example of temperature variation for PWO APD H.G. signals (TASC 11th layer, Lower PWO)
1MIP = p0 + p1*T(℃)

## Slide 11: High/Low Gain

## Slide 12
High gain VS low gain in TASC (first layer)
p 350GeV
In fitting with a linear function, these signals are removed.
Correlation between High Gain and Low Gain for the TASC First layer PWO

## Slide 13: Event Selection to derive the correlation between H.G. and L.G.
Low gain [ADC count]
High gain [ADC count]
difference from fitting graph
make histogram
High gain VS low gain in TASC first layer
The distribution of differences from fitting graph
(red curve: Gaussian curve)
Event data within 3 sigma are selected.
3σ
3σ
removed
removed

## Slide 14
High gain VS low gain in the TASC first layer PWO (before event selection)
High gain VS low gain in the TASC first layer PWO (after event selection)
PMT-H.G. = p0 + p1*PMT-L.G.
PMT Low Gain (ADC count)
PMT Low Gain (ADC count)
PMT High Gain (ADC count)
PMT High Gain (ADC count)

## Slide 15: Variation of the correlation between High gain and Low gain in the TASC first layer PWOs
PMT-H.G. = p0 + p1*PMT-L.G.
P1
P0

## Slide 16: Correlation between APD High Gain and APD Low Gain in TASC (2nd – 12th layer)
p 350GeV
APD high gain VS APD low gain of a PWO in TASC (2nd ~12th layer)
Fitting with a linear function
1000
6000
(APD-H.G.) = p1*(APD-L.G.)

## Slide 17: Examples of correlation among APD-H.G., APD-L.G., PD-H.G., and PD-L.G.
Correlation between APD High Gain and APD Low Gain
Correlation between APD Low Gain and PD High Gain
Correlation between PD High Gain and PD Low Gain
Overflow
By using Proton 350GeV data
(*) Red line: preliminary fitting results.
Unused in this analysis.
(*) In this comparison with M.C., PD H.G. and PD L.G. are not used.

## Slide 18: Temperature dependence on APD Low Gain/PD High Gain
(*)TASC 2nd layer, bottom PWO

## Slide 19: Energy Deposit distribution for each TASC layer

## Slide 20: Energy Deposits. In each TASC layer
1layer
12layer
10
102
103
100
Total Energy Deposit for TASC
Comparison with M.C. Simulation
Positron 30 GeV
Vertical axis: Counts in linear scale
Horizontal axis: MIPs in log scale

## Slide 21: Back up

## Slide 22: Convolved function of Gaussian to Landau distribution
Landau distribution
Landau function
Gauss Function
Gauss distribution
Gauss function <= Pedestal distribution
・・・
・・・
Convolution
Most Probable Value (MP)

## Slide 23: Signal Readout of TASC
① APD + High Gain Amp.
② APD + Low Gain Amp.
③ PD + High Gain Amp.
④ PD + Low Gain Amp.
APD,PD: multi-gain amp.
=> Dynamic range: >10^6
10-1
100
101
102
103
104
105
106
[MIP]
4
40
400
4000
Effective ADC [ch]
①
②
③
④
APD
PD
ch | APD/PD gain | Acceptance surface | Amp. Gain ratio | Total Gain ratio
①APD+High Gain Shaper | 50 | 18 | 30 | 2.7×104
②APD+Low Gain Shaper | 50 | 18 | 1 | 9.0×102
③PD+High Gain Shaper | 1 | 1 | 30 | 30
④PD+Low Gain Shaper | 1 | 1 | 1 | 1
13.4mm
19.7mm
PD
2.4×2.4mm2
APD
10×10mm2
S10937-9351
(Hamamatsu Co.)
Comparison of each Gain to PD+Low Gain (=1)
×
×
×
×
×
×
×
×
=
=
=
=

## Slide 24: Variation of the correlation between H.G. and L.G. in the TASC First layer PWOs