## Slide 1: Position Calibration
2012.3.16
Y. Akaike

## Slide 2: Detector Configuration
Movable plate
Beam
Z
X
side view
Movable plate
Beam
Z
Y
Top view
SIA
SIA
IMC box
TASC box
IMC box
TASC box
Strip
Strip
CHD
SciFi
CHD
SciFi
PWO
PWO
Trig
Trig
32.8cm

## Slide 3: Analysis method
Correction for Si-Strip alignment
Track reconstruction
Correction for SciFi position
Correction for PWO position
Correction for CHD position
Derivation of beam profile
Get accurate position of each detector relative to Si-Strip from reconstructed track
Correction for Trig. Scin position
Define track position on Strip layers as beam incident position

## Slide 4: Correction for Si-Strip alignment
Si-Strip
Luminous strip
correction
⇒ Correct misalignment between Si-Strip layers using most
luminous fiber of each layer

## Slide 5: Correction for Si-Strip alignment
Position distribution of most luminous fiber of each Si-Strip
（horizontal ：position [cm], vertical : counts）
Peak value of fitting function (Gaussian)
※ used data
Proton 350GeV
（H4A.951）
 | X-layer |  | Y-layer | 
 | Peak value [cm] | Difference from Layer 1 [cm] | Peak value [cm] | Difference from Layer 1 [cm]
Layer １ | 0.174 | 0 | 0.437 | 0
Layer ２ | 0.123 | 0.051 | 0.432 | 0.005
Layer ３ | -0.045 | 0.219 | 0.940 | -0.503
Layer ４ | -0.339 | 0.513 | 0.703 | -0.266
Layer １：black
Layer ２：red
Layer ３：green
Layer ４：blue
Adjust position of Layer 2-4 with reference to Layer 1
X-layer
Y-layer

## Slide 6: Track reconstruction using Si-Strip
ΔE > 0.3MIP in at least 3 layers
Select most luminous fiber from each layer
as a track pass point candidate
Fit by straight line
R.M.S of residuals < half width of fiber
Reconstruction complete
worst candidate = most luminous fiber
Choose second luminous fiber instead of most luminous one
Exclude the fiber from candidates
Drop out
Yes
Yes
Yes
No
No
No

## Slide 7: Estimation of beam profile
Define reconstructed track position on Si-Strip layers as beam incident position
ex：beam profile of proton 30GeV （RunNo.288）
projection
projection
※use the adjusted position with
reference to Layer 1

## Slide 8: Correction of CHD position
Get accurate position of CHD from the distribution of track position on Si-Strip
※use muon data, height of movable plate = 131.56mm
Detection Efficiency
Track Position [cm]
Detection efficiency of CHD0
(ΔE in CHD > 0.5MIP )
 | Center pos. [cm] | Size [cm]
CHD0 | 0.412 | 2.95
CHD1 | 0.438 | 2.96
CHD2 | 0.760 | 3.13
CHD3 | 0.760 | 3.13
c
μ
※ set a, b, c,μ,σ as free parameter

## Slide 9: Correction for SciFi position
Get accurate position of SciFi layers from the distribution of track position on Si-Strip
※use muon data, height of movable plate = 131.56mm
Detection Efficiency
Track Position [cm]
Detection efficiency of SciFi layer1
(sum of dE in SciFi layer > 0.5MIP)
 | Center pos. [cm] | Size [cm]
Layer1 | 0.529 | 3.10
Layer2 | 0.580 | 3.13
Layer3 | 0.524 | 3.08
Layer4 | 0.562 | 3.13
Layer5 | 0.567 | 3.13
Layer6 | 0.569 | 3.12
Layer7 | 0.574 | 3.12
Layer8 | 0.505 | 3.11
c
μ
※ set a, b, c,μ,σ as free parameter

## Slide 10: Correction for PWO position
Get accurate position of PWO from the distribution of track position on Si-Strip
Detection Efficiency
Track Position [cm]
Detection efficiency of PWO in Layer1
(dE in PWO > 0.5MIP)
 | Center pos. [cm] | Size [cm]
Layer1 | 0.326 | 1.91
Layer2 | 0.292 | 1.88
Layer3 | 0.294 | 1.85
Layer4 | 0.286 | 1.87
Layer5 | 0.291 | 1.87
Layer6 | 0.320 | 1.88
Layer7 | 0.306 | 1.88
Layer8 | 0.312 | 1.85
Layer9 | 0.310 | 1.85
Layer10 | 0.315 | 1.87
Layer11 | 0.329 | 1.85
Layer12 | 0.388 | 1.98
c
μ
※use muon data, height of movable plate = 131.56mm
※ set a, b, c,μ,σ as free parameter

## Slide 11: Correction for trigger scintillator position
Draw a map of track position on trigger scintillator
Projection X, Y (horizontal : position, vertical : counts)
※ use muon150GeV
Position of trig. scin.
X-position
Y-position
Position [cm]
Position [cm]
4.392cm
4.9044cm
-0.366cm
-0.4758cm
Estimate center position difference between trigger Scintillator and Si-Strip layer1,
and size of the scintillator

## Slide 12: Size of each detector
 | Standard [cm] | estimated [cm]
CHD0 | 3.0 | 2.95
CHD1 | 3.0 | 2.96
CHD2 | 3.2 | 3.13
CHD3 | 3.2 | 3.13
CHD
 | Measured [cm] | estimated [cm]
Layer1 | 3.20 | 3.10
Layer2 | 3.20 | 3.13
Layer3 | 3.19 | 3.08
Layer4 | 3.18 | 3.13
Layer5 | 3.10 | 3.13
Layer6 | 3.21 | 3.12
Layer7 | 3.19 | 3.12
Layer8 | 3.20 | 3.11
IMC
 | Standard [cm] | estimated [cm]
Layer1 | 1.9 | 1.91
Layer2 | 1.9 | 1.88
Layer3 | 1.9 | 1.85
Layer4 | 1.9 | 1.87
Layer5 | 1.9 | 1.87
Layer6 | 1.9 | 1.88
Layer7 | 1.9 | 1.88
Layer8 | 1.9 | 1.85
Layer9 | 1.9 | 1.85
Layer10 | 1.9 | 1.87
Layer11 | 1.9 | 1.85
Layer12 | 2.0 | 1.98
TASC
Analytical size estimated from Si-Strip data is ~2 % smaller than actual measured value

## Slide 13
CHD：0.982
PWO：0.984
SciFi：0.978
Total：0.978
Estimated / measured
Estimated / measured
Estimated / measured
Estimated / measured

## Slide 14
Assuming the detector incline,
  L’ / L = cos θ ≒ 0.98
  θ = 11.5 degree
  ← too big?
  Strip-pitch is exactly 732 um?
  The beam direction is spread?
  An assumption that the beam direction is exactly vertical is bad?
L
L’
θ

## Slide 15: END