## Slide 1
Commissioning of the new XPS
SPECS system
October 2011 – May 2012
Section:
Surface Coating &Chemistry
Elise USUREAU

## Slide 2: Contents
XPS presentation
Commissioning
Comparison with older system (ESCA5400) + Issues
Future work
Conclusion

## Slide 3: XPS presentation
XPS: X-ray Photoemission Spectroscopy
EK = hn – EB
– F
Photoelectric effect
Photoelectrons from Soft X-rays
 Surface analysis (0.3 – 3.0 nm)
Sweep in Kinetic Energy to know the Binding Energy of the photoelectrons, which is specific to the elements:
 Spectrum: elemental composition…

## Slide 4: Parts of XPS system
HSA: Hemispherical Analyser
Two hemispheres (R1, R2) concentric
Potentials applied (V1 <V2)
 Electrons arriving at S can reach F
if they have the correct energy EK
 HSA at constant energy: Pass Energy EP
Analyser Components
Entrance and Exit slits: Energy resolution
Iris aperture: Lateral resolution
Entrance lens system: angular and spatially resolved studies through lens modes
Combinations possible
Detector Assembly : 9 MCD
(Multi Channel Detectors)

## Slide 5: Parts of XPS system
Source / Monochromator
XR Sources:
Al: Eexc = 1486.74 eV; Pmax = 400 W
Ag: Eexc = 2984.3 eV; Pmax = 600 W
 spot size on the sample: 1x3.5 mm2
Advantages of monochromatic XR
Monochromator principle:
Bragg diffraction
Settings:
Source position
Cristal settings (Focus, Roll, Pitch)
Non-monochromatic
rays
Monochromatic
rays

## Slide 6: Contents
XPS presentation
Commissioning
Comparison with older system (ESCA5400) + Issues
Future work
Conclusion

## Slide 7: Commissioning
Energy Resolution: Minimum FWHM (monochromator)
with good intensity (Analyzer MCD)
 | Nominal (Ag3d5/2) | C1s | Au4f7/5 (slit Ø1, Iris 5) | Au4f
FWHM (eV) | 0.6 | 0.52 | 0.37 | 0.76
Intensity (kcps) | 400 | 99 | 0.03 | 266
Ratio | 667 | 160 | 0.08 | 350

## Slide 8: Commissioning
Energy scale calibration
  Gain setting: stretch the scale – G↗ D↘
  Offset setting: shift – offset ↗ EB↗
EB_Cu-2p3/2 | 932.67 eV
EB_Au-4f7/5 | 83.98 eV
D EB | 848.69 eV
Gain = stretch parameter

## Slide 9: Commissioning
Chart: z change
| 25.650000000000002 | 25.9 | 26.1 | 26.5 | 27 | 27.2 | 27.3 | 27.4 | 27.45 | 27.5 | 27.55 | 27.6 | 27.650000000000002 | 27.7 | 27.75 | 27.8 | 27.9
iris20-3 | 94.5111 | 96.4667 | 100.7778 | 95.1556 | 99.3111 | 92.7778 | 87.1333 | 66.4889 | 52 | 38.1556 | 25.4444 | 16.1333 | 10.2889 | 6.0889 | 3.7556 | 2.0889 | 0.51111
Ir30-7.20 | 97.3991 | 99.1414 | 100.0842 | 99.5912 | 99.7377 | 94.753 | 90.0651 | 83.3679 | 72.9507 | 67.5642 | 61.2502 | 53.2153 | 47.0023 | 40.4958 | 34.6595 | 22.8163 | 13.9888 | 9.0949 | 7.4242 | 3.9935 | 2.7228 | 1.8237 | 0.62047 | 0.072558 | 0.021395
iris10-1 | 93 | 102 | 101 | 101 | 98 | 92 | 61 | 33 | 9 | 4
Combination | dx (mm) | dz (mm) | Nominal
Slit 7x30 – Iris 30 mm | 1.2 | 0.49 | 0.7 (slit 7)
Slit Ø3 – Iris 20 mm | 0.3 | 0.24 | 0.30
Slit Ø1 – Iris 10 mm | 0.12 | 0.18 | 0.1 (iris 5)
Spatial Resolution
  Square gold foil on stainless steel
  At “High Magnification” Lens Mode for different slit/iris combinations
  Displacement in x & z directions (AuSt-st)
  Au-4f7/5 line measurement for each value
  Area calculation (after Shirley background subtraction) then plotted according to direction value
d = from 80% to 20% max intensity

## Slide 10: Contents
XPS presentation
Commissioning
Comparison with older system (ESCA5400) + Issues
Future work
Conclusion

## Slide 11: Comparison
Non-monochromatic source
Signal to noise/Intensity –
Energy Resolution – Spatial Resolution:
 | Old | New
FWHM (eV) | 1 | 0.76
Signal to noise | 12 | 22
Source spot area (mm2) | 100 | 1x3.5
Photon density | 5.1013 ph.s-1.cm-2 | 1013 ph.s-1.cm-2
Spatial resolution | 0.4x0.6 mm | 0.12x0.18 mm
 | Old | New
Anode | Mg | Al (shift -0.5eV)
EP (eV) | 17.9 | 10
Chart: Au4f line comparison
Chart: Au4f line scaled

## Slide 12: Practical comparison
Larger sample
Bigger chamber
Local heating
Disadvantage: Analysis time (load lock venting time + software to set)
Others disadvantages:
  Peak energy instability < 0.1 eV
  More frequent calibration: DVS monthly
Detector voltage = CEM (channeltrons) voltage
Need to be adjusted according to dose
Maximum 3.5 kV then CEM need to be changed

## Slide 13: Quantification
Concentration of an element A in the sample
Transmission function
  Specific to the system
  Non-linear but different approximations
  Approximation within software: Tr = a.EP.(EP/EK)b
    Original settings: a = 1; b = 0.5
    Acquisition software: change a, b
    Data handling software: any function
  Need to be calculated according to lens mode
r
Peak intensity
of the element
Incident XR flux
Concentration
of the element
Cross-section
of the element
Inelastic Mean
Free Pass
Transmission Function

## Slide 14: Transmission function
Series of measurements
Cu sample: 2s, 2p, 3s, 3p lines
Variable EP: 5, 10, 20, 50, 100 eV
Lens Modes surveys: Large Area, Medium Magnification, Medium Area  only Medium Magnification is “linear”
Area calculations: after Shirley background subtraction
Plot: log(A/EP) = f (log(EP/EK)) for each line then scale
Slope = b of Tr = a.EP.(EP/EK)b
Chart: Transmission Function measurements - MM
| -1.7435803122533731 | -2.0446103079173543 | -1.4425895047075803 | -1.0447278616652149 | -0.74381538793854918
2p3/2 | 4.8585 | 4.7249 | 4.9885 | 5.1124 | 5.1473
2p1/2 | 4.5564 | 4.4224 | 4.6908 | 4.8048 | 4.8357
3s | 3.3258 | 3.1304 | 3.5949 | 3.6813 | 3.8081
3p | 3.7817 | 3.5736 | 3.9352 | 4.1406 | 4.2749
2s | 4.1095 | 3.9767 | 4.2835 | 4.3322 | 4.3945
Chart: Data scaled
| -1.7435803122533731 | -2.0446103079173543 | -1.4425895047075803 | -1.0447278616652149 | -0.74381538793854918 | -1.7277364017102956 | -2.0287663973742767 | -1.4267470501538471 | -1.0288883182874333 | -0.72793958422123128 | -2.1347952660374347 | -2.4359844383874951 | -1.8339085320160176 | -1.4359844383874949 | -1.134986271060995 | -2.1497239559596459 | -2.4508000970370758 | -1.8487247244494156 | -1.4507539516236272 | -1.1497547201133966 | -1.5912761348676479 | -1.8920277830146961 | -1.2901348216037585 | -0.89219481293172098 | -0.59144305778691408
all data | 3.934 | 3.8004 | 4.064 | 4.1879 | 4.2228 | 3.9403 | 3.8063 | 4.0746 | 4.1887 | 4.2195 | 3.6814 | 3.486 | 3.9505 | 4.0369 | 4.1637 | 3.6911 | 3.483 | 3.8446 | 4.05 | 4.1843 | 3.9784 | 3.8456 | 4.1524 | 4.2011 | 4.2634

## Slide 15: Quantification
Comparison
  Same sample for two systems: Cu + O + Ag + C
  XPS-Specs: sample set at 45° to the analyser (same configuration as XPS old)
  XPS old: Mg-source
  Moreover: Cu(OH)2 visible with XPS Specs
Component | XPS old | XPS new
C | 56.4 | 51.7
Cu | 13.1 | 11
O | 26.3 | 25.7
Ag | 4.2 | 11.6
Chart: Comparison XPS systems with same sample

## Slide 16: Contents
XPS presentation
Commissioning
Comparison with older system (ESCA5400) + Issues
Future work
Conclusion

## Slide 17: Future work
RSF (relative sensitivity factor) calculations & verification
  RSF = sA.lA.Tr specific to the system
  Measure intensity of pure elements and compare with s.l literature values with measured Tr
  Extrapolate to other elements: RSF database
Ag-source commissioning
To be installed
  UPS source
  Ion gun (Ar gas line already installed)

## Slide 18: Conclusion
First part of commissioning done
Energy calibration EB_Cu2p3/2 = 932.66 eV / EB_Au4f7/5 = 83.98 eV
Energy resolution as expected
Spatial resolution = 0.12 x 0.18 mm
Transmission function (Medium Magnification, Iris = 15 mm):
Tr = a.EP(EP/EK)b : a = 1, b = 0.41
New system much more efficient than the older
Monochromatic X-rays
Better resolution, signal to noise, precision
However need more attention (time, calibration)
Good perspectives
Future install more devices, precise quantification (RSF)

## Slide 19: Merci de votre attention…
Questions?
Acknowledgments: SCC section, especially colleagues in the lab…