## Slide 1: A Multimodal Ultrasonic Probe Featuring Time of Flight PET in Diagnostic and Therapeutic Endoscopy
APD Array
PADJK
NINO
LSO
APD Array
PADJK
NINO
LSO
Fig. 5: Functional diagram of the photon detection and signal processing layout
APD Array
PADJK
NINO
LSO
Fig. 5: Functional diagram of the photon detection and signal processing layout
Endo-TOFPET-US
FP7 Project, Call Health 2010
© 2008 CERN, for the benefit of the CMS Collaboration
“The research leading to these results has received funding from the European Union Seventh Framework Program (FP7/2007-2013) under Grant Agreement No. 256984. This research is also supported by the Marie-Curie Actions ITN: PicoSEC-MCNeT under Grant Agreement No. 289355.”
Notes: Introduction:
From physics to health? US stands for ultrasound not united states;
CMS calorimeter a real marvel, fully crystal-built
Endoscopic probe: go from 80 tons to grams and build a miniature calorimeter; as small as my small finger tip;
learn from HEP; use tools and instruments from HEP (NINO from ALICE used in TOF, crystals from CMS, FEDC05 amplifiers from ATLAS, etc); Vast pool of knowledge and competence in scintillating material;
TOF extremely useful, temporal collimation, space points, used in TOTEM, ALICE, comes at no extra charge (except for ALICE)
this research also useful for next generation of accelerators (CLIC);
Endo-TOFPET-US addresses new biomarkers specific for prostatic and pancreatic cancers
Collaboration – called consortium – of academia, medical sector and industry, perhaps a novelty in joint ventures
All partners plus one more (STM) have also engaged in the Marie-Curie Actions Program and actively participate in an Initial training network; education oriented
Our research entirely “sponsored” funded by the European Commission.

## Slide 2: The Consortium
APD Array
PADJK
NINO
LSO
APD Array
PADJK
NINO
LSO
Fig. 5: Functional diagram of the photon detection and signal processing layout
APD Array
PADJK
NINO
LSO
Fig. 5: Functional diagram of the photon detection and signal processing layout
Chart: Partners in this project:
| Medicine | Industry | Academia
Partners | 3 | 3 | 6
Notes: Very exciting environment involving research & academia and industry and medical sector

## Slide 3: EndoTOFPET-US: The Principle

## Slide 4: EndoTOFPET-US: The Principle

## Slide 5: EndoTOFPET-US: The Principle
External PET Plate
Notes: Get large field of view;
we in HEP call it: cover as much of the solid angle as possible

## Slide 6: EndoTOFPET-US: The Principle
External PET Plate

## Slide 7: EndoTOFPET-US: The Principle
External PET Plate

## Slide 8: EndoTOFPET-US: The Principle
External PET Plate
Conventional Back-Projection

## Slide 9: EndoTOFPET-US: Why TOF?
External PET Plate
With TOF Constraint
t1
t2
Target CTR: 200ps  30mm FWHM

## Slide 10: Project Objectives
Development of new biomarkers;
First clinical targets: pancreatic & prostatic cancers;
Develop a dual modality PET-US endoscopic probe with...
  Spatial resolution: 1mm;
  Timing resolution: 200ps FWHM coincidence;
  High sensitivity to detect 1mm tumors in a few minutes;
  Energy resolution: sufficient to discriminate against Compton events .

## Slide 11: PET Configuration
Asymmetric:
  One PET detector close to ROI
    incomplete (non-2p) projections;
     Simulation and reconstruction challenge.
Endoscopic:
  One PET head inside the body
     Extreme miniaturization;
     Background from close organs (e.g. heart, bladder);
    Varying geometry (body & organ motion)

## Slide 12: The Endoscopic Probe
Ultrasonic (US) probe
PET head
e.m. tracking sensor
PET head FOV
US FOV
Object under study
External plate

## Slide 13: Two Different Probes...
Ultrasonic (US) probe
PET head
e.m. tracking sensor
PET head FOV
US FOV
Object under study
External plate

## Slide 14: Two Different Probes...
22.00
22.00
Option 1: prostate
 |  | Prostate
Crystal Matrix | [mm2] | 14 x 15
Fiber length | [mm] | 10
Fiber pitch x/y | [µm] | 780/800
# Fibers in x/y |  | 18x18 (324)
Diffractive optics | [mm] | 2
SPAD array thickness | [mm] | 0.75
PCB thickness | [mm] | 1
# Readout layers |  | 1
Total thickness | [mm] | 18
Diameter of detector | [mm] | 23
Length of detector | [mm] | 22
Notes: What looks big on slide and screen is in fact only as large as a thumb or finger.

## Slide 15: Two Different Probes...
22.00
22.00
Option 1: prostate
Option 2: pancreas
22.00
 |  | Prostate | Pancreas
Crystal Matrix | [mm2] | 14 x 15 | 7 x 15
Fiber length | [mm] | 10 | 5 (9)
Fiber pitch x/y | [µm] | 780/800 | 780/800
# Fibers in x/y |  | 18x18 (324) | 9x18 (162)
Diffractive optics | [mm] | 2 | 2
SPAD array thickness | [mm] | 0.75 | 0.75
PCB thickness | [mm] | 1 | 1
# Readout layers |  | 1 | 1
Total thickness | [mm] | 18 | 13
Diameter of detector | [mm] | 23 | 15
Length of detector | [mm] | 22 | 22

## Slide 16: External PET Plate
205 mm
205 mm
15 mm
64 x 64 Crystals Total
8 x 8 Crystals/”Matrix”
3 x 3 x 15 mm3
8 x 8 crystals
SiPM “8x8-Module”, i.e. four close-butted 4 x 4 MPPC monoliths

## Slide 17: Technical Challenges
Calls for innovative solutions
  Very thin crystal pixels/fibers for the internal probe;
    for high granularity and sub-millimeter spatial resolution;
  Temporal collimation with TOF: CTR < 200ps FWHM;
    for background rejection outside ROI of 30mm;
  Digital light detection: SiPM with single SPAD readout;
    for single optical photon counting and ultimate timing resolution;
  Diffractive optics - light concentrators ;
    for overcoming loss of sensitive area and optimizing light collection;
  High level of integration for electronics & mechanics;
    For miniaturization (±5mm precision);

## Slide 18: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 19: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 20: Photodetectors: Choices
∑
a-SiPM
The a-SiPM:
low rise time;
high capacitance;
reasonable fill factor (FF);
mature technology;
commercially available
time over threshold discr.
standard (HP)TDC readout.

## Slide 21: a-SiPM versus d-SiPM
TDC
TDC
TDC
TDC
d-SiPM
“endoTOFPET”
∑
a-SiPM
The a-SiPM:
low rise time;
high capacitance;
reasonable fill factor (FF);
mature technology;
commercially available
time over threshold discr.
standard (HP)TDC readout.
very low rise time;
individual SPAD readout
   single photon counting
   optimum timing
high functionality
ambitious/risky
novel technology
optimized for endoscope
The d-SiPM:
Notes: d-SiPM:
intelligent detector/chip
functionality: active quenching, control afterpulsing, counters (1,1.5, 2 bits)
very attractive features

## Slide 22: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 23: a-SiPM: Test Scenario
☛ For an intrinsic SiPM timing evaluation, see poster by Stefan Gundacker/CERN;
Coincidence time resolution (CTR) measured with scintillating “reference” crystals and the NINO amplifier/discriminator;
Use high-BW scope (LeCroy DDA 735Zi, 40GS/s) or HPTDC.
NINO
SiPM
LSO
LSO
SiPM
NINO
22Na-Source
(2 x 2 x 10 mm3)
CTR measured with two same SiPMs:
Crystals are wrapped and mounted “small-face-to-SiPM” with grease;
Spectra are refined through photo-peak selection.
Sample CTR Spectrum
NINO chip:
Fast discrimin. (3GHz BW)
Signal peaking time: <1ns
Output time jitter: ≤ 25ps
Differential architecture

## Slide 24: a-SiPM: Evaluation & Selection
Suitable SiPMs are commercial MPPCs by Hamamatsu Photonics;
Evaluation of photodetectors via measurement of the (CTR):
MPPC S10931- | # SPADs | Fill Factor [%] | VManuf. [V] | Voptimum [V] | DCR*) [MHz] | NINO Thr. [mV] | CTR FWHM [ps]
-025P | 14’400 | 30.8 | 71.49 71.44 | 73.0 | 3.2 | 150 | 340±9
-050P | 3’600 | 61.5 | 72.11 72.09 | 72.4 | 1.1 1.0 | 100 | 220±4
-100P | 900 | 78.5 | 70.81 70.87 | 70.3 | 9.0 9.5 | 300 | 280±9
*) DCR = Dark Count Rate
☛
3mm
3mm
Note: Optimization of SiPMs done with non-optimized crystals.
DCR [MHz]
CTR FWHM [ps]
CTR FWHM [ps]
MPPC S10931-050P
S. Gundacker et al., “A Systematic Study to Optimize SiPM Photo-Detectors for Highest Time Resolution in PET” (TNS-00225-2011)
Bias Scan
Threshold Scan

## Slide 25: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 26: d-SiPM for endo-TOFPET-US
Capability to gather the statistics of the first individual photons that reach a cluster;
  Intrinsically best timing performance  attractive for TOF;
Balance functionality and PDE;
Design must adapt to process-specific effects:
  high DCR from tunneling and trap-assisted noise;
  after-pulsing;
  lower PDP than with conventional a-SiPMs.
Requires multi-parameter optimization/simulation.
*) M. Fishburn & E. Charbon, “System Tradeoffs in Gamma-Ray Detection Utilizing SPAD Arrays and Scintillators”, IEEE-TNS, VOL. 57, NO. 5, OCTOBER 2010. S. Seifert et al., “The lower bound on the timing resolution of scintillation detectors”, Phys. Med. Biol. 57 (2012) 1797–1814
Notes: d-SiPM favored for internal probe;
get the statistics of the first 48 individual photons that reach the cluster (assuming one Gamma event per cluster-column or 24 assuming two Gamma events per cluster column;
Modified d-SiPMs (Philips) can give you the first photon or the second or the third or the fourth, so it is far from the 20 photons you need to achieve the CR lowerbound (see Fig. 4 in Seifert et al.);
Improvement of one parameter easily offset by the degradation of an associated parameter;
AMS process
low PDP: Well, PDP is not to be confused with QE. In fact, PDP = QE * Prob(avalanche). The latter is dependent on the impact ionization process. The energy required to ionize a hole is indeed higher that that of the electron. So, you are right about your reasoning, however there is a difficulty in having electrons as minority carriers since it would imply a different design (which we tried) that has some major shortcomings in terms of DCR, cross-talk, and after pulsing.
p-n-junction vs. n-p-junction in CMOS

## Slide 27: d-SiPM: Sensor Floor Plan
Added functionality results in a maximum fill factor of ~50%.
“Pancreas” Imager: 9 x 18 clusters
1 Cluster .8 x .8 mm2 16 x 26 SPADs
14.4mm
7.2mm
1 LSO Fiber .8 x .8 x 10mm3
50um
30um
Test Structures
(4 x 4 or 4 x 6)
pads
TDC Array
Sensitive
Strips
Inactive Zones

## Slide 28: d-SiPM: Characteristics & Results
Characteristic Parameters of d-SiPM- “endo-TOFPET-US” Test Structure |  | Commercial d-SiPM
Cluster Pitch [μm] | 800 | 4000
# SPADs/cluster | 416 | 6400
Maximum Fill Factor [%] | 50 | 77
PDP [%] @ 430nm | 32 | 31
PDE [%] | 15 | 24
# TDCs / Cluster Column | 48 | 1
# Time Of Arrivals (TOA) / cluster | 48 | 1
TDC Resolution or LSB [ps] | 51.8 | -
Clock Frequency [MHz] | 25 | 200
First evaluation results:
DCR (0 °C): 25kHz/SPAD  DCR (40 °C): 80kHz/SPAD
NEED COOLING OF d-SiPM!
Masking of noisy pixels (“screamers”): DCR ↘ but also PDE ↘
Lowering excess bias: DCR ↘ but also PDE ↘
Boost PDE optically!
Example of multi-parameter chip optimization:

## Slide 29: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 30: Diffractive Optics: Concept
Objective: Recover light lost (50%) in the dead zones of the d-SiPM
d-SiPM (50% FF)
Crystal

## Slide 31: Diffractive Optics: Concept
Light “Concentrator”
Objective: Recover light lost (50%) in the dead zones of the d-SiPM
  Glue an optical grating between crystal and photodetector;
  Match pitch of grating with that of d-SiPM (50µm);
d-SiPM (50% FF)
Crystal
Optical Grating/
Lenticular Lens
500µm
Crystal
d-SiPM
grating
Notes: direct laser writing process, micro-structuring

## Slide 32: Diffractive Optics: Concept
Objective: Recover light lost (50%) in the dead zones of the d-SiPM
  Glue an optical grating between crystal and photodetector;
  Match pitch of grating with that of d-SiPM (50µm);
  Create ‘differential’ light pattern on the SPAD surface only;
  Expect transmission gain of 1.3 (from simulations).
Light “Concentrator”
Light Pattern
d-SiPM (50% FF)
Crystal
Optical Grating/
Lenticular Lens
500µm
Crystal
d-SiPM
grating
Notes: grating is miniscule, handling with micro-manipulators and microscopes,

## Slide 33: Diffractive Optics: Tests
CCD
Crystal Matrix (4x4)
Tests with Proteus/AGILE crystal matrix:
interface crystal-grating: air
interface grating-CCD: air
all crystals fully wrapped (3M-ESR Vikuiti)
Glass plate
Bare crystal- matrix & glass
0.71mm
X-Ray Generator
‘Air’ contact
Sony, Philips, Eureka
Notes: CCDs from Sony, Eureka (D), webcam, Philips

## Slide 34: Diffractive Optics: Tests
CCD
Crystal Matrix (4x4)
☛ Average gain in peaks = 1.26
Tests with Proteus/AGILE crystal matrix:
interface crystal-grating: air
interface grating-CCD: air
all crystals fully wrapped (3M-ESR Vikuiti)
Grating
Crystal-matrix & optical grating
25μm
w/o grating
with grating
0.71mm
X-Ray Generator
‘Air’ contact
Sony, Philips, Eureka
Bare crystal- matrix & glass

## Slide 35: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 36: Crystal Performance
Crystal | Size [mm3] | Decay Time [ns] | Light Yield [Ph/MeV] | CTR [ps] FWHM
LuAG:Ce | 2 x 2 x 8 | 60 | 15’000 | 872 ± 50
LuAG:Pr | 2 x 2 x 8 | 20 | 10’400 | 672 ± 30
Crystal performance is “gauged” by light yield (LY) and coincidence time resolution (CTR):
Crystal (2x2x10mm3) | Manufacturer | Decay Time [ns] | Light Yield [Ph/MeV] | CTR [ps] FWHM
LGSO:Ce | Hitachi | 42.50 | 19’500 | 176
LYSO:Ce | SIPAT | 42.60 | 19’000 | 206
LYSO:Ce | Proteus | 41.40 | 18’000 | 175
LYSO:Ce | CPI | 45.00 | 17’400 | 184
LSO:Ce | PML | 47.00 | 16’300 | 206
LSO:Ce | CTI | 43.50 | 23’000 | 190
LSO:Ce:Ca | Agile | 32.00 | 15’700 | 170
CTR measurements were made with the Hamamatsu MPPC S10931-050P;
all crystals were fully wrapped and coupled with grease to the MPPC.
PMT: Photonis XP 2020 Q
Light yield measured with PMT (linear!); crystals wrapped and mounted “small-face-to-PMT” with grease; source: 167Cs.
E. Affray et al., Nuclear Science Symposium and Medical Imaging Conference (NSS/MIC), 2011 IEEE, 10.1109/NSSMIC.2011.6154402

## Slide 37: Crystal Performance
Crystal | Size [mm3] | Decay Time [ns] | Light Yield [Ph/MeV] | CTR [ps] FWHM
LuAG:Ce | 2 x 2 x 8 | 60 | 15’000 | 872 ± 50
LuAG:Pr | 2 x 2 x 8 | 20 | 10’400 | 672 ± 30
Crystal (2x2x10mm3) | Manufacturer | Decay Time [ns] | Light Yield [Ph/MeV] | CTR [ps] FWHM
LGSO:Ce | Hitachi | 42.50 | 19’500 | 176
LYSO:Ce | SIPAT | 42.60 | 19’000 | 206
LYSO:Ce | Proteus | 41.40 | 18’000 | 175
LYSO:Ce | CPI | 45.00 | 17’400 | 184
LSO:Ce | PML | 47.00 | 16’300 | 206
LSO:Ce | CTI | 43.50 | 23’000 | 190
LSO:Ce:Ca | Agile | 32.00 | 15’700 | 170
slow!
UV!
best!
CTR measurements were made with the Hamamatsu MPPC S10931-050P;
all crystals were fully wrapped and coupled with grease to the MPPC.
PMT: Photonis XP 2020 Q
Light yield measured with PMT (linear!); crystals wrapped and mounted “small-face-to-PMT” with grease; source: 167Cs.
E. Affray et al., Nuclear Science Symposium and Medical Imaging Conference (NSS/MIC), 2011 IEEE, 10.1109/NSSMIC.2011.6154402
Crystal performance is “gauged” by light yield (LY) and coincidence time resolution (CTR):

## Slide 38: Small & Large Crystals
0.75 x 0.75 x 10mm3
Light Yield & E-Resolution
16 LYSO crystals, 4 x 4 matrix; 80µm 3M-ESR Vikuiti wrapping; Uniform irradiation with 137Cs; Dry contact with PMT.
3 x 3 x 15mm3

## Slide 39: Small & Large Crystals
0.75 x 0.75 x 10mm3
Light Yield & E-Resolution
16 LYSO crystals, 4 x 4 matrix; 80µm 3M-ESR Vikuiti wrapping; Uniform irradiation with 137Cs; Dry contact with PMT.
3 x 3 x 15mm3
Coincidence Time Resolution
187±27ps FWHM
235±4ps FWHM

## Slide 40: Small & Large Crystals
0.75 x 0.75 x 10mm3
Light Yield & Energy-Resolution
16 LYSO crystals, 4 x 4 matrix; 80µm 3M-ESR Vikuiti wrapping; Uniform irradiation with 137Cs; Dry contact with PMT.
3 x 3 x 15mm3
Coincidence Time Resolution
187±27ps FWHM
235±4ps FWHM
De-convolution yields 212ps FWHM small vs. large (target value: 200ps)

## Slide 41: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 42: Integration: Crystal Matrices
Endoscope Matrix:
Photo of MPPCs courtesy of Hamamatsu Photonics
External Plate Matrix:
80µm 3M ESR Gap
12.4 mm
Crystals: 3 x 3 x 15 mm3
S11827-3344MG
Monolithic (4 x 4) MPPC array:
each with 3 x 3mm2 area
SPAD pitch: 50µm
Fibers: .75 x .75 x 10 mm3
3.4 mm
80µm 3M ESR
Gap
CMOS (4x4) array:
each with 800 x 800µm2 area
SPAD pitch 50µm (x), 30µm (y)
192 TDCs total

## Slide 43: System Integration: External Plate
T-sensor
HV filter
4 x 4 - MPPC
monolithic array on
a custom-made PCB
Sensitive detector: 16 x16 detector units (4096 ch.) with ~100µm gap;
Electronics: dedicated readout chips on PCBs (512 ch./PCB) ;
Cooling: detector stabilized to room temperature via Peltier elements (expected 30mW/ch);
Tracking: detector mounted on a robotic arm for mechanical tracking (6D info, <1mm accuracy);
Curved geometry: facing the organ during diagnose (r = 21cm).
Electronics PCBs
Cooling
23 cm
22 cm
Notes: The design guidelines chosen in this design are:
Modularity: interchangeable parts to reduce costs and required spares
Flexibility: allow parallel development of novel components minimizing final time for integration
Minimum dead material between crystals

## Slide 44: Summary
Frontline research in the domain of:
  digital photodetectors;
  scintillators & optical systems;
  medical instrumentation.
Large knowledge and technology transfer between HEP, industry and medicine;
  Defines a roadmap for the development of a new generation of multimodal endoscopic probes.
Thanks to the endo-TOFET and PicoSEC-Coll.
We still seek ESR applications for the Marie-Curie ITN
  Please contact us or our Marie-Curie-Homepage:
  http://picosec.web.cern.ch/picosec/home.html
Notes: The design guidelines chosen in this design are:
Modularity: interchangeable parts to reduce costs and required spares
Flexibility: allow parallel development of novel components minimizing final time for integration
Minimum dead material between crystals

## Slide 45: End
Notes: The design guidelines chosen in this design are:
Modularity: interchangeable parts to reduce costs and required spares
Flexibility: allow parallel development of novel components minimizing final time for integration
Minimum dead material between crystals

## Slide 46: Detector R&D
1.) Photodetectors
  The analog SiPM (a-SiPM)
  The digital “Endo-TOFPET” SiPM (d-SiPM)
2.) Diffractive Optics
3.) Scintillators
4.) Integration
5.) Readout and Data Acquisition

## Slide 47: Event Rates and DAQ
The hit rate in the external plate is expected to be 10 kHz per crystal. The total hit rate is 40 MHz.
The event rate in the internal probe is of the order of 1 MHz.
The total average data rate from the PET plate is 2 Gb/s. The available bandwidth is 3.2 Gb/s (4 x LVDS links at 800 Mb/s each)
The total average rate from the internal probe is of the order of 1 Gb/s.  Use 2 LVDS links at 800 Mb/s each.
Random coincidence rate is estimated at 250 kHz (assuming 6.25 ns coincidence window). The true coincidence rate is assumed to be of the same order. Therefore we assume that the total trigger rate is 0.5MHz.
For a trigger rate of 0.5MHz the storage rate is relatively modest (about 100 Mbyte/s).
The stored data is reduced by a factor 10 after offline filtering with a coincidence window of 600 ps.
Notes: The design guidelines chosen in this design are:
Modularity: interchangeable parts to reduce costs and required spares
Flexibility: allow parallel development of novel components minimizing final time for integration
Minimum dead material between crystals