## Slide 1: The Silicon Tracking System of the CBM at FAIR: detector development and system integration.
A. Lymanets1,2, E. Lavrik1, H.-R. Schmidt1
University of Tübingen1, Kiev Institute for Nuclear Research2
KINR
TIPP ’14, Amsterdam

## Slide 2: Facility for Antiproton and Ion Research
magnets: 300 Tm bending power
high-intensity DC beam:
109 ions/s at CBM
max. beam energies:
  heavy ions: 45 GeV/u
  protons: 90 GeV
2012: World intensity record for low Charge State Heavy Ions (e.g., U28+)

## Slide 3: CBM experiment
Dipole
magnet
Ring Imaging
Cherenkov
Detector
Resistive
Plate
Chambers
(TOF)
Electro-
magnetic
Calorimeter
Silicon
Tracking
System
Projectile
Spectator
Detector
(Calorimeter)
Micro Vertex
Detector

## Slide 4: Tracking challenge
Central Au+Au@25AGeV UrQMD event
Au+Au at 10 MHz interaction rate
Up to 700 charged particles/evt
Track densities up to 30 cm-2/evt
Need fast and radiation hard detectors as well as fast data acquisition system for online event selection

## Slide 5: Silicon Tracking System
2133k r/o channels
212 FEE blocks at 140 W (10 FEBs with 14 W each) + 60 W (HUB chip)
FEE heat dissipation of 42.4 kW
Total power dissipated by sensors after irradiation ~ 1 W  Convective cooling
212 FEE blocks
106 ladders
8 stations
Cooling R&D:
Open blown CO2 system
1 kW CO2 system TRACI-XL
Plans for 50 kW system

## Slide 6: Open CO2 system
Built to validate heat exchanger design. Currently commissioned and running.

## Slide 7: Experimental setup
Humidity sensor
FEB box with
heating elements
Pressure sensors
Chiller
Induction coil
Control panel

## Slide 8: Heat exchanger studies
FEM calculation (Solid Works)
200 W input power
1 mm thick fins
200 W heat load by resistors
1.8 m long tube with ø1.8 mm (optimal parameters)
Obtain temperature distributions
Check thermal contact
Goal: optimize heat exchanger/cooling fins geometry. Guarantee thermal neutrality (-5 ˚C).
Simulation studies
Prototype

## Slide 9: Very first results
Pin
input power Pin: 0, 45, 80, 100, 140, 200 W
liquid CO2 input temperature Tin: -40 °C
nominal flow to neutralize 200 W:
1 g/s = 27 ln/min
saturation temperature THX is reached at flows > 15 ln/min
Tin
Tout
THeatExchanger

## Slide 10: Very first results II
32 °C
36 °C
42 °C
38 °C
32 °C
18 °C
-34 °C
Flow: 30 ln/s, input power: 200 W
Much higher than -5 °C
(not thermally neutral)
probable cause: insufficient thermal contact between heat exchanger and FEB box

## Slide 11: Simulations studies
Fin width | all aluminum | air gap (100 µm) | thermal grease gap | grease with Cu-FEB Box | grease with all Cu
1 mm | -7.05 | 35.1 | -3.97 | -18.2 | -18.5
2 mm | -21.4 | 24.7 | -18.3 | -26.2 | -26.6
3 mm | -26.9 | 20.9 | -23.4 | -29 | -29.4
simulations confirm experimental observations
good thermal coupling and wider fins (2 mm) should allow complete thermal neutralization
experimental verification underway
Temperature maximum

## Slide 12: Simulations with detailed layout
Temperature maximum °C
fin width | all aluminum | air gap | thermal grease gap | grease with Cu FEB box | grease with all Cu
1 mm | -5.09 | 34.2 | 7.18 | -9.54 | -10
2 mm | -11.4 | 23 | -8.92 | -20 | -20.5
3 mm | -18.4 | 18.7 | -15.7 | -24.1 | -24.6
4 mm | -22.8 | 15.8 | -20 | -26.8 | -27.3
realistic placement of heat sources suggests 2-3 mm fins

## Slide 13: New measurements: improved thermal contact

## Slide 14: New measurements: T @ 100 W
Flow: 20.5 ln/s,
applied heat power: 100 W
SolidWorks simulations predict -19.4 °C highest temperature in this regime
-13.5 °C
-11 °C
-11 °C
-10 °C
-15 °C
-34 °C
-40.3 °C
Solidworks reproduces measurement trends correctly.
Thermal neutrality is achieved with thicker fins.

## Slide 15: TRACI-XL Main components
Condensing unit:
Easily detachable
FU to adapt to different load situations
Technical specifications
Power: 1kW @ -40 ⁰C
Power supply: three-phase
Weight almost 600 kg
Dimensions: 1132x1153x1900
PLC controlled
1132
1153
1900
Control cabinet:
2 Layer (HV + LV)
HMI for friendly User experience
GSI+Tübingen
Transfer line to Remote control box + Experiment
Top Exhaust:
Extract pump + electronics heat
CO2 Front Panel:
Easy removable for maintenance
Armaflex foam insulated
Next slide
Drip Pan:
Tray to detect CO2 leakages
Aluminum profiles frame:
Item profiles 8 40x40 mm
With TRACI-XL you will be able to test 7 FEE blocks of 140 W
For example ¼ Stations 4, 5 or 6 
Thermal box is foreseen to have this size or maybe half station size with dummy loads instead the rest of blocks

## Slide 16: Commercial CO2 power plant
No oil-free setup needed
No particular temperature stability requirement
50 kW externally cooled liquid CO2 system with 1 kW standby power
Company: Hafner-Muschler (http://www.hafner-muschler.de/)
System similar to CERN 2PACL principle with liquid CO2 (not transcritical)

## Slide 17: Conclusions
CO2 based cooling system is required to cool 42 kW of power.
Heat exchanger design has been optimized using simulation approach
Verified by the measurements
1 kW TRACI-XL system is in production. Commissioning to start in July
Industry partners identified for a final STS cooling system.