## Slide 1: Study of the Radiation Damage of Hamamatsu Silicon Photo Multipliers
Wander Baldini
Istituto Nazionale di Fisica Nucleare and
Universita’ degli Studi di Ferrara
Technology and Instrumentation in Particle Physics, Amsterdam June 2-6 2014

## Slide 2: Outlook
Silicon Photo Multipliers
The Louvain Cyclotron Facility
Experimental Setup & Measurements
Results
Conclusions

## Slide 3: Silicon Photo Multipliers
SiPMs are recently developed semiconductor photo detectors composed by a matrix of Geiger Mode Avalanches Photo Diodes (GMAPD)
The working point is a few Volts above the breakdown voltage
The output signal is the sum of the charge produced by each pixel
The shape of the active area and the dimension of the pixels can be customized according to user’s needs

## Slide 4: Silicon Photo Multipliers
Charge spectra
High gain ≈ 106
No sensitivity to magnetic fields
Small-sized
Single photon resolution
Low bias voltage (<100V)
Drawbacks:
High dark count (typically ≈ 100s kHz at room temperature, 0.5 p.e. threshold )
  greatly improved in the last few years
Radiation damage
    requires extensive studies for applications in high radiation environment (like High Energy Physics)
Main topic of this presentation

## Slide 5: The Louvain Cyclotron 110 Facility
50 MeV deuterons on Be target
Continuous spectrum up to 50MeV
Peak at ≈23MeV
Filters for g and low energy neutrons (98% purity)
Neutron fluence precisely controlled
Cooling box up to -25 C°
n

## Slide 6: SiPM Hamamatsu Sample
We tested 16 devices of 8 different types and 3 categories:
  A, B: conventional MPPC like devices
  C,E,F,G,H: “low after pulse” devices
  D: special device with a new structure
All devices had the same geometry:
  1x1 mm2 active area on plastic support
  50x50 mm2 pixel size
Two devices for each type

## Slide 7: Measurements
For each SiPM we measured:
  I-V characteristic curve
  Acquired a sample of 100k waveforms, from which we:
    Estimate the dark rate
    Measure the Gain
The above measurement were done after each irradiation step:
  Before irradiation
  5 x 108 1MeV-eq n/cm2
  5 x 109 1MeV-eq n/cm2
  5 x 1010 1MeV-eq n/cm2
The temperature was monitored and the measurements corrected offline
I-V curve
acquired waveform
dark spectra

## Slide 8: Experimental Setup
The SiPMs were mounted on a custom made PCB:
  Bias the SiPMs
  Connect the SiPMs with the DAQ
  Monitor the SiPMs temperature (4 pt100 probes)
The SiPMs were protected and light tightened through an aluminum cap fixed on the PCB
I-V characteristic curves were measured with a Keithley pico-Ammeter
Waveforms were acquired thanks to a CAEN Digitizer (2Gs, 12 bits)
SiPM holder
Neutron beam line
n
SiPMs

## Slide 9: Results: Radiation Effect at a Glance
The radiation damage is very clear looking at the dark noise:
Before irradiation
5 x 108 neq/cm2
5 x 109 neq/cm2
5 x 1010 neq/cm2

## Slide 10: Results: I-V Characteristic Curves
Standard type:
Low afterpulse
New structure
Low afterpulse
In general all the devices show an increase of a factor ≈100 in the dark current
The breakdown point become less sharp and starts a few volts earlier
No significative differences among the devices

## Slide 11: Results: Dark Spectra vs Dose
Before irradiation
5 x 108 neq/cm2
5 x 109 neq/cm2
5 x 1010 neq/cm2

## Slide 12: Results: Gain vs Dose
Before irradiation
5 x 108 neq/cm2
The Gain can be measured as long as the p.e. peaks are visible, i.e.
up to 5 x 109 neq/cm2
5 x 109 neq/cm2

## Slide 13: Results: Gain vs Dose
Standard type
Low afterpulse
New structure
The Gain reduction is ≈ 6.5% for the standard and low afterpulse types, about 11% for the new structure type
Initial gain of the new structure SiPM was ≈ 10% higher
Temperature effects have been corrected
G0 ≈ 33 mV/p.e.
G0≈ 31 mV/p.e.
G0≈ 36 mV/p.e.

## Slide 14: Results: Dark Noise
The rate was measured counting the average number of signals above 0.5 p.e. in the acquired time window (2 ms)
The dark count rate increases ≈ factor 10 at each irradiation step
No significative differences among the devices
standard device
low afterpulse
0.5 p.e. threshold
0.5 p.e. threshold

## Slide 15: Conclusions
16 SiPM Hamamatsu have been irradiated on a 23MeV neutron beam in Louvain-la-Neuve
The devices were of 3 types: standard MPPC-like, low afterpulse, new structure
The integrated 1MeV equivalent doses were:
  5 x 108 1MeV-eq n/cm2
  5 x 109 1MeV-eq n/cm2
  5 x 1010 1MeV-eq n/cm2
After each step and for each device we acquired:
  the characteristic I-V curve
  a sample of 100k waveform, to measure gain and dark count
According to our measurements, no significative differences in the main parameters have been observed
A more detailed analysis is ongoing

## Slide 16: Spare

## Slide 17
NIEL (Non Ionizing Energy Loss) curve for silicon