## Slide 1: Real-time Imaging of prompt gammas in proton therapy using improved Electron Tracking Compton Camera (ETCC)
TPC
Pixel Scintillator Arrays (PSA)
RI reagent
Electron Tracking Compton Camera (ETCC) & its Medical application
Improvement in new ETCC in SMILE-II
Result from Beam test
Summary
T.Tanimori1 H.Kubo1, A.Takada1, S.Iwaki1, S.Komura1, .Matsuoka1, , K.Miuchi2, T.Mizumoto1, Y.Mizumura1, K.Nakamura1, S.Nakamura1, M.Oda1. J.D.Parker1, T.Sawano1 S.Sonoda1, D.Tomono1, S.Kurosawa3
1)Dep. of Phys, Kyoto Univ., Kyoto, Japan, 2) Dep.of Physics Kobe University, Japan,
3) Institute of Material Research, Tohoku University, Japan
3/June/2014 TIPP2014@Amsterdam

## Slide 2: Electron Tracking Compton Camera (ETCC)
W(I-131:364keV)
2
direction of photon by photon with SPD useful for Noise cut. & Clear Imaging
Noise Reduction by Kinematics(a)+dE/dx
Large FoV >4str(good for monitoring)
Modular structure
Event Selection
  dE/dx background cut
dE/dx
SMILE-I (2006)
10cm-cube ETCC
Old Medical ETCC based on SMILE-I (2006) balloon model
New one based on SIMLE-II (2013)
gas
mPIC with 400mm pitch
GSO pixel scintillator
Gas TPC

## Slide 3: Variety of RI applications in ETCC
 | Ce-139 | Cr- 51 | Ba-133 | I- 131 | Au- 198 | Na- 22 | F- 18 | Cu- 64 | Cs-137 | Mn-54 | Fe- 59 | Zn-65 | Co- 60
Energy [keV] | 167 | 320 | 354 | 364 | 410 | 511, 1275 | 511 | 511 | 662 | 835 | 1095, 1292 | 1116 | 1173, 1333
Life | 137.6 day | 27.7 day | 10.52year | 8.01 day | 2.6 day | 2.609 Ｙear | 109.8 min | 12.70 hour | 30.04 year | 312.1 day | 44.5 day | 244 day | 5.271 year
Energy dynamic range : 167 – 1333 keV.
ETCC
Zn-65-Porphyrin (1.2MeV)
SPECT
PET
W (364keV)
6cm
Cal.source
Ba. 335keV
Zn-65-Porphyrin (1.1MeV)
I-131 364keV
Rainbow： 511keV（FDG)
Orange： 365keV(MIBG)
(I-131:364keV)
Thyroid gland phantom from one direction
Notes: 3
3
3

## Slide 4: Improvement from 1stMedicalETCC
1st Medical 10cm-cube ETCC
Observation Time: ~3 hours
Pos. Res.: 8-5mm(FWHM)
(8o at 662keV GSO)
Energy range: 150keV-2MeV
FoV：2 0x20cm＠10cm front
Max. Trigger rate <30Hz
Due to ~10% tracking efficiency
& slow Electronics
~1 m
GSO
GSO
TPC
GSO
30cm
proton
electron
Tacking. Eff. -> 100% in SMILE-II
Fast electronics x 100 and more
Good Angular resolution ->5o GSO
SMILE-II (2013)
30cm –cube-ETCC
SMILE-I
New Med. ETCC Head

## Slide 5
  New Tracers using new RI with its decay time similar to biological decay time; Visualizations of immunity and enzyme : (FDG for visualization of metabolism)
  Multi RI Tracer Image 99mTc +18F
  3D imaging from one directional observation for operation supporting
  Imaging for beam therapy (proton, ion, neutron)
Features of ETCC for Molecular Imaging & Nuclear Medicine

## Slide 6: Imaging Test in 30cm ETC
Noise reduction by Energy loss rate dE/dx
Energy Cut
⇒ Continuum fully gamma events selected by dE/dx cut
Diffuse gamma
137Cs(～0.85MBq)
Preliminary

## Slide 7: Performance of 30cm-cube ETCC
ARM 5.3o (@662 keV)
Old ETCC ~10o
Reliable efficiency
(Good consistency with MC )
Good ARM res.
Simple Analysis (dE/dx)
High contrast image
large Eff.Area
15 deg.
30 deg.
60 deg.
90deg.
Wide Field of View (6 sr)
LaBr3 use 10cm front
99mTc -> ~5mm
511keV-> ~2mm
Aｒ１．５atm +2 RL.GSO
Eff.Area 1cm2 =10-2 efficiency in 10cm ETCC
CF4 3atm
Preliminary
Preliminary

## Slide 8: Miss reconstruction events in Compton Camera
ETCC perfectly remove miss reconstructed events
A: Multi scattering in F.D.
B: electron escapes from F.D.
Ｃ：electron hits to B.D.
Forward detector
Backward Detector
E2
γ
E1
Recoil electron hits B.D.
Forward detector
Backward Detector
Recoil electron escapes
E2
γ
E1
γ
E1
Forward detector
Backward Detector
E2
γ
E1
Mult scattering in F,D.
A
B
C
ETCC easily detects one recoil electron fully contained in TPC
Usual CC hardly distinguish
those events from good Compton events

## Slide 9: Imaging Improvement by SPD
(SPD=200°)
Ryan, J. M., NewAR, 48, 199 (2004).
Usual CC
:ETCC
⇒ ～3times better contrast image
RI
artifact
Preliminary
no use Electron track
use Electron track

## Slide 10: First Proton Therapy Beam-on Imaging @2009
511keVγ
800-2000keVγ
Bragg Peak
First Imaging of continuum gammas at Beam-on by ETCC using SMILE-I technology.
But intensity <0.01nA &
Trigger rate <50Hz !!
Kurosawa et al. (2012)
800-2000 keV
>2000keV
511 keV
All energy
Simulation
Notes: 10

## Slide 11: New 30cmETCC beam test using 140MeV p
Ring Cyclotron @RCNP Osaka Univ.
・Similar intensity of neutron to gammas (~1/10)
Liquid Scintillator
keV
gamma
neutron
Scintillator
ETCC
TPC
6x6 PSAs
3x6 PSAs
30cm
GSO Scintillator
Shield Plate
150cm
30cm
137Cs
0.7MBq
water
Plastic
Scintillator
140MeV
Proton
>0.1nA
Liquid
Scintillator
70cm
ETCC
Proton beam
H2O Target
140MeV p beam Intensity >01.nA
Trigger rate 0.3-1kHz
p +n
MIPs
ETCC dE/dx

## Slide 12: Imaging test in intense BG environment
↑ Top
beam
Water Target
> 1500 keV
Bragg peak
beam
511 keV ± 10%
ETCC Energy Spectrum
dE/dx cut
1MeV
Direct gamma-ray Back Projection Image
Raw data
gamma
662 keV
background 1/10 by ｄE/dx cut
Sexcess @ 511, 662 keV
137Cs
Beam line
ETCC detected gamma ray correctly with same efficiency as no beam.
3MeV
Trigger rate ~500Hz
Preliminary
Preliminary
Measured time 14min.
~2000 gammas >1.5MeV
150cm

## Slide 13: What are needed in Practical use
Beam Therapy Imaging
Present Status
-> 3 gammas/sec >1.5MeV
Development factor s
1. Distance 1.5m ->40cm x 10
2. Efficiency develop. x 50
3. Beam Intensity x ~10
Total 5x103
Expected gammas >1MeV
1.5x104 /sec.
Expected Trigger
500x (10x50) x10= ~106 Hz
But
Energy threshold >0.8 MeV -> 1/4
+ limited Field of View x1/4
=> Expected Trigger rate 105 Hz
DAQ Improvement
VME bus (~2KHz) -> network x >102
2KHz x 102 ~2x105Hz
Multi heads => several 104Hz
Possible by present hardware !
Bragg peak position determination
1cm position resolution for 1 gamma
-> 100 gammas for <1mm accuracy per second
at least several x102 gammas/sec >1MeV is needed !

## Slide 14: Imaging for Boron chemicals in BNCT
10B+n -> 7Li* +a (94%)
7Li* -> 7Li +g(478keV)
Imaging of Boron chemicals(BC)
BC+ 18F for PET
However, no guarantee for BC+18F with same distribution as BC in body
PET image
Only ETCC may image this g under neutron background !
Imaging System of BC in mice is inevitable for developing new BCs
Solution:
Small neutron source ( 3MeV proton LINAC+ modulator) + ETCC
Boron Chemicals (BCs)

## Slide 15: Way to Practical Use of ＥＴＣC
GOAL for Medical ETCC head in Practical Use
Position Resolution -> 2-3mm in human
Observation Time -> 10 minutes for human
High Efficiency -> ~1% (3atm CF4 gas in 10cm ETCC)
High counting rate >50KHz at one head
3-6 ETCC ring heads system for CT imaging
Near future developments
New Medical ETCC is now OK( ~1KHz operation) and 2nd one soon
Following tests using 2 ETCC heads will be done in 1-2 years.
Multi tracer 3D image with 99mTc and FDG in mouse
10B imaging in mouse
High through put rate imaging under intense radiation for particle therapy
Detail of SMILE=II is presented in 4th evening
in II.b Astro & Space session1
New Medical ETCC head

## Slide 16: Observation of time variation
ETCC
ETCC
No source
~20 min
54Mn (835 keV, 1 MBq)
~3 m
~3 hours
0.025
0.020
0.015
0.010
0.005
SMILE-II clearly detected
gamma-ray source with only 10 minutes.
~10 min/frame
835 keV±10%
30˚
60˚
+0.01 mSv/h

## Slide 17: Difference in Noise reduction
Noise reduction in Classical CC
Only Cut for E1 or E2, or E1/E2
Anyway, some of real gamma events are lost
E0
E1
E2
q
Noise reduction in ETCC
Only cut for dE/dx,
dE/dx is not used for event reconstruction,
Good dE/dx reduces backgrounds without loss of real gamma events
γ
E1
dE/dx
E2
TPC