## Slide 1: TRAINING COURSE on radiation dosimetry:
Instrumentation 1 –
Gas detectors / Part 1
Anthony WAKER,
University of Ontario Instutute of Technology
Wed. 21/11/2012, 15:00 – 16:00 pm, and 16:30 – 17:30 pm

## Slide 2: Gas-Filled Detectors
One of the oldest and most widely used radiation detectors
Gas-filled detectors sense the direct ionization created by the passage of charged particles caused by the interaction of the radiation with the chamber gas
      Ion Chambers
      Proportional Counters
      Geiger-Mueller Counters

## Slide 3: Basic Components of AN Ionization chamber
Common Fill Gases: Ar, He, H2, N2, Air, O2, CH4, TE
Notes: Basic elements of I.C. - enclosed volume of gas in region where E-field is applied by an external voltage.
Basic process – radiation interacts with gas molecules generating e-ion pairs. E-field separates positive and negative (e) ions and are collected at the electrodes. Measuring the current generated by these pairs – IONIZATION CURRENT due to radiation interaction with sensitive gas volume.

## Slide 4: Ionization in Gases
  To create an ion pair, a minimum energy equal to the ionization energy of the gas molecule must be transferred
  Ionization energy between 10 to 25 eV for least tightly bound electron shells for gases of interest
    Competing mechanisms such as excitation leads to incident particle energy loss without the creation of ion pair
  W-value: average energy lost by incident particle per ion pair formed
  Typical W-values are in the range of 25 – 35 eV/ion pair
Notes: b/c incident particle also loses energy by other mechanisms, ex. excitation

## Slide 5: Charge Collection
  Under steady-state irradiation, rate of ion-pair formation is constant
  For a small test volume, rate of formation will be exactly balanced by rate at which ion pairs are lost from volume due to recombination, diffusion or migration from the volume.

## Slide 6: Charge Collection
For ions in a gas,
  +ve and –ve charges are swept towards their respective electrodes with
  By increasing concentration of ions within the gas volume decreases suppressing volume recombination within the gas volume.
Notes: E/p – reduced electric field
Ex. At 1 atm, E=10^4 V/m, v_drift ~1m/s This is a very long transit time in a detector volume with cm range, Transit time = 10 ms.
So what can we do? - Collect electrons – much faster drift velocities.

## Slide 7: Recombination
Two types of recombination:
  Columnar (or initial) recombination
      Increases with LET of Radiation
  Volume recombination
      Increases with dose-rate
Notes: Recombination rate, alpha – recombination coefficient, n+ density of +ve ions, n- density of –ve ions
Columnar – at incident particle track
Volume – after diffuse away from track location.

## Slide 8: Charge collection

## Slide 9: Ionization Chamber – FARMER Chamber

## Slide 10: Ionization Chambers
What ionization current would we expect to measure in a Farmer chamber placed in a high energy photon radiotherapy beam where the dose rate to air is 10 Gy per minute.

## Slide 11: Insulators and Guard Rings
  Typical ionization currents are extremely small.
  Require good insulation and guard rings to ensure leakage current does not interfere with ionization current
Notes: Insulators – high resistivity synthetic plastics
Radiation damage can occur when exposed to high radiation fields – use ceramics in detector design b/c ceramics have a higher resistance to radiation damage.

## Slide 12: Ionization chamber dosimetry

## Slide 13: Dosimetry with ionization chambers

## Slide 14: Fano’s Theorem
In an infinite medium of given atomic composition exposed to a uniform field of indirectly ionizing radiation, the field of secondary radiation is also uniform and independent of density of the medium, as well as density variations from point to point.
This means that if an ionization chamber is constructed of a wall material and filled with gas of the same atomic composition the dose to the wall material will be the same as the dose measured to the gas regardless of the size of the chamber

## Slide 15: Ionization Chamber Dosimetry - calibration

## Slide 16: Ionization Chamber Dosimetry

## Slide 17: Dosimetry with gas ionization devises
TISSUE EQUIVALENT PROPORTIONAL COUNTERS TEPC

## Slide 18: Atomic Composition of tissue and TE gas
  ICRU Tissue (Muscle) atomic composition by % weight
H | C | N | O
10.2 | 12.3 | 3.5 | 72.9
Methane based
  CH4 (64.4% partial pressure)
  CO2 (32.4% partial pressure)
  N2 (3.2% partial pressure)
  By %weight: H (10.2); C (45.6); N (3.5); O (40.7)
Propane based
  C3H8 (% partial pressure)
  CO2 (% partial pressure)
  N2 (%partial pressure)
  By %weight: H (10.3); C (56.9); N (3.5); O (29.3)

## Slide 19: Tissue equivalent plastic
A150 TE-plastic atomic composition by % weight
  The main tissue equivalent plastic used in dosimetry is A150.
  The atomic composition of A150 is close to tissue but has a higher percentage by weight of carbon, which makes it conductive.
H | C | N | O
muscle (10.2) | muscle (12.3) | muscle (3.5) | muscle (72.9)
10.1 | 77.6 | 3.5 | 5.2

## Slide 21: Gas-gain in proportional counters
  A proportional counter is a gas-ionization device consisting of a cathode, thin anode wire and fill-gas. Ionization in the fill gas is multiplied providing an amplified signal proportional to the original amount of ionization.

## Slide 22: Gas Gain
  The gas-gain achievable in a proportional counter is determined by the first Townsend coefficient α for the counter fill gas used
  α itself depends on the reduced electric field in the counter, which is determined by the applied voltage and counter geometry

## Slide 23: Gas Gain
  To a first approximation the relationship between the logarithm of gas-gain and applied anode voltage is linear
Chart: Relative Gas Gain for Propane Based TE Gas at Pressures 3.25 (graph 1), 6.5 (2), 16.25 (3), 26 (4), 32.5 torr, Relative to the Measurement Made at 32.5 Torr and Vanode 100 V
| 100 | 200 | 300 | 400 | 500 | 600 | 700
 | 2.01 | 3.1 | 4.23 | 5.27 | 6.36
 | 0.85 | 1.71 | 2.71 | 3.76 | 4.79 | 5.82 | 6.88
 | 0.52 | 1.15 | 1.94 | 2.82 | 3.76 | 4.72 | 5.71
 | 0.3 | 0.85 | 1.5 | 2.29 | 3.1 | 3.99 | 4.93
 | 0 | 0.55 | 1.25 | 1.98 | 2.78 | 3.62 | 4.5

## Slide 24: Simulation using a gas cavity

## Slide 25: Site-size simulation
  Energy deposited in the gas cavity by a charged particle crossing the cavity equals energy deposited in tissue site by an identical particle
Et
Eg

## Slide 26: Site-size simulation
tissue
gas

## Slide 27: Site-size simulation
  The density of the gas in the cavity is adjusted to equal the ratio of the tissue site diameter to the gas cavity diameter
Density of Gas
Diameter of Gas Cavity
Density of Tissue Site (1000 kg.m-3)
Diameter of Tissue Site

## Slide 28: example
  What is the density of propane TE gas required for a 1 cm cavity to simulate a tissue sphere of 1 µm.
= 10-4 . 1000 kg.m-3
= 0.1kg.m-3

## Slide 29: example
  What pressure of propane TE gas is required for a density of 0.1 kg.m-3
At 20 oC:
(41.72 torr)

## Slide 30
Rossi Counter

## Slide 32: Tepc calibration
Internal alpha sources

## Slide 33: Internal source calibration
  In crossing the TEPC an alpha particle will lose an average amount of energy that can be calculated using range-energy data

## Slide 34: Internal source calibration
  Each alpha particle crossing the counter generates a pulse height proportional to the energy deposited; the mean of this distribution is associated with the mean energy deposited in the counter

## Slide 36: example
Applied Voltage 750V; amplifier gain 10
Mean energy lost 168.48 keV
Mean chord-length 1.33 µm
Channel 3835 corresponds to 126.68 keV/µm
Calibration Factor for amplifier setting of 10 (126.68/3835) = 0.03303 keV/µm/channel
  Using range-energy data for propane based TE gas for a 2 micron simulated diameter and a Cm-244 internal alpha source of energy 5.8 MeV.

## Slide 37
Frequency distribution measured in an Am-Be field with a 2” REM500 TEPC with simulated diameter 2µm and calibration factor 1.641keV/µm/chn

## Slide 38
Frequency x lineal energy: Dose Distribution d(y) with calibration factor 1.641keV/µm/chn

## Slide 39
y.f(y) data plotted in equal logarithmic intervals, 50 per decade

## Slide 41: TEPC measurable quantities

## Slide 42: Measureable Quantities – Ambient Dose Equivalent
Estimated directly from the measured event-size spectrum
Determined from the shape of the event-size spectrum and assuming Q(y) = Q(L)

## Slide 43: TEPC Measureable Quantities – Absorbed Dose

## Slide 44
From ICRP 60

## Slide 45
Assuming that Lineal Energy y is equal to Linear Energy Transfer L

## Slide 46: TEPC Response – Kerma

## Slide 47: TEPC ENERGY RESPONSE – Quality Factor

## Slide 48: Detectors other than tepcs
Gas Electron Multipliers

## Slide 49: GEMs

## Slide 50: Gas Electron Multiplier
  Operates as proportional counter except multiplication takes place between the top and bottom surfaces of the GEM structure through microscopically etched holes

## Slide 52
Dubeau and Waker Radiat. Prot. Dosim. 128(4), 2008

## Slide 53
Dubeau and Waker Radiat. Prot. Dosim. 128(4), 2008

## Slide 54: Gem summary
  GEMs can be configured to operate as TEPC and have the advantage of
    smaller physical size for each detecting element and smaller simulated diameters for improving dose equivalent response (better LET spectrometers)
    Potential for basis as a personal neutron dosimeter
  Particle tracking capability depending on read-out pattern of anode
  Much work still to be done!