## Slide 1: TRAINING COURSE on radiation dosimetry:
Instrumentation 2 – Solid state detectors
Hybrid pixel detectors and their applications
Thilo MICHEL, University of Erlangen-Nuremberg
Thu. 22/11/2012, 10:00 – 11:00 am

## Slide 2: Agenda
The Medipix/Timepix detector
Energy response
Application: Imaging spectroscopy at high flux
Application: Material resolved imaging
Application: Dosimetry
2

## Slide 3: The Medipix2/3 and Timepix detectors: hybrid photon counting pixel detectors
ASIC/Sensor:
  Development: International Collaboration with seat at CERN
  Bump-bonded with Pb/ Sn
  65536 pixels in 256 columns and 256 rows
  Pixel pitch: 55 µm
  Size of the matrix: 14 mm (2 cm2)
  0.25 µm CMOS
Sensor:
  Materials: Si, GaAs, CdTe
  Bias voltage: 150 V (300 µm Si)
  2x2-version (Quad)
14 mm
E
3

## Slide 4: Step 1a: photoabsortion + relaxation
GaAs (Z=31, 33)
Photon
10 keV
12 keV
Energy not suffcient to remove e- from K-shell
Longer penetration depth
Production and detection of fluorescence photons and Auger electron emission
Smaller penetration depth
Broader transversal profile
x [mm]
y [mm]
z [mm]
x [mm]
y [mm]
z [mm]

## Slide 5: Step 1b: Continuous energy loss of electrons
Contiouous energy loss: Bethe-Bloch-Fomula
For silicon: electron path length approximately
Electron-hole pairs are generated

## Slide 6: Step 2: Drift and diffusion of released charge carriers in semiconductor
Charge generation term
Charge carrier density change
Drift term
Diffusion term
Limited lifetime term
Transport of charge carriers described by differential equation

## Slide 7: Step 2: Drift and diffusion of charge carriers
Electric field lines
Electric potential
Holes for example
Electrons for example
+ 150 V (for example)
Pixel electrodes at approx. 0 V
Charge might be shared between several pixels (charge sharing, charge splitting)
Attention: Do not forget the second charge carrier type

## Slide 8: Step 3: signal generation in pixel electrode
Charge above electrode influences mirror charge in all pixel electrodes
Charge carrier moves -> mirror charge changes
A current flows in or out the pixel electrode
+
Pixelelectrode at is put to norm weighting potential
Other electrodes to 0:
Solve Laplace equation:
Obtain weighting-Field:
Weighting potential
Influenced current
Attention: electrons AND holes influence both a current with same polarity!!!

## Slide 9: Step 3: The weighting field above a pixel electrode
Electric field
Weighting field
Central pixel
Neighbours
Attention: The charge carrier type drifting to pixels contributes most!!!

## Slide 10: Step 3: the pulse-forms of the influenced currents
Electric field
I [µA]
Q [C]
t [ns]

## Slide 11: Step 3: Pulseform depends on point of generation of charge carriers
I [µA]
Start near pixel electrode
Drift time t [ns]
Start near common electrode

## Slide 12: Step 4: integrate current and process digitally
Pixel electronics
  Charge-Sensitive Preamplifier
  1 discriminator (minimum threshold approx. 1000 e- = 3.6 keV in Si and 4.4 keV in CdTe)
  1 counter per pixel
Operation modes
  Counting
  Time-Over-Threshold
  Time-Of-Arrival
Pixel electronics
Principle
14 bits
Readout zone block diagram
Pixel cell
12

## Slide 13: Advantages and disadvantages of the Medipix detector in X-ray imaging
Noiseless dark field image
No readout noise
No blooming
No afterglow
Very linear at low rates
High position resolution
High frame rate
Energy sensitive
Advantages
Small size
Charge-Sharing
Multiple counting reduces DQE
Rate limitations at high flux (1010 photons/cm2/sec)
Disadvantages
13

## Slide 14: Working principle
Analogue
Threshold
V
t
*by Winnie Wong
14

## Slide 15: Working principle
Analogue
Threshold
V
t
Counting (Medipix2,3 & Timepix)
Time of Arrival (ToA) (Timepix)
Time over Threshold (ToT) (Timepix)
End frame
Start
Frame
15

## Slide 16
16
Deposited energy in keV
Pixelwise energy calibration for ToT mode necessary
Notes: 16

## Slide 17: Agenda
The Medipix/Timepix detector
Energy response
Application: Imaging spectroscopy at high flux
Application: Material resolved imaging
Application: Dosimetry
17

## Slide 18: Does it measure the impinging photon energy ?
Detector response affected by
Collected charge
Intensity
Electronic noise
Eγ
Compton scattering
EFl.
Eγ – EFl.
(Escape peak)
Fluorescence photon production
Charge Sharing
EFl.
18

## Slide 19: Threshold scan for 60 keV photons onto silicon sensor
19

## Slide 20: Differentiated threshold-scan = response function
Photoeffect
Diffusion, electron track length, Compton scattering
20

## Slide 21: Response of a CdTe Medipix to X-rays
In pixel deposited energy in keV (from tresholdscan)
Number of events
* Measurement
Simulation
Primary photon
energy
Noise edge
21
1.6 mm CdTe sensor
(220 µm pitch)
Fluorescence lines of Cd and Te
Escape- peaks

## Slide 22: Simulation of response has to take many effects into account
Elektromagnetic interactions
Secondary radiation (fluorecences)
Tracking on µm-scale
Fano-Noise
Charge-Carrier respulsion
Isotropic diffusion during drift
Limited charge carrier lifetime
Electron+hole signal contribution
Electric field in doped CdTe
Electrode size, pitch (Weighting potential)
Electronics noise
Threshold dispersion
Physics models needed for:
Experiments* / Simulations
1.6 mm CdTe, 700 V
*) Experiments are performed with a CdTe detector made by the Freiburger Materialforschungszentrum
22

## Slide 23: Response matrix for silicon sensor
Complete set of response functions
23

## Slide 24: Agenda
The Medipix/Timepix detector
Energy response
Application: Imaging spectroscopy at high flux
Application: Material resolved imaging
Application: Dosimetry
24

## Slide 25: Spectrum of deposited energies is a convolution of response with impinging spectrum
Measurement (Threshold scan)
Response-Matrix
Impinging spectrum
Possibilities to get impinging spectrum:
Pseudo Inverse Matrix
Spectrum-Stripping
Bayesian deconvolution
25
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 26: Differentiated threshold scans can be used for deconvolution, BUT ...
Differentiate
26
Deconvolution
Measurement
Simulation
Result
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 27: …it is statistically better to use threshold scans directly in Bayesian deconvolution
with
27
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 28: A difficulty: threshold dispersion among pixels
Measurement (N40 X-ray quality)
Global Kalibration of thresholds
Image at 42,5 keV threshold
28
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 29: Improvement by pixelwise threshold calibration
Global calibration (64k pixels and 1 calibration)
Pixelwise calibration
(64k calibrations)
29
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 30: Spectral imaging with a brachytherapy X-ray tube
Tube-collimator: 10 cm
Collimator-detector: 30 cm
Hole in collimator: 50 µm
Collimator thickness: 1 mm
=> Magnification: 2.99
Acc. voltage: 50 kV
Current: 40 µA
Measurement time: per threshold: 60 s overall: 4,4 h
30
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 31: Camera obscura principle
Timepix
Collimator
e-
2.8 mm
g
Tube
31
Brachytherapy- X-ray tube
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 32: The electron beam moves around
32
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 33: Look from the side: Misalignment of the target …
g
33
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)

## Slide 34: … causes differences in the spectrum
34
Courtesy: Peter Sievers (Physikalisch Technische Bundesanstalt PTB)
left ROI
right ROI

## Slide 35: Agenda
The Medipix/Timepix detector
Energy response
Application: Imaging spectroscopy at high flux
Application: Material resolved imaging
Application: Dosimetry
35

## Slide 36: Principle of Material Reconstruction
X-Ray transmission through a compound object
mass attenuation coefficient
areal density
index for basis materials
Detector
source
36

## Slide 37: A method of material reconstruction without explicit spectrum reconstruction
Number of counts in energy deposition interval
Sum over energy depositions in interval
Sum over primary energies in spectrum
Sum over materials
Vary to get best agreement of Mb and measurement
37

## Slide 38: Measurement details
Tested with the MARS scanner together with the University of Canterbury, Christchurch, New Zealand
Medipix2 MXR, 300 µm Si sensor
Energy calibration with 241Am, Gd- and Mo K-edges
Threshold equalisation mask generated with flatfield at THL approx. 35keV and not at the noise floor
Source: 0.1mA @75kV tungsten anode X-Ray tube
Acquisition at 4 THL values: 8, 20, 33 and 45keV with 2, 3, 4 and 5s duration respectively
3 detector positions and 180 projections (360 degrees)
38

## Slide 39: Experimental setup
Medipix2 MXR
Mouse inside
Perspex tube
0.5mm Al Filter
X-Ray source
39

## Slide 40: Material Reconstruction: Projections
Photon counting image
Non-iodine (water) image
Iodine image
40

## Slide 41: Material Reconstruction in CT: Brightness = density of base materials (g/cm3)
Photon counting at
8 keV threshold
water /
non iodine
iodine
Brightness = areal density of materials
41

## Slide 42: Agenda
The Medipix/Timepix detector
Energy response
Application: Imaging spectroscopy at high flux
Application: Material resolved imaging
Application: Dosimetry
42

## Slide 43: The personal dose equivalents Hp(10) and Hp(0.07)
Photon energy dependence of personal dose equivalents 1)
Definition of the personal dose equivalents
Dose:
Hp(d): dose in soft tissue in a certain depth d [mm] of the ICRU slab phantom (30 x 30 x 15 cm3 PMMA)
To be measured with an active personal dosemeter:
  Hp(10): Deep dose
  Hp(0.07): Surface dose
1) Data compiled from ICRU-Report 57, 1998
Hp(d) / Fluence [pGy x cm2]
Hp(0.07) / Fluence
Photon energy [keV]
Hp(10)
Hp(0.07)
Example: Hp(0.07) = 1 nSv
  10 keV: 140 photons/cm2
  60 keV: 2000 photons/cm2
  200 keV: 800 photons/cm2
43

## Slide 44
44

## Slide 45: Solution: estimation of the dose equivalent with the measured number of counts in energy intervals
Monochromatic irradiation with Ej
Calibrate with jmax different photon energies
Apply pseudoinverse
of the „counts matrix“
Known doses
Number of counts in each energy deposition bin
Unknown calibration factors from counts to dose
Measure/simulate the number of counts in energy deposition intervals
Best estimation (maximum-likelihood) of the calibration factors
Energy deposition [keV]
Counts
N1(60)
N2(60)
N3(60)
45

## Slide 46: Calculation of conversion factors and energy bins
Assess quality of dose reconstruction on test spectra
Coupled set of equations
Apply pseudoinverse matrix
Iterate
E1
E2
E3
E4
E5
N1
N2
N3
N4
N5
Collection of monoenergetic responses
E1
E2
E3
E4
E5
N1
N2
N3
N4
N5
46

## Slide 47: Extension of energy range beyond 200 keV only possible with filters above pixel detector,
Uncovered
Filter 2
Filter 1
Reason: energy deposition spectra look the same for higher energies due to dominance of Compton scattering in silicon
low energies
medium energies
high energies
47

## Slide 48: Response to perpendicular irradiation
3 detectors with 2 filters and 3 x 16 energy bins
IEC limit (+67%)
IEC limit (-29%)
48

## Slide 49: Angular dependence of deep dose Hp(10; α)
Hp(10,0°)/Fluence in pSv*cm2
Hp(10, a)/Hp(10, 0°) @fixed fluence
49

## Slide 50: Under angle of incidence the sensor effectively becomes thicker and smaller
Number of counts at 60 degrees / Number of counts at 0 degrees
Photon energy in keV
50

## Slide 51: Dose reconstruction with 220 micron pixel at high dose rates
Doserate of 1 Sv/h:
111 kHz @14 keV -> Dtmean = 9 μs
ToT-duration@14 keV = 1.4 μs
-> significant pile up expected
A
Number of counts /
pixel / μSv
51

## Slide 52: Pile-up in preamplifier (analog side of electronics)
Two events above threshold are recorded as one event but with higher energy
Two events originally below threshold are detected by coincidence
52

## Slide 53: How does pile-up influence the measured spectrum ?
Deposited energy in keV
Number of counts
1 Sv/h
2 Sv/h
0 Sv/h
53

## Slide 54
New pixel detector Dosepix
54
Non-linear ToT(E)
Feedback from digital to analog pixel electronics
Dynamic range of test pulsers too small
Sensitive to temperature changes
Power consumption too high
Pixel suffer from charge sharing (too small)
More functionality needed in pixel
Dead time during readout
3.6 mm
Readout area
3.6 mm
Disadvantages Timepix for dosimetry
New detector:
16 x 16 Pixel with 220 µm x 220 µm area
Notes: 54

## Slide 55: May 2011: the first sample without sensor
55
3.5 mm

## Slide 56: Pixel electronics (CERN, IBA, Uni Erlangen)
Patent: M. Campbell (CERN), W. Wong (CERN), X. Llopart-Cudie (CERN), R. Ballabriga-Sune (CERN), T. Michel, G. Anton, M. Böhnel, : „Radiation Monitoring Device'', EP 07114122.0 (2007)
56

## Slide 57: Hybrid photon counting pixel detectors
... have their roots in high energy physics (e.g. ATLAS, CMS, ALICE,...)
... can have intelligent pixel cell electronics
... are used in combination with Si, CdTe, GaAs sensor layers
... can measure energy and time of arrival
... can have high position resolution
... can process events in high flux individually
... allow spectrometry in high flux fields
... allow identification of particle types by track signature
57