## Slide 1
Heinz Graafsma
Head of FS-DS; DESY-Hamburg
and University of Mid-Sweden; Sundsvall
3D Needs and Opportunities
in X-ray Detection.

## Slide 2: The Detector Challenge:
1900 1960 1980 2000
PETRA-3
Second generation
First generation
X-ray
tubes
ESRF (2000)
ESRF (1994)
Synchrotron Sources
brilliance

## Slide 3: But is there hope ?
Europe has a long tradition in developing detectors for science; at least in part thanks to CERN
Europe has a large number of Storage Ring and FEL sources; which are centrally run  this allows for centralized and larger scale instrumentation development
There is a surprisingly large number of Hybrid Pixel Array Detector systems; mostly photon-counting
PSI/SLS and CPPM-Marseille over 10 years ago started to use HEP HPAD detector technologies for dedicated photon science systems.

## Slide 4: (some) major running systems:
PILATUS + EIGER
XPAD
MYTHEN + GOTTHARD
PIXIRAD
pnCCD
MEDIPIX
….

## Slide 5: (some) major systems under development
MEDIPIX
  LAMBDA (DESY)
  EXCALIBUR (DIAMOND)
AGIPD (for EXFEL) (3 German + 1 Swiss partner)
LPD (for EXFEL) (2 British partners)
DSSC (for EXFEL) (7 German + Italian partners)
JUNGFRAU (for SWISSFEL) (PSI)
PERCIVAL (for low-E FEL applications) (DESY, STFC, Elettra, Diamond)
…

## Slide 6: The FEL-Challenge: Different Science
Completely new science
Fast science 100 fsec
“Single shot” science
x109

## Slide 7
K. J. Gaffney and H. N. Chapman, Science 8 June 2007
What is the Goal ?

## Slide 8: The European Free Electron Laser

## Slide 9
4.5 MHz
What are the Challenges ?

## Slide 10
The Adaptive Gain Integrating Pixel Detector
08/04/2013
C1
Leakage comp.
C2
C3
Discr.
Control logic
Trim
DAC
Vthr  VADCmax
Analogue encoding
Normal Charge sensitive amplifier
High dynamic range:
Dynamically gain switching system
Extremely fast readout (200ns):
Analogue pipeline storage

## Slide 11
At the P10 beamline
Beam direction (coming from sample)
AGIPD sensor
Chip tester box
It took about 1 ½ hours to set up, after about 2 hours we saw the first image
Not in the picture: Sample, Alexanders PC, people, …

## Slide 12
Looking at the direct beam
photons/second
Logarithmic color scale
Gain switching experimentally proven
104 photons / pulse
Single photon sensitivity
4.5 MHz frame rate

## Slide 13
Some XPCS measurements
Scientific quality data obtained
Complete system proven to work
Calibration proven to be adequate
photons/second
Logarithmic color scale

## Slide 14
The Analogue Storage
Sensor
ASIC per pixel
ASIC periphery
Chip output driver
Mux
HV
+
-
THR
DAC
SW
CTRL
Analog Mem
Analog Mem
CDS
RO Amp
Calibration circuitry
Adaptive gain amplifier
352 analog memory cells
…
…
RO bus (per column)
Pixel matrix
Notes: Mention random vetoing possibility.

## Slide 15
AGIPD 1.0 Pixel electronics
200 x 200 micron2 pixels
352 storage cells + veto possibilities.
Minumum signal (noise) = 330 electrons = 0.1 photon of 12 keV
Maximum signal = 33 106 electrons = 104 photons of 12 keV
4.5 MHz frame rate
64 x 64 pixels per ASIC
2 x 8 ASICs per module (128x512 pixels, no dead area)
4 modules per quadrant

## Slide 16
Smaller pixels with
  Reduced dynamic range (no gain switching);
  Reduced number of frames?
New Technologies including:
  3D CMOS: second layer (more functionality)
What could be next?

## Slide 17
3D-CMOS
1) bump-bond to the detector (TSV)
2) amplification & double sampling
3) store the analog signal in a memory cell in 1st/2nd layer (544 memory depth)
4) signal readout
5x5 mm2 double-tier dice
MPW, via CMP
submitted 2011
to be delivered early 2013
Standard Cells

## Slide 18
Smaller pixels with
  Reduced dynamic range (no gain switching);
  Reduced number of frames?
New Technologies including:
  3D CMOS: second layer (more functionality)
  65 nm CMOS technology (more functionality)
  3D CMOS + 65 nm CMOS
What could be next?
Notes: HGF Portfolio activity

## Slide 19: HGF-Detector Program; Pillar-1 : The “Helmholtz-Cube”
What we have today
Replace standard sensor with: 3D and edgeless sensors, as well as High-Z
Replace standard bump-bonds with new interconnect techniques
Replace standard ASIC and wire bonds with thinned ASICs and TSV as well as ball-grid arrays
Replace standard ASIC with 3D-ASICs
Develop new high speed IO’s
Employ innovative materials

## Slide 20
new sensors
Thinned and 3D ASICs
Innovative materials
Advanced interconnect
High-speed I/O
HGF-Detector Program; Pillar-1 : The “Helmholtz-Cube”
What we want tomorrow