## Slide 1: Electro-Optic Bunch Profile Monitor for the CERN-CTF3 probe beam
R. Pan, T. Lefevre , S.P. Jamison, W.A. Gillespie
CERN
STFC Daresbury Laboratory
University of Dundee
BI Day----Convention Centre of Archamps
6th of Dec. 2012

## Slide 2: Outline
Summary
Resolution
EO bunch profile monitor system
EOSD principle and Simulation
Introduction for CALIFES

## Slide 3
Energy | 200 MeV
Train duration max. | 140ns
Train rep. rate | 5 Hz
Bunch length | 8-10 ps (1.4 ps)
Bunch rep. rate | 1.5 GHz
Bunch charge | 0.1–0.05nC (0.6 nC)
4
Introduction ---- CALIFES
1
2
3
5
6
4
single klystron
Standing-wave photo-injector
3 travelling-wave structures (LIL)
RF network
Complete set of diagnostics
to Two Beam Test Stand (TBTS)
Existing bunch profile monitor:
Deflecting cavity (bunch head downward, tail upward)
Bunch length measurement with an acceleration structure (bunch head decelerated, tail accelerated)

## Slide 4
Coulomb field of e-bunch
Coulomb field temporal profile
β=0
β=0.9
β=0.9999
High energy , Coulomb field temporal profile is approximately the bunch temporal profile
Broadening of profile:
Coulomb field of one electron

## Slide 5
Simulation: Coulomb field of e-bunch
Coulomb field temporal profile and broadening
Radial offset from single electron
Electrons’ density distribution within one bunch
Convolution
For high energy beam (200 MeV):
Broadening rate < 2% @ 10 mm
2%
r0: the distance far away from e-bunch
Closer
Further
Damage to crystal
Crystal survives, low Coulomb field

## Slide 6
Simulation: EOSD
Electro-Optical Spectral Decoding:
Linear chirped optical pulse
Polarization variation caused by Coulomb field—laser nonlinear effect
Polarization  Intensity, by two crossed polarizers
I(λ)  I(t)
---- rotation matrix
----Jones matrix for half waveplate
----Jones matrix for quarter waveplate
EO bunch profile measurements:
--EO spectral decoding
--EO temporal decoding
--EO spatial encoding
--EO up conversion

## Slide 7
EO effect:
Coulomb field
Simulation: EO phase mismatching
Laser
Frequency mixing
Polarization variation
induce
Wider bandwidth, better resolution !

## Slide 8
Simulation: EOSD results and limitation
Other parameters:
Laser wavelength: 780nm
Laser pulse energy: 1.5nJ
Pulse duration: 150fs
Crystal thickness: 1mm
Distance: 5mm
distortion
Short bunch----fast temporal modulation---- spectral content
---- t~λ mapping

## Slide 9
EO monitor Design for CALIFES
Stretcher:
SF11 glass
100fs----3.24ps
Actuator
Cameras
Rotation stage

## Slide 10
Laser system
Laser system
Custom Erbium fiber laser with pulse picker
Laser head and controller
Fundamental wavelength | 1560 nm
Laser output power | > 350 mW
Pulse width | < 120 fs
Repetition rate | 37.482 MHz
Second-harmonic wavelength | 780 nm
Laser output power | > 120 mW
Pulse width | < 120 fs
Beam size (1/e2) | Typ. Ø 1.2 mm (780 nm) Typ. Ø 3.5 mm (1560 nm
Beam divergence | < 1 mrad (780 nm) < 2 mrad (1560 nm)
Synchronization box
PD in
RF 1
RF 2
Phase out
PID in
Piezo control
Motor control
PID: Proportional Integral Differential
75 MHz
37.5 MHz

## Slide 11
OTR
e-
ICCD
Camera
Optical system
Streak
Camera
Photodiode
(5 ns)
3 ns min. exposure
signal (fibre, 20m, ~100 ns)
Delay stage (6 fs step, 0.75 ns range)
Amplifier
Oscillator
Pulse picker
Pulse generator
RF (74.963 MHz)
RF (74.963 MHz, 13.3ns)
RF (37.4815 MHz, 26.6 ns)
Phase shifter
5 Hz rep.
Trigger from the accelerator
Ref (74.963 MHz)
Timing and synchronization scheme

## Slide 12
EO monitor Design for CALIFES
Chamber design
<100>
<010>
<001>
<-110>
<-1-10>
First chamber includes a mirror
The second chamber includes a mirror, a crystal and a OTR screen
Crystal orientation
OTR

## Slide 13
Laser/OTR Transfer Line
We transfer the 780nm fs laser from our laser lab (building 2010-1-002) to CLEX, and transfer the OTR photons from CLEX to our lab.
Laser lab
Streak camera room
CLEX
CLEX

## Slide 14
Laser/OTR Transfer Line 2
Installation:
Take the box1 and box2 for example
Mirror/lens cover box

## Slide 15
Laser/OTR Transfer Line 3

## Slide 16
Expected Resolution
1. Distance between crystal and e-beam
~10 fs at r=5 mm
2. The frequency response of crystal (material and thickness)
for 1 mm ZnTe: ~333 fs ~1/(3THz)
3. EOSD limitation (Laser pulse duration and chirped duration)
~550 fs (100 fs3 ps)
4. Resolution of spectrometer and CCD ~40 fs (512 pixels)
Grating

## Slide 17
Summary & Outlook
EO technique is a non-destructive testing technique, and has the feasibility in fs resolution.
Based on numerical simulation, the resolution of this system is expected to be sub-picosecond.
The system waiting for being installed during the winter shut down. First measurement in the beginning of next year.
Summary:
Outlook:
This research project has been supported by a Marie Curie Early Initial Training Network Fellowship of the European Community’s Seventh Framework Programme under contract number (PITN-GA-2008-215080-DITANET)
Resolution improvement study
System simulation, designs of laser system, transfer line, chambers, detection system and timing and synchronization have done.

## Slide 18
Thank You !

## Slide 19
[1] X. Yan, A. M. MacLeod, W. A. Gillespie, G. M. H. Knippels, D. Oepts, A. F. G. van der Meer and W. Seidel, Phys. Rev. Lett. 85, 3404 (2000).
[2] I. Wilke, A. M. MacLeod, W. A. Gillespie, G. Berden, G. M. H. Knippels, and A. F. G. van der Meer, Phys. Rev. Lett. 88, 124801 (2002).
[3] A. Azima, S. D¨usterer, H. Schlarb, J. Feldhaus, A. Cavalieri, D. Fritz, and K. Sengstock, In Proceedings of the EPAC 2006, Edinburgh, Scotland, 2006, p. 71.
[4] G. Berden, S. P. Jamison, A. M. McLeod, W. A. Gillespie, B. Redlich, and A. F. G. van der Meer, Phys. Rev. Lett. 93, 114802 (2004).
[5] S. P. Jamison, G. Berden, P. J. Phillips, W. A. Gillespie, and A. M. MacLeod, Appl. Phys. Lett. 96, 231114 (2010).
[6] J. D. Jackson. Classical Electrodynamics (John Wiley and Sons, Inc., 1999).
[7] S. P. Jamison, A. M. MacLeod, G. Berden, D. A. Jaroszynski, W. A. Gillespie, Opt. Lett. 31, 1753 (2006).
[8] S.P. Jamison, G. Berden, W.A. Gillespie, P.J. Phillips, A.M. MacLeod, in Proceedings of EPAC08, Genoa, Italy, 2008, p. 1149.
[9] J. Breunlin, Diploma thesis, University Hamburg, 2011.
[10] B.Steffen et al., Phys. Rev. ST Accel. Beams 12, 032802 (2009).
[11] A. Yariv. Quantum electronics (1989).
References