## Slide 1: The new BLM System for the Injector Complex
BI-BL:
M. Alsdorf; M. Kwiatkowski; O. Bitterling; W.Vigano’; C. Zamantzas
BI-SW:
E. Angelogiannopoulos; S. Jackson
Project Team
Support from BI-BL:
E. Effinger; J.Emery; E. Nebot Del Busto; S. Grishin; G. Venturini
BI Day 2012

## Slide 2: Introduction
The design of the new Beam Loss Monitor System is needed to upgrade the current system on the Injectors.
Project Objectives:
=> Build a generic, highly configurable and high performing BLM system.
=> Design the Acquisition part able to accept several detector types.
=> Use reprogrammable parts to target all injectors’ requirements.

## Slide 3: System Architecture

## Slide 4: Acquisition Crate
Acquisition module (BLEDP)
Up to 8 modules with 8 channel each
Main panel
Ref. current Input
LEDs
Power switch
Control Unit
Later version w/ advanced remote functions
Custom Backplane
See next slide

## Slide 5: Backplane of the Acquisition Crate
CUSTOM BACKPLANE
Routing of 64 input cables to the 8 BLEDP cards.
48Vdc Power Distribution and protection by Circuit Breakers .
Redundant Power supply single failure tolerant.
Current Supply monitoring.
Internal calibration current Source.
Relays Multiplexer.
Geographical addressing .
Chip ID.
Temperature and Relative Humidity measurement.
Low Leakage design.
Standard Vias and True hole footprint
Modified Vias and True hole footprint
Hole
High Insulation Capacitor

## Slide 6: Keeping Down the Power Supply Noise
Acquisition Module
Ac/Dc Converter
Backplane
Protections+ Filters
48VDC
48VDC
FPGA + Logic
Analog
5VDC
Filter
Filter
Filter
LVDS
Filter
220VAC
Filter
Dc/Dc 1.2VDC
LDO 2.5VDC
Filter
Filter
Several PCB layers are used as ground planes.
The star point connection is done on the Dc/Dc converters’ ground.
Dc/Dc
3.3VDC
Dc/Dc
5VDC
3.3VDC
Filter
Filter

## Slide 7: Acquisition module (BLEDP)
Backplane connection
Analogue inputs, power and control
SFP connectors
Gigabit optical and/or Ethernet links
FPGA
Altera Cyclone IV
JTAG connection
Local programming and diagnostics
Acquisition
digitisation of 8 channels
Currently verifying version 2 of the
printed circuit board
Circuit Breakers
Stop the failure propagation
Local power supplies
Separated for Digital and Analog

## Slide 8: Acquisition principle (FDFC & DADC)
The input channel circuit is able to measure current input from 10pA to 200mA.
The measurement of the current input is performed by two different techniques:
Fully Differential Current to Frequency Converter (FDFC) used in the range 10pA to 30mA
Direct ADC acquisition (DADC) used in the range 80µA to 200mA
No gain change required: The switch between the 2 ranges is managed by the FPGA.

## Slide 9: Acquisition Channel Architecture

## Slide 10: FDFC data processing
The number of accumulated counts are combined with the ΔADC values to calculate the integrated loss over a 2 μs period.
0
Time
counts
2 μs
2 μs
2 μs
ADC values
thhigh
thLow

## Slide 11: FDFC Data Merger
Algorithms
Two types of data acquired: Counts & ADC values.
Algorithms have been implemented to merge those values for all possible acquisition cases.
Next step currently under testing: implementation of further algorithms to reduce the noise.

## Slide 12: Simulation Model
PSPICE simulation model of the FDFC
Models built using imported files
from component’s supplier

## Slide 13: Algorithm Verification with Matlab
PSpice Output
Analog Signal + 2 Comparator Pulses
Data created by PSPICE
Data interpolated by MATLAB
FDFC out converted in Ampere (Current value)
Algorithm Verification with Matlab
Full correspondence with reality
MODIFICA!!!!

## Slide 14: System Survey and Diagnostic Reader
System Survey and Diagnostic Reader
100% of the System Functions have Diagnosis
Read-out status and diagnostic information from connected chips or sensors with every timing event.
All readouts are afterwards pre-stored, packaged and forwarded to logging.
Additional interfaces can be easily added to the reader.
All readouts can be monitored for failsafe actions.

## Slide 15: Processing Mezzanine (BLEPM)
2 SFP modules
gigabit Optical links and Ethernet
FPGA
Cyclone IV 150K logic elements
Connectors
to the mainboard
DAB64x
JTAG connection
Local programming and diagnostics
SRAM
Dc/Dc converter
And filters

## Slide 16: Acquisition & Processing
Integration | Minimum detectable current | Minimum detectable dose [Gy/s] | 
 |  | Detector: IC | Detector: LIC
2 μs | 31.250 nA | 5.787E-04 | 7.234E-03
1 ms | 62.250 pA | 1.153E-06 | 1.441E-05
400 ms | 0.156 pA | 2.889E-09 | 3.611E-08
1200 ms | 0.052 pA | 9.630E-10 | 1.204E-08
Different Processing per target.
Selectable integration periods.
E.g. : 2 µs, 400 µs, 1 ms and 1.2 s
Threshold values unique per channel

## Slide 17: 3D Models

## Slide 18: Test Installation in the PS
5 type of detectors
are connected to BLEDP cards
Pre-Series Manufacturing

## Slide 19: Test Installation in the PS
Ethernet Hub +
JTAG Chain programmer

## Slide 20: Acquisitions
Ionization Chamber
PEPII on MU16
500.000 acquisition points

## Slide 21: Thank you