## Slide 1: The HIE-ISOLDE Faraday Cup
BI Day
December 6, 2012
Alejandro Garcia Sosa
W. Andreazza, E. Bravin, E. Daniel Cantero, M. Fraser, Y. Kadi, D. Lanaia, D. Voulot, F. Wenander

## Slide 2: Contents
Introduction to HIE-ISOLDE
The short diagnostic boxes
Review of Faraday cups
The HIE-ISOLDE Faraday cup
Simulations
Experimental Tests
Conclusions

## Slide 3: 1. Introduction to HIE-ISOLDE
Energy upgrade
10 MeV/u
Construction of SC LINAC
+ service buildings
Intensity upgrade
LINAC4+PSB
Design Study of target area, Class-A lab and beam lines
Beam quality upgrade
RFQ cooler and buncher
Solid state lasers for RILIS
Higher mass resolving
power HRS

## Slide 4: 1. Introduction to HIE-ISOLDE
Energy upgrade
10 MeV/u
Construction of SC LINAC
+ service buildings
Intensity upgrade
LINAC4+PSB
Design Study of target area, Class-A lab and beam lines
Beam quality upgrade
RFQ cooler and buncher
Solid state lasers for RILIS
Higher mass resolving
power HRS

## Slide 5: 1. Introduction to HIE-ISOLDE
Energy upgrade
10 MeV/u
Construction of SC LINAC
+ service buildings
Intensity upgrade
LINAC4+PSB
Design Study of target area, Class-A lab and beam lines
Beam quality upgrade
RFQ cooler and buncher
Solid state lasers for RILIS
Higher mass resolving
power HRS

## Slide 7
2. The short diagnostic boxes

## Slide 8
Class 100
ISO 5
Short diagnostic box
Vacuum valves
Steerer magnet
2. The short diagnostic boxes
Notes: Intertank distance: 320 mm from valve to valve, 420 including the flanges of the tanks.

## Slide 9
Currently being developed by CERN and Added Value Solutions (AVS) with the support of the Spanish government (CDTI)
The device fits in the 90 mm long space available and contains
  A 45 degree scanning blade with 2 slits (H, V)
  HIE-ISOLDE Faraday cup
  Solid-state detector (Absolute energy, ToF)
  Collimators blade
  Attenuating and stripping foils
2. The short diagnostic boxes

## Slide 10
Beam

## Slide 14: REX vs HIE - DBs
REX - Faraday cup (length: 59 mm)
MCP (beam profile - image) or Si detector (energy and TOF)
Collimators wheel (attenuating and stripping foils as well)
HIE - Faraday cup (14 mm)
Scanning slits (beam profile)
Collimators (optics)
Stripping foils in some cases
Solid-state detector in a few boxes (absolute energy, TOF)
The most important device that needs to be tested is the FC

## Slide 15
3. Review of Faraday cups
Destructive measurement
Measures absolute beam current
The escape of electrons increases the value read in the picoammeter.
Ion-induced electron emission:
  Low energy electrons (Ee < 20 eV)
  High energy electrons (Ee ~ keV for MeV/u ions)

## Slide 16: 4. The HIE-ISOLDE Faraday cup
30
14
26
59
(Distances in mm)
REX-ISOLDE
HIE-ISOLDE
ISAC 2

## Slide 17: Electrons captured and retained Intensity: solid angle of the signal plate
In a simplified model: W2  retained electrons W1  lost electrons
Signal plate geometry

## Slide 18: 5. Simulations
REX-ISOLDE Faraday cup potential distribution -60 V

## Slide 19: 5. Simulations
REX-ISOLDE Faraday cup secondary e- tracking

## Slide 20: 5. Simulations
More uniform potential distribution
Enhanced effectiveness
HIE-ISOLDE Faraday cup: -500V +100V

## Slide 21: 5. Simulations

## Slide 22: 6. Experimental tests

## Slide 23: Definition of I0
stable beam A/Q =4
Vsignal plate = 0 V
Vrepeller < 0 V (variable).
Signal Plate
Repeller

## Slide 24: Biasing the repeller ring
stable beam A/Q =4
Vsignal plate = 0 V
Vrepeller < 0 V (variable).
Signal Plate
Repeller ring
Ground ring
The short cup does not measure the same current value and needs much higher repelling voltage.

## Slide 25: Biasing the signal plate
stable beam A/Q =4
Vsignal plate > 0 V (variable).
Vrepeller = 0 V
Signal Plate
Repeller ring
Ground ring
High leakage current
Beam current does not reach a current plateau increasing the voltage

## Slide 26: Biasing repeller ring & signal plate
Vsignal plate > 0 V (variable)
Vrepeller = -500 V
Signal Plate
Repeller ring
Ground ring
stable beam A/Q =4

## Slide 27: 7. Conclusions
Current measurements do not agree with the nominal beam current using the present design
Further improvements in the design are in progress
Beam profile measurements are unaffected in principle by a change in the design.
Notes: Beam profile measurements are unaffected in principle by a change in the design.

## Slide 28: Acknowledgements
The ISOLDE Collaboration
The HIE-ISOLDE Project Team and groups within CERN Accelerator and Technology Sector
The Swedish Knut and Alice Wallenberg Foundation (KAW 2005-0121)
The Belgian Big Science program of the FWO (Research Foundation Flanders) and the Research Council K.U. Leuven
The CATHI Marie Curie Initial Training Network: EU-FP7-PEOPLE-2010-ITN Project number 264330.
The Spanish Programme “Industry for Science” from CDTI

## Slide 29: Thank you!
Work supported by the CATHI Marie Curie actions under contract GA-PITN-2010-264330.