## Slide 1: BI Day 2012 The Beam Position System for LINAC4
Michael Sordet BE-BI-PI
Linac4 BPM, 06.12.2012

## Slide 2: Presentation index
5. décembre 2012
2
Michael Sordet BE/BI
Linac4 accelerator
BPM on Linac4
Mechanical details
Electronic details
Results
To-do list
Conclusion
Questions

## Slide 3: Linac4 accelerator
5. décembre 2012
3
Michael Sordet BE/BI
Ion species H-
Output energy 160MeV
Bunch frequency 352.2MHz
Max. repetition rate 2Hz
Beam pulse length 400µs
Mean pulse current 40mA
Number of particles per bunch 1.14 ·109

## Slide 4: BPM layout and specs for Linac4
5. décembre 2012
4
Michael Sordet BE/BI
Line | # of Monitors | Beam position Resolution | Beam position Accuracy | Rel. Beam Intensity Resolution | Beam phase Resolution | TOF Resolution | Comments
L4D | 2 | 0.1mm | 0.3mm | 1% of peak current | 0.5 | 1 per thousand | Ø34mm
L4C | 7 |  |  |  |  |  | Ø39mm
L4P | 6 |  |  |  |  |  | Ø39mm
L4T to PSB | 10+17 |  |  |  |  |  | Ø100 + 140mm
Observable 1: Beam position
Observable 2: Beam intensity
3 MeV
50 MeV
102 MeV
160 MeV
Transfer Line to PSB
H- source
LEBT
RFQ
MEBT (chopper)
DTL
CCDTL
PIMS
Dump Line
L4L ~ 12 m
L4D ~18 m
L4C ~22 m
L4P ~14 m
L4T- LT- LTB-BI ~178 m
L4Z ~10 m
Observable 3: Beam phase in reference to RF
Observable 4: Energy via TOF between two BPMs

## Slide 5: Transfer function of the test stand BPM
5. décembre 2012
5
Michael Sordet BE/BI
Shorted stripline with added capacitance
Linear on working range
Better for linear phase measurement
b-dependent
Ø67mm

## Slide 6: Dealing with low-b beams
5. décembre 2012
6
Michael Sordet BE/BI
PU sensitivity changes with b
These simulations confirm Shafer’s theorem (1994)
The values obtained from the simulations will be implemented in the software
Transfer Line BPM Ø140mm
RFQ
DTL1
DTL3
PIMS12
Β=1

## Slide 7: Many different BPM variants
5. décembre 2012
7
Michael Sordet BE/BI
DTL
Beam aperture: 34 mm
Electrode length: 117 mm
TRANSFER LINE
Beam aperture: 100 mm
Electrode length: 140 mm
3MeV
Beam aperture: 60mm
Electrode length: 67mm

## Slide 8: Lab test bench
5. décembre 2012
8
Michael Sordet BE/BI
Wire technique with standing waves
Optical sensor for mechanical centering : 0.01mm
352MHz sine wave with a N.A.
Data acquisition with steps of 0.1mm
Gives sensitivity and electrical offset

## Slide 9: BPM sensitivity at 352MHz
5. décembre 2012
9
Michael Sordet BE/BI
Body Shorted Striplines
@352MHz | Simulations | H plane | V plane
Sensitivity[mm] | 75.4 | 75.88 | 75.24
Elec. Offset [mm] | - | -0.31 | 0.34
Transfer Line BPM Ø140mm

## Slide 10: Acquisition chain
5. décembre 2012
10
Michael Sordet BE/BI
FEATURES
Input Low pass filtering @1GHz
Down-mixing with LO
Variable gain selection
Band-pass filtering @22MHz
FPGA for calibration and gain control
Sampling frequency = 4x IF
ENOB ADC 13.4 bits
LSB ADC 76µV
A
D
C
1
6
B
i
t
s
FESA
Class

## Slide 11: Signal processing
5. décembre 2012
11
Michael Sordet BE/BI
Phase and intensity calculation with I/Q sample per channel
Calculation for each PU
  Position
  Phase
  Intensity

## Slide 12: Calibration procedure
5. décembre 2012
12
Michael Sordet BE/BI
Set the phase and the gain per channel
  Inject sine wave into the direct channel(1)
  Numerical gain and phase correction
Set the length of cables per PU
  Inject sine wave into the cable(2)
  Signal back after reflection on the PU(3)
  Numerical gain and phase correction

## Slide 13: Calibration : raw data
5. décembre 2012
13
Michael Sordet BE/BI
4 electrode-calibration
Each channel space with 100µs
Direct and reflected sine wave pulse
I/Q sample series
With artificial offset
100µs
100µs
100µs
Direct Pulse
Reflected Pulse
H+
H-
V+
V-

## Slide 14: Calibration : processed data
5. décembre 2012
14
Michael Sordet BE/BI
4 electrode-calibration
Direct and reflected sine wave train
One magnitude point for one quadruplet (I,Q,-I,-Q)
Gives a set of scaling factors per electrode:
[Amplitude, Phase]
[A1..4, P 1..4]
Direct Pulse
Reflected Pulse
H+
H-
V+
V-
800ns

## Slide 15: Commissioning with proton beam
5. décembre 2012
15
Michael Sordet BE/BI
Injection
lines
Linac4
Linac2
LTB.UMA20
Transfer lines from Linac-2 to Booster
Upgrade with Linac4 type BPM : shorted stripline(present BPMs are 40 years old)
Status : commissioning of one monitor with protons since Sept. 2012
Bunching frequency 202MHz

## Slide 16: First raw beam data
5. décembre 2012
16
Michael Sordet BE/BI
TOF beam
4 channels
Electronics card output
Difference between simulation and real life : Amplitude is 45% less than expected

## Slide 17: Position measurement
5. décembre 2012
17
Michael Sordet BE/BI
Horizontal position
Vertical position
New BPM measurement
Old BPM measurement
Calibration pulse
Beam
Artifact
TOF beam
40µs
20µs Head

## Slide 18: Intensity measurement
5. décembre 2012
18
Michael Sordet BE/BI
BPM measurement
BCT measurement
40µs
150mA average
TOF beam
Beam current [A]
Calibrate with BCT

## Slide 19: Phase measurement
5. décembre 2012
19
Michael Sordet BE/BI
First phase measurement
5° phase jump
Phase between one BPM and the RF reference

## Slide 20: To-do list
5. décembre 2012
20
Michael Sordet BE/BI
Finish commissioning test with the LTB-UMA-20
Be ready for H- beam in the 3MeV test stand in January 2013
LS1change all the old L2 BPM and install new acquisition system
Be ready for H- beam test in L4 tunnel in September 2013

## Slide 21: Summary and conclusion
5. décembre 2012
21
Michael Sordet BE/BI
Encouraging first results with beam
Very useful test with Linac2 for the future commissioning of the Linac4
Not expected position measurement excursion, we need to do a MD to test that
Intensity measurement with 45% less signal at the output of the BPM
5° phase jump for phase measurement to be understood and corrected

## Slide 22: Acknowledgements
5. décembre 2012
22
Michael Sordet BE/BI
Project leader Jocelyn Tan
Michael Ludwig for the software
Delphine Gerard for the mechanics
RF group
All the BI-PI section
Drawing office EN/MME
Thank you for your attention
QUESTIONS ???

## Slide 23
5. décembre 2012
23
Michael Sordet BE/BI
Extra Slide

## Slide 24: Sensibility measurement
5. décembre 2012
24
Michael Sordet BE/BI
Constant sum between ±15mm
Linear difference between ±15mm
Linear electrodes between ±15mm
The PU is linear on the desired range

## Slide 25: Analog card block diagram
5. décembre 2012
25
Michael Sordet BE/BI

## Slide 26: Lab measurement
5. décembre 2012
26
Michael Sordet BE/BI
Dynamic range
Noise measure
Gain & Bandwidth

## Slide 27: Analog board characteristics
5. décembre 2012
27
Michael Sordet BE/BI
 | Measure
Gain | 0dB to 30.8 dB
Central frequency | 22.01 MHz
Bandwidth | 10 MHz
Output range | ±2.68 V
Output thermal noise | 700 µV
Crosstalk Signal Cal to IF 22 MHz | -50.4 dBm/0.7mV
Crosstalk LO 330 MHz | -57.2 dBm/0.3mV

## Slide 28: Acquisition chain
5. décembre 2012
28
Michael Sordet BE/BI

## Slide 29: Dealing with low-b beams
5. décembre 2012
29
Michael Sordet BE/BI
Sensitivity at 352MHz
b
Transfer impedance changes with b and bunch length
PU sensitivity changes with b
These simulations confirm Shafer’s theorem (1994)
The values obtained from the simulations will be implemented in the software

## Slide 30: Intensity and position measurement
5. décembre 2012
30
Michael Sordet BE/BI