## Slide 1: OUTCOME OF FIRST CAVITY TESTS
With help of T. Junginger, N. Schwerg, K. Liao, P. Maesen, M. Gourragne, A. Benoit, O. Brunner
SPL Seminar 2012 at CERN

## Slide 2: OUTLINE
Cavity Preparation for RF test
  - Antenna adjustment
  - Hydrogen desorption
  - Electro polishing and rinsing
  RF facilities at SM18
  - 2K RF test stand
  - Cavity and diagnostic system mounting on the insert
  - RF test control room
  - Automatic measurement program development
  RF measurement setup
  RF tests on Saclay (β=0.65) and SPL (β=1) 704MHz mono-cell cavities
  Integration of the 5 cells on the new vertical RF test stand
  Outlook

## Slide 3: Cavity Preparation for RF test

## Slide 4: SPL MONO CELL ANTENNA ADJUSTEMENT
• Input Antenna:
Linput = 17 mm
QExtInput = 3.75E+9
REF: SPL Mono-Cell Antenna Adjustment, N. Schwerg, T.Junginger BE/RF, EDMS 1240961
Measurement setup
• Pick-up Antenna:
Lpick-up = 86 mm
QExtPick-up = 2.11E+11
fres = 704.1972MHz

## Slide 5: HYDROGEN DESORPTION TREATMENT
Bake out under Vacuum
24 h at T = 650°C after reach the maximum outgasing of hydrogen
p = 1.1 E-7 mbar at the end of the process
Ref: A. Mongeluzzo TE/VSC

## Slide 6: ELECTRO-POLISHING AND HPWR RINSING
1st : 31 μm
Tinlet 11.6C Toutlet 12.5C
Time 7h
Potential 12.2V I= 29 to 50A
2nd : 32 μm
Tinlet 9.5C Toutlet 10C
Time 7h17
Potential 12.2V I = 18 to 45A
3nd : 22 μm
Tinlet 10.5C Toutlet 11.5C
Time 5h
Potential 12.2V I = 20 to 45A
Nb removed 100 μm (Spec 200 μm)
TOC > 40ppb
Resistivity 18MOhm.cm
Pcav = 1.4e-2 mbar
Electro polishing setup
Rinsing setup
Rmq: Observation of cracks after electropolishing (J. Chambrillon, O. Capatina)
Ref: L. Ferreira TE/VSC

## Slide 7: RF facilities at SM18

## Slide 8
2K VERTICAL TEST STAND AT SM18
 = 1 mono-cell mounted on the insert
Cryostat (Height: 2.4m) cover with RF and cryogenics connections
New cryostat 4m next year

## Slide 9
Thermal sensors
1- TT850
2- TT851
3- TT852
4- TT816 (520mm from the first screen)
5- TT813 (370mm from the first screen)
6- TT810 (1340mm from the first screen)
7- TT811 (700mm from the first screen)
8- TT814 (430mm from the first screen)
9- TT815 (270mm from the first screen)
Heaters
10- Heater DN100 EH817 (connection AB (2-6)) (R=90)
11- Heater EH818 (connection CD (5-6)) (R=100)
Thermocouples
12- TC818 (thermocouple CD, 5-6 connectors)
13- TC817 (thermocouple AB)
He Level
14- LT 810 (L=1150 mm)
15- LT811 (L=1150 mm)
LT811 = 760mm in order to cover the cavity
CRYOGENIC INSTRUMENTATION
2
1
3
4
14
1080
30
8
9
6
7
5
10
10
11
13
12
15
1420
210
690
670

## Slide 10
leak
Second mounting
Parameters checked before cool down:
Leak detection: 1.10-8 mbar l/s
Pcav= 2.10-7 mbar at 25°C
fres = 704.1972 MHz at 25°C
CAVITY MOUNTING FOR RF TEST
We found a Leak with a rate of 10-6 mbar l/s
We opened the cavity in the clean room ISO 5
Pb Copper Gasket
No time for a new HPWR (due to the cryo shutdown in SM18)
We lost the benefit of the HPWR for the RF test

## Slide 11: OSCILLATING SUPERLEAK TRANSDUCER
20 OSTs were installed to the equator and the Iris
No cavity quench during the mono-cell RF measurement
Localized hot spot (default) with 2nd sound propagation in He superfluid method
Ref: K. Liao BE/RF

## Slide 12: RF TEST FACILITY CONTROL ROOM

## Slide 13: TEST CONTROL PANEL PROGRAM
Improvement: automatic control of the PLL, τ measurement
Program was done by A. Benoit BE/RF

## Slide 14: RF measurement setup

## Slide 15: RF MEASUREMENT SETUP
A
PFWD
PRFL
D.C.
Directional Coupler
P.M.
Power Meter
Cavity
Coupler
PTXM
D.C.
Phase Shifter
RF Box
Amplifier
PIN Switch/Att
Frequency meter
Attenuator
FM Modulation
Signal
Generator
F.M.
In high Q system it is necessary to have a feedback system to power the cavity at his resonance frequency.

## Slide 16: CALCULATION OF Q0 AND EACC
Calibration point measurements
Time decay  of the Ptrans
Q0 calculation
Pinc, Pref, Ptrans measured with power meters
with
and
Eacc calculation
and
Eacc

## Slide 17: RF tests on Saclay (β=0.65) and SPL (β=1) 704MHz mono-cell cavities

## Slide 18: 704 MHz SACLAY CAVITY 0.65 AT 2K
Strongly overcoupled big error (uncertainty) ≈ 30%
Min Eacc  18MV/m from different method
Validation of the electro polishing , rinsing, mounting, RF test stand…
Test parameters:
Pcav = 1.3E^-10 mbar
PHe= 28 mbar
T= 2K
fres= 700,65549 MHz
Chart: Saclay cavity RF test at 2K
| 0.33434974383070426 | 0.42237771344415864 | 0.52929890409448199 | 0.66865330380850119 | 0.84761960818485926 | 1.0597440693227262 | 1.3402968434894393 | 1.6776485846424471 | 2.1169031774689073 | 2.6497261683163797 | 3.3473489986979539 | 4.1995329660507021 | 5.2444626257354603 | 6.5117991213265913 | 8.3023491765311501 | 10.416000565243891 | 13.08280881180908 | 14.578102343024577 | 16.188290707164924 | 18.205432249300245 | 20.216268398080267 | 22.423375765735809 | 24.643420906333315 | 24.277334086239222 | 25.246489288849371 | 26.193950036826315 | 26.375517824402294 | 26.043598700734975 | 26.558344179755917
After EP | 23717725700.6292 | 24002870378.2127 | 23866973328.747 | 23454425552.2809 | 24194596423.6193 | 26163692890.9916 | 26164613106.591 | 26353595170.0803 | 24669632119.7168 | 23386090966.2311 | 22876736505.6863 | 22659119266.8789 | 22497542000.5076 | 23604653020.2572 | 21015012955.9822 | 19664396692.9145 | 16871227314.8028 | 15771906512.2474 | 15198727908.1926 | 13819475414.7944 | 13062405369.7849 | 12309068553.6359 | 11609675680.1084 | 11582481259.5931 | 11066570326.6793 | 10784847789.8779 | 10907120433.123 | 11097925757.0309 | 10827997296.7942
After EP from emitted power | 25586557530.9679 | 25894170106.8839 | 25747565085.8827 | 25302510718.1048 | 26101003154.5758 | 28225254049.5556 | 28226246773.2962 | 28430119627.7521 | 26613469160.8396 | 25228791726.6893 | 24679302818.1454 | 24444538487.4183 | 24270229783.7134 | 25464575319.0967 | 22670885261.0607 | 21213847552.073 | 18200591142.5474 | 17014649652.4311 | 16396307594.1206 | 14908377270.7201 | 14091654094.7245 | 13278958306.3556 | 12524456959.0874 | 12495119761.5071 | 11938557765.1136 | 11634636976.6441 | 11766544059.9937 | 11972383838.1835 | 11681186437.3759
Before EP | 1248749959.2527 | 1299325918.9444 | 1254342772.0939 | 1169146387.055 | 1220434095.2168 | 758639366.6605 | 388519399.5825
Saclay measurement | 69220000000 | 69210000000 | 67200000000 | 70450000000 | 71780000000 | 71260000000 | 71590000000 | 71360000000 | 68920000000 | 65230000000 | 64220000000 | 61570000000 | 58280000000 | 54390000000 | 48450000000 | 40970000000 | 33170000000 | 26430000000 | 21040000000 | 16700000000 | 13790000000 | 9490000000 | 6760000000 | 4990000000 | 3860000000 | 3050000000 | 2550000000 | 2330000000 | 2220000000 | 2250000000

## Slide 19: SPL β=1 MONO-CELL CAVITY TEST AT 2K
Test parameters:
Pcav = 10E^-10 mbar
PHe= 24 mbar
T= 2K
fres= 705,3152 MHz
- Calibration points were taken at 4 and 10 MV/m and agree better than 2 %.
- The cavity was strongly over coupled (Qext=6*10^8). This results in the rather big error bars, especially for Q0
- For Eacc a lower estimate can be made from an emitted power measurement. This technique yields a value agreeing with the lower limit of the horizontal error bars given in the graph
- Residual resistance is about 20 n, which is a decent value for this frequency
- Only mild medium Q-slope, no sign of Q-disease.
- Field emission starting at 13 MV/m, clearly identified from x-ray measurements (higher fields up to 2 mSv/h)
- He processing was done unfortunately the RF Input connector has been burnt (too much power)
Chart: Eacc (MV/m)
| 3.2546733649585482 | 4.356210247826696 | 4.8915420801660012 | 5.6177724729257577 | 6.702010792219899 | 7.6811461844317188 | 8.4679865503636549 | 9.123523479984275 | 10.244681127536516 | 11.903225467101111 | 11.91205297651282 | 12.480808778870488 | 13.377409054878425 | 13.996917794364451 | 14.467208076063265 | 14.844851444850818 | 15.34766302606878 | 15.578803287351349 | 16.16066406211359 | 15.84378582937145 | 16.597604810037272 | 16.72156736500645 | 16.896838507933904 | 16.984851502602901 | 17.838991668020096 | 3.9434906340999949 | 4.6034947867138909 | 5.20024595532107 | 5.9371825863364487 | 6.7086282790853025 | 6.309772677764828 | 5.5471398168585857 | 6.695659598393755 | 7.1708920272857695 | 8.097839225791482 | 9.1867974319468448 | 9.7623688683526719 | 11.052828081802677 | 1.2733367807051057 | 1.810520083545724 | 2.281856666241314 | 2.8827259413768953 | 3.6701355873100803 | 1.1452886195390994 | 1.0143621870739354 | 0.80363580340538643 | 0.50781684195825716
Q0(Eacc) at 2K | 10349270000 | 10110840000 | 10058050000 | 9957655000 | 9565304000 | 9465089000 | 9435052000 | 9341452000 | 9246393000 | 8431722000 | 8472983000 | 7877996000 | 6787013000 | 5708389000 | 4870455000 | 4200525000 | 3446491000 | 3102236000 | 2329106000 | 2681572000 | 1840587000 | 1700961000 | 1516156000 | 1011926000 | 890661900 | 9758274000 | 9757949000 | 9685537000 | 9510898000 | 9377871000 | 9413743000 | 9631090000 | 9582198000 | 9388959000 | 9271684000 | 9124778000 | 9050989000 | 8820321000 | 10456060000 | 10420240000 | 10122130000 | 10008120000 | 10002980000 | 10667030000 | 10758520000 | 10524040000 | 10596200000
Calibration point | 10483790000
Specifications:
Q0=10^10 at 25 MV/m

## Slide 20: INTEGRATION OF THE 5 CELLS ON THE NEW VERTICAL RF TEST STAND
New support for the cavity and for diagnostic system (OST)
Temperature mapping : octopus using Allen-Bradley resistors temperature sensor (K.Liao)
New cryostat Height = 4m and New cryo line at SM18 (P. Measen)

## Slide 21: OUTLOOK
Refurbishment and development of the cavity preparation, diagnostic system, RF test facilities:
Electro polishing, Rinsing, RF test stand, New automatic measurement program, 2nd sound Measurement, Temperature mapping, optical diagnostic
Result of the First test of the SPL mono cell cavity :
  Q0= 10E^10 until 13 MV/m and field emission appears at Eacc> 13MV/m with X-Ray (maybe due to the fact that we opened the cavity due the leak)
Next steps:
Integration of the mono-cell and 5 cells cavities in the new cryostat
New RF test on the SPL mono-cell cavity β=1
5 cells RF test
Thanks for your attention

## Slide 23: CAVITY COOLDOWN