## Slide 1
Secondary Electron Yield (SEY):
Pulsed Mode Development
M. Mensi, H. Neupert, M. Taborelli

## Slide 2: Outline
Secondary Electron Emission (SEE)
Physics of the SEE: Overview
Pulsed Mode SEY
Motivations
General Principle
System’s Development: Flowchart & Interface
Results
System’s Characterization: StSt as Benchmark
Insulating Samples: Continuous vs Pulsed SEY
Outlook and Conclusion

## Slide 3: SEE measurement: Physics of SEE
Secondary Electron Emission
SEY
Primary e- Energy [eV]
SEY =
Primary e-
Secondary e-
Secondary Electron Yield
dmax
SEE Energy Spectrum
Figure after Goldstein et al. 1981
Backscattered e-
X-rays
Cathodo-luminescence
Secondary e-
Primary Electron Beam
SE escape depth
Excitation
Volume
Sample
Secondary e- escape depth: typ. 1nm

## Slide 4: Pulsed Mode SEY: Motivations
Two major effects to be considered
Electron conditioning
Chart: Approx. Dose [C/mm2]
| 3.3299999999999998E-7 | 9.9999999999999995E-7 | 3.3299999999999999E-6 | 1.0000000000000001E-5 | 3.3300000000000003E-5 | 1E-4 | 3.3300000000000002E-4 | 1E-3 | 3.3300000000000001E-3 | 0.01 | 0.02
100eV | 1.84 | 1.78 | 1.67 | 1.58 | 1.5 | 1.41 | 1.35 | 1.34 | 1.37 | 1.33 | 1.33
268eV | 1.88 | 1.81 | 1.73 | 1.62 | 1.52 | 1.44 | 1.39 | 1.38 | 1.37 | 1.36 | 1.35
518eV | 1.76 | 1.74 | 1.69 | 1.62 | 1.52 | 1.45 | 1.44 | 1.4 | 1.39 | 1.39 | 1.395
1000eV | 1.89 | 1.81 | 1.79 | 1.65 | 1.54 | 1.47 | 1.44 | 1.41 | 1.385 | 1.42 | 1.42
Electron Stimulated Desorption: sample cleaning
From time to time: C increase on the sample’s surface
Electron conditioning of StSt
Charge accumulation
Insulator sample = capacitor: V=Q/C: charging effect
DV =
e0eRS / d
Q
++++
Electron Beam
SE escape depth
Excitation
Volume
Insulator
Sample
- - - -
SEY
PE Energy [eV]
1
d < 1
SEY
PE Energy [eV]
1
d > 1
Backscattered e-
- - - -
Secondary e-
Electron Beam
SE escape depth
Excitation
Volume
Insulator
Sample
++++
LOW DOSE
LOW DOSE

## Slide 5: Pulsed Mode SEY: Working Principle
SEY
PE Energy [eV]
Pulse Gen.
e-gun controller
SEY =
PE
SE
Ccoll+Csample
Ccoll
=
DC+AC
I/U
Oscilloscope
I/U
I/U preamp
DC+AC
Bias Networks
AC
+
-
AC
+
-
Sec. e-
Sample
e-gun
Prim. e-
SEY =
PE
SE
Icoll+Isample
Icoll
=
- +
+ -
Isample
Icoll
Continuous Measurement
Pulsed Measurement

## Slide 6: Pulsed Mode SEY: Bias Network
Techniques commonly used for continuous SEY measurements cannot be transposed to Pulsed mode SEY
What is a Bias Network
Bias Networks decouple the AC component of the signal from the AC + DC
No batteries in the signal path
Bias Networks should not affect the SEY since their theoretical behavior is linear with the charge
Discrete electronics components contain parasitic elements, i.e.:
Rreal = Rth + Lth + Cth …
AC
AC + DC
+
-
C
R
V
0 V

## Slide 7: Pulsed Mode SEY: Flowchart
Next E Step
Reset instruments
Data Processing
Initialize instruments
Save data
Set Instruments
Pulse and Acquisition

## Slide 8: Pulsed Mode SEY: Flowchart
Next E Step
Reset instruments
Data Processing
Initialize instruments
Save data
Set Instruments
Pulse and Acquisition
Set the e-gun parameters:
Energy 0 – 2000 eV
Focus = f(Energy)
First Anode = f(Energy)
Typ. Spot area 1mm2
SEY
PE Energy [eV]
d

## Slide 9: Pulsed Mode SEY: Flowchart
Next E Step
Reset instruments
Data Processing
Initialize instruments
Save data
Set Instruments
Pulse and Acquisition
Arm oscilloscope trigger
Pulse (single): typ. 1 – 10ms
Get waveforms from scope

## Slide 10: Pulsed Mode SEY: Flowchart
Next E Step
Reset instruments
Data Processing
Initialize instruments
Save data
Set Instruments
Pulse and Acquisition
Average the DC base level
Subtract offset to data points
Integrate Pulse
0 V
Charge
Sample and collector charge
Primary and Secondary C
SEY =
PE
SE
Ccoll+Csample
Ccoll
=
1 – 10ms

## Slide 11: Pulsed Mode SEY: LabVIEW Interface
Graphical User Interface: Main window and additional data

## Slide 12: Results: System’s Characterization
Chart: Primary electron energy [eV]
| 0 | 40.448980000000013 | 80.897959999999998 | 121.34690000000001 | 161.79589999999999 | 202.2449 | 242.69390000000001 | 283.1429 | 323.59179999999992 | 364.04079999999999 | 404.4898 | 444.93880000000001 | 485.38780000000008 | 525.83669999999972 | 566.28570000000002 | 606.73469999999998 | 647.18370000000004 | 687.63259999999991 | 728.08159999999998 | 768.53059999999994 | 808.9796 | 849.42849999999999 | 889.87749999999994 | 930.3264999999999 | 970.77549999999997 | 1011.224 | 1051.673 | 1092.1220000000001 | 1132.5709999999999 | 1173.02 | 1213.4690000000001 | 1253.9179999999999 | 1294.367 | 1334.816 | 1375.2650000000001 | 1415.7139999999999 | 1456.163 | 1496.6120000000001 | 1537.0609999999999 | 1577.51 | 1617.9590000000001 | 1658.4079999999999 | 1698.857 | 1739.306 | 1779.7550000000001 | 1820.204 | 1860.653 | 1901.1020000000001 | 1941.5509999999999 | 1982
 | 0.85225 | 1.0013 | 1.3502 | 1.5495 | 1.5705 | 1.5954 | 1.647 | 1.6633 | 1.6724 | 1.6652 | 1.6172 | 1.6108 | 1.5941 | 1.5605 | 1.5254 | 1.5354 | 1.4806 | 1.4736 | 1.4452 | 1.4321 | 1.4056 | 1.3791 | 1.3741 | 1.3547 | 1.3227 | 1.3018 | 1.2913 | 1.2453 | 1.2367 | 1.2092 | 1.1885 | 1.1676 | 1.1769 | 1.1692 | 1.1788 | 1.1537 | 1.1505 | 1.1548 | 1.1481 | 1.1463 | 1.1182 | 1.1181 | 1.1168 | 1.1001 | 1.1048 | 1.0932 | 1.0947 | 1.0773 | 1.0743 | 1.0615
 | 0.85317 | 0.96434 | 1.3311 | 1.5225 | 1.5229 | 1.5674 | 1.6012 | 1.6432 | 1.6457 | 1.6403 | 1.6048 | 1.6115 | 1.5803 | 1.5655 | 1.5255 | 1.4993 | 1.4775 | 1.4575 | 1.4351 | 1.4005 | 1.4014 | 1.3825 | 1.3581 | 1.3258 | 1.3129 | 1.2883 | 1.2661 | 1.2398 | 1.2145 | 1.2031 | 1.1875 | 1.1649 | 1.16 | 1.1543 | 1.1492 | 1.1556 | 1.1433 | 1.1498 | 1.1291 | 1.1223 | 1.1171 | 1.1153 | 1.1038 | 1.09 | 1.0874 | 1.0811 | 1.0888 | 1.068 | 1.0697 | 1.048
 | 0.85337 | 0.96441 | 1.3376 | 1.5166 | 1.5178 | 1.5608 | 1.6053 | 1.6091 | 1.6361 | 1.6442 | 1.636 | 1.5981 | 1.5911 | 1.5567 | 1.5346 | 1.5008 | 1.4816 | 1.4627 | 1.4415 | 1.4213 | 1.3928 | 1.3773 | 1.3515 | 1.3366 | 1.3331 | 1.295 | 1.2598 | 1.2338 | 1.2042 | 1.1935 | 1.1808 | 1.1774 | 1.1553 | 1.1553 | 1.1463 | 1.1544 | 1.1458 | 1.1416 | 1.1245 | 1.1265 | 1.1172 | 1.1088 | 1.1091 | 1.0886 | 1.0953 | 1.0802 | 1.078 | 1.0747 | 1.072 | 1.0558
 | 0.85008 | 0.97116 | 1.3435 | 1.5167 | 1.5234 | 1.584 | 1.5965 | 1.6286 | 1.653 | 1.6588 | 1.618 | 1.5921 | 1.5681 | 1.5629 | 1.5314 | 1.4963 | 1.4636 | 1.4784 | 1.4457 | 1.4089 | 1.3928 | 1.3589 | 1.3563 | 1.3431 | 1.3211 | 1.3025 | 1.2624 | 1.2369 | 1.2188 | 1.1892 | 1.1841 | 1.1732 | 1.1509 | 1.1407 | 1.1403 | 1.1618 | 1.149 | 1.1417 | 1.1291 | 1.1175 | 1.122 | 1.1114 | 1.1002 | 1.0879 | 1.0862 | 1.0804 | 1.0767 | 1.0714 | 1.0726 | 1.0642
 | 0.9117 | 0.94453 | 1.3416 | 1.519 | 1.4992 | 1.5852 | 1.5964 | 1.6205 | 1.6295 | 1.6489 | 1.6396 | 1.6231 | 1.5798 | 1.5685 | 1.5222 | 1.4954 | 1.4765 | 1.4633 | 1.4351 | 1.4117 | 1.4062 | 1.3926 | 1.3475 | 1.3283 | 1.3063 | 1.2793 | 1.2586 | 1.2393 | 1.2119 | 1.202 | 1.174 | 1.1711 | 1.1698 | 1.1648 | 1.1541 | 1.1452 | 1.1412 | 1.141 | 1.1303 | 1.1263 | 1.1334 | 1.1141 | 1.1079 | 1.0957 | 1.0827 | 1.0942 | 1.0818 | 1.0774 | 1.0723 | 1.0637
 | 0.85411 | 0.9605 | 1.3223 | 1.5032 | 1.5302 | 1.5645 | 1.5945 | 1.6274 | 1.6378 | 1.6228 | 1.6212 | 1.6084 | 1.5945 | 1.5434 | 1.5267 | 1.4853 | 1.4811 | 1.4387 | 1.435 | 1.4066 | 1.3891 | 1.3826 | 1.3611 | 1.3462 | 1.3205 | 1.2883 | 1.2511 | 1.2333 | 1.2152 | 1.188 | 1.1708 | 1.1763 | 1.1454 | 1.1581 | 1.1493 | 1.144 | 1.1402 | 1.138 | 1.1376 | 1.1264 | 1.1308 | 1.1178 | 1.1071 | 1.1113 | 1.0948 | 1.0758 | 1.0879 | 1.0753 | 1.0611 | 1.0702
 | 0.86353 | 0.97903 | 1.3405 | 1.5111 | 1.529 | 1.566 | 1.6248 | 1.6173 | 1.6278 | 1.6702 | 1.6383 | 1.5976 | 1.58 | 1.5615 | 1.5453 | 1.5116 | 1.4811 | 1.463 | 1.4353 | 1.4231 | 1.3966 | 1.3718 | 1.3575 | 1.3543 | 1.3174 | 1.2922 | 1.2617 | 1.2369 | 1.2091 | 1.1874 | 1.1736 | 1.1744 | 1.1622 | 1.1524 | 1.1377 | 1.1645 | 1.1455 | 1.1445 | 1.1337 | 1.1367 | 1.1168 | 1.1123 | 1.1067 | 1.1035 | 1.0874 | 1.0757 | 1.0936 | 1.0764 | 1.0853 | 1.0683
 | 0.86131 | 0.97529 | 1.3335 | 1.5119 | 1.5132 | 1.5654 | 1.628 | 1.6374 | 1.6361 | 1.6646 | 1.6377 | 1.6005 | 1.5728 | 1.5477 | 1.5174 | 1.4988 | 1.466 | 1.4659 | 1.4267 | 1.4066 | 1.3995 | 1.3624 | 1.3427 | 1.3343 | 1.3012 | 1.2833 | 1.2551 | 1.2265 | 1.2073 | 1.1882 | 1.1619 | 1.1571 | 1.1667 | 1.1592 | 1.1557 | 1.1507 | 1.1436 | 1.1338 | 1.1398 | 1.116 | 1.1188 | 1.1069 | 1.1002 | 1.0993 | 1.086 | 1.0841 | 1.0931 | 1.0633 | 1.074 | 1.0693
 | 0.86026 | 0.96181 | 1.3236 | 1.5002 | 1.528 | 1.5651 | 1.6227 | 1.6343 | 1.6495 | 1.645 | 1.6271 | 1.6193 | 1.5903 | 1.5382 | 1.519 | 1.507 | 1.4754 | 1.4412 | 1.4491 | 1.416 | 1.3925 | 1.3787 | 1.3495 | 1.3373 | 1.318 | 1.2807 | 1.2553 | 1.2331 | 1.2108 | 1.1896 | 1.1651 | 1.1732 | 1.1626 | 1.155 | 1.1678 | 1.1439 | 1.1394 | 1.1392 | 1.1335 | 1.1105 | 1.1138 | 1.1166 | 1.1175 | 1.097 | 1.0965 | 1.0924 | 1.0741 | 1.0744 | 1.0782 | 1.0635
 | 0.8609 | 0.93432 | 1.3373 | 1.4763 | 1.5308 | 1.567 | 1.6035 | 1.6404 | 1.6301 | 1.6239 | 1.6088 | 1.6165 | 1.5856 | 1.5473 | 1.5242 | 1.5002 | 1.4767 | 1.4444 | 1.4282 | 1.3996 | 1.3862 | 1.3821 | 1.3469 | 1.3262 | 1.3167 | 1.2907 | 1.2588 | 1.235 | 1.2119 | 1.1931 | 1.1769 | 1.1646 | 1.1628 | 1.1615 | 1.1543 | 1.1474 | 1.1437 | 1.1405 | 1.1229 | 1.1253 | 1.1053 | 1.1056 | 1.1009 | 1.0965 | 1.0966 | 1.0773 | 1.0819 | 1.0818 | 1.0682 | 1.0587
 | 0.83574 | 0.95632 | 1.3283 | 1.5376 | 1.5207 | 1.5637 | 1.6112 | 1.6202 | 1.6353 | 1.6533 | 1.6291 | 1.6065 | 1.5839 | 1.5542 | 1.5315 | 1.4946 | 1.4722 | 1.461 | 1.431 | 1.4075 | 1.3928 | 1.3546 | 1.3485 | 1.3337 | 1.3109 | 1.2877 | 1.2552 | 1.231 | 1.1917 | 1.1887 | 1.1635 | 1.1529 | 1.1582 | 1.1507 | 1.1398 | 1.1427 | 1.143 | 1.1394 | 1.14 | 1.1297 | 1.1115 | 1.1092 | 1.1078 | 1.099 | 1.0965 | 1.0871 | 1.0838 | 1.0816 | 1.0722 | 1.0719
 | 0.86042 | 0.96117 | 1.3339 | 1.5115 | 1.5215 | 1.5689 | 1.6084 | 1.6279 | 1.6381 | 1.6472 | 1.6261 | 1.6074 | 1.5826 | 1.5546 | 1.5278 | 1.4989 | 1.4752 | 1.4576 | 1.4363 | 1.4102 | 1.395 | 1.3744 | 1.3519 | 1.3366 | 1.3158 | 1.2888 | 1.2584 | 1.2346 | 1.2095 | 1.1923 | 1.1738 | 1.1685 | 1.1594 | 1.1552 | 1.1495 | 1.151 | 1.1435 | 1.141 | 1.132 | 1.1237 | 1.1187 | 1.1118 | 1.1061 | 1.0969 | 1.0909 | 1.0828 | 1.084 | 1.0744 | 1.0726 | 1.0634
StSt as benchmark sample
Measurement characteristics
Pulse width: 1 - 10ms
PE current: 10 - 100nA
Dose/pulse: -10-12 – -10-14 C/pulse
Cum.dose : -3 x 10-11 – -3 x 10-13 C
Continuous vs Pulsed Cum.Dose: 104-106
Signal Noise : 10nA p-p @500MHz BW
SEY Noise @ dmax : 0.15 p-p
SEY StDev @ dmax: 0.021
Continuous vs Pulse offset @ dmax : 0.06
Chart: Primary Energy [eV]
| 50 | 144.73679999999999 | 239.47370000000001 | 334.21050000000002 | 428.94740000000002 | 523.68420000000003 | 618.42099999999994 | 713.15789999999993 | 807.89469999999983 | 902.63149999999996 | 997.36839999999984 | 1092.105 | 1186.8420000000001 | 1281.579 | 1376.316 | 1471.0530000000001 | 1565.789 | 1660.5260000000001 | 1755.2629999999999 | 1850
Pulsed Mode | 1.3686 | 1.9656 | 2.0943 | 2.1249 | 2.0404 | 1.9339 | 1.8092 | 1.7356 | 1.657 | 1.5798 | 1.5239 | 1.4259 | 1.3792 | 1.3243 | 1.3166 | 1.2752 | 1.251 | 1.2215 | 1.2076 | 1.1902
Continuous Mode | 1.1006 | 1.7561 | 1.9835 | 2.0348 | 1.9869 | 1.9011 | 1.8096 | 1.7236 | 1.6498 | 1.578 | 1.5242 | 1.4796 | 1.4343 | 1.3912 | 1.3569 | 1.3264 | 1.2911 | 1.2555 | 1.2252 | 1.2008
Chart: Primary Energy [eV]
Repetitions @300eV, continuous
Chart: Primary Energy [eV]
Repetitions @300eV, Pulsed

## Slide 13: Results: Insulators
PC#17: H plasma contamination a-C coating
Continuous vs Pulsed-mode SEY on Insulators
Chart: Primary Electron Energy [eV]
| 0 | 58.294110000000003 | 116.5882 | 174.88239999999999 | 233.1765 | 291.47059999999982 | 349.7647 | 408.05880000000002 | 466.35289999999998 | 524.64699999999982 | 582.94119999999975 | 641.23530000000005 | 699.52940000000001 | 757.82349999999997 | 816.11759999999981 | 874.4117 | 932.70590000000004 | 991 | 1049.2940000000001 | 1107.588 | 1165.8820000000001 | 1224.1769999999999 | 1282.471 | 1340.7650000000001 | 1399.059 | 1457.3530000000001 | 1515.6469999999999 | 1573.941 | 1632.2349999999999 | 1690.529 | 1748.8230000000001 | 1807.1179999999999 | 1865.412 | 1923.7059999999999 | 1982
pulsed 1 | 0.89949 | 1.2269 | 1.3302 | 1.4306 | 1.4752 | 1.4673 | 1.4199 | 1.3877 | 1.3187 | 1.2657 | 1.2171 | 1.167 | 1.1262 | 1.0916 | 1.0687 | 1.0276 | 0.98883 | 0.97022 | 0.96618 | 0.9597 | 0.94078 | 0.93442 | 0.92208 | 0.91052 | 0.90964 | 0.90526 | 0.91206 | 0.87379 | 0.87288 | 0.875 | 0.86539 | 0.85597 | 0.85251 | 0.84905 | 0.84464
pulsed 2 | 0.73152 | 1.2298 | 1.2862 | 1.4457 | 1.4989 | 1.4785 | 1.429 | 1.3804 | 1.3289 | 1.2657 | 1.218 | 1.1701 | 1.1172 | 1.0835 | 1.0642 | 1.0237 | 0.9988 | 0.95566 | 0.94713 | 0.95026 | 0.94194 | 0.93579 | 0.91854 | 0.91871 | 0.90256 | 0.90158 | 0.89097 | 0.88154 | 0.86577 | 0.86738 | 0.86267 | 0.85925 | 0.85397 | 0.85491 | 0.84992
pulsed 3 | 0.99555 | 1.2436 | 1.3199 | 1.4365 | 1.4726 | 1.4438 | 1.4439 | 1.3651 | 1.3094 | 1.2737 | 1.206 | 1.1553 | 1.1241 | 1.0812 | 1.0495 | 1.0139 | 0.99056 | 0.96864 | 0.95497 | 0.94331 | 0.94591 | 0.94997 | 0.92388 | 0.92357 | 0.9143 | 0.89605 | 0.88977 | 0.87561 | 0.87108 | 0.86981 | 0.86039 | 0.85899 | 0.86162 | 0.85347 | 0.8531
continuous 1 | 1.3853 | 1.1866 | 1.3824 | 1.3307 | 1.2849 | 1.253 | 1.224 | 1.198 | 1.1727 | 1.1569 | 1.1438 | 1.133 | 1.1215 | 1.1102 | 1.0973 | 1.0704 | 1.0429 | 1.0184 | 0.99606 | 0.97519 | 0.9545 | 0.93561 | 0.9169 | 0.90033 | 0.88492 | 0.86993 | 0.85624 | 0.84367 | 0.83215 | 0.8216 | 0.81144 | 0.80166 | 0.79402 | 0.78796 | 0.78231
continuous 2 | 1.018 | 1.1866 | 1.3821 | 1.3523 | 1.2947 | 1.2548 | 1.2202 | 1.1884 | 1.1672 | 1.1521 | 1.1407 | 1.1276 | 1.1175 | 1.1071 | 1.0967 | 1.0695 | 1.042 | 1.0181 | 0.9958 | 0.97411 | 0.95397 | 0.9348 | 0.91683 | 0.90013 | 0.88441 | 0.86985 | 0.85615 | 0.84347 | 0.83217 | 0.82128
Ar bleed pressure : 4.7e-2 [Torr], H signal current @ RGA: 1.6e-9 [A] (at/or close to chemical regime), Power: 600 [W]
500x
5000x
SEM Morphology
More insulating samples are actually not measurable, work in progress…

## Slide 14: Outlook and Conclusion
Outlook
Decrease sample’s charge density by e-beam defocusing
Understand e-gun behavior
Better characterization of the measurements
Find optimal parameters, i.e. pulse width, dose, U/I gain
Add the possibility to save all data, i.e. every pulses
Conclusion
Hardware and Software for a pulsed-mode SEY setup were implemented
Preliminary characterization and limitations of the system were assessed

## Slide 15
Thank you for your attention