## Slide 1: R&D and simulations on gain stability and IBF for the ALICE GEM-TPC upgrade
Taku Gunji
Center for Nuclear Study
The University of Tokyo
For the ALICE TPC Upgrade Collaboration
RD51 mini week at CERN, Dec. 3-5, 2012

## Slide 2: Outline
ALICE GEM-TPC upgrade
R&D Status of gain stability
R&D Status of Ion back Flow
Simulation study of Ion Back Flow
Summary and Outlook

## Slide 3: ALICE GEM-TPC Upgrade
LoI of the ALICE upgrade
  https://cdsweb.cern.ch/record/1475243/files/LHCC-I-022.pdf
  Endorsed by the LHCC
High rate capability
  Target: 2MHz in p-p and 50kHz in Pb-Pb collisions
Plan for the ALICE-TPC upgrade
  No gating grid and continuous readout
    Inherited the idea from PANDA GEM-TPC [arXiv:1207.0013]
  MWPC readout will be
  replaced with GEM.
  Ne(90)/CO2(10)
  Ed=0.4kV/cm

## Slide 4: ALICE GEM-TPC Upgrade
Major issues for the GEM-TPC upgrade
  dE/dx resolution for the particle identification
    ~5% for Kr by PANDA GEM-TPC.
    Beamtest of prototype GEM-TPC at CERN PS-T10
      Detailed presentation by P. Gasik
Electronics (PCA16+ALTRO: loan from
the LCTPC. Thanks!!) & RCU
TPC Gas Vessel & GEM-Stack
very preliminary
of dE/dx for p/e
(no calibration,
no tracking)

## Slide 5: ALICE GEM-TPC Upgrade
Major issues for the GEM-TPC upgrade
  Stability of GEM (gain, charge up, discharge, P/T)
    Measurements in the lab.
    Test with the prototype at ALICE cavern in 2013. (p-Pb)
  Ion back flow to avoid space-charge distortion
    Requirement < 0.25%
    Test bench in CERN, Munich, and Tokyo
    Simulations to search for the optimal solutions

## Slide 6: R&D of gain stability
Measurement setup
  Single wire chamber as reference
  Monitor humidity
Mass flow
meters
GEM
Single-wire
chamber
Hygro-
meter
PC
NIM
CAMAC
PC
Sealed shielding box flushed with N2, containing GEM
Single wire chamber used as reference
Single/double GEM
gain~900-2000
Current density ~ 2-7nA/cm2
V. Peskov
J. Reinink

## Slide 7: Gain stability
2 GEMs (cylindrical holes) in Ar/CO2(10). Sr90 source
  4-5% variation of GEM and wire chamber current
    4-5% variation was compatible with temperature variation (T=23~24.5).
  Gain stays stable to within 1% after a few hours
  Humidity: 56-73 ppm. Gain~900 & current density~1.8nA/cm2
GEM current/Wire chamber current

## Slide 8: Gain stability
Gain~2000 & current density ~ 7nA/cm2
  Stability is ~3%
Next is to measure stability with 3 GEMs under Ne/CO2

## Slide 9: Ion back Flow
High rate operation (50kHz), continuous readout (no gating grid), and online calibration/clustering
  Need to minimize field distortion by back drifting ions
  Target: IBF ~ 0.25% at gain 1000-2000
Direction of IBF R&D
  Use standard GEMs and optimize by asymmetric electric field
  Use exotic GEMs (Flower GEM, Cobra GEM, MHSP)
  Simulations to search for optimal solutions

## Slide 10: Measurement at CERN/TUM/CNS
Systematic measurement of IBF
  Using 3 layers of standard GEMs
  Rate and gain dependence under various field configurations
Setup at CERN (RD51-Lab.), TUM, and CNS
  Various rate of X-ray gun
  Simultaneous measurement of IBF/energy resolution (TUM/CNS)
  Readout currents from all electrodes (TUM)
TUM: Technical University Munchen
CNS: Center for Nuclear Study, Univ. of Tokyo

## Slide 11: Rate dependence of IBF
Changing X-ray tube current and absorber filter
  Covering charge density= 1000-40000nA/(10cm2)
Clear rate dependence on IBF (0.1%~1%)
  Absorption of ions at GEM3 gets larger for higher rate
Due to space charge?
Ar(70)/CO2(30)
Y. Yamaguchi
C. Garabatos

## Slide 12: X-ray position dependence
Rate dependence
for 4cm case
Changing X-ray position from top of GEM1
  Different local charge density due to diffusion
IBF gets better for smaller distance between X-ray and GEM1 top (larger local charge density).
Due to space charge?
Ar(70)/CO2(30)

## Slide 13: Drift space dependence
Changing drift space from 80mm to 3mm.
  Different interaction rate
Clear difference in IBF due to different interaction rate.
  IBF gets better for larger interaction rate.
  IBF=2~5% for lower rate (not so much dependeing on rate).
Ar(70)/CO2(30)

## Slide 14: VGEM dependence
Due to space charge?
Ar(70)/CO2(30)
VGEM dependence for different drift space
IBF depends on VGEM.
  Steeper dependence for 80mm case.
  (even if rate dependence is small for 3mm, VGEM dependence is visible..)

## Slide 15: Space charge and IBF
Nions=104
Nions=105
Simulation study by garfield simulation.
  Presented at the last RD51 meeting on Oct. 2012.
Put many Ions above GEM1 (Ed=0.4kV/cm)
IBF strongly depends on Nions (>104). Space charge may play an important role for IBF.
Nions=0, 102, 103,
104, 2x104
Ions at [0, 100um]
above GEM1
T. Gunji

## Slide 16: Space charge and IBF
More dynamical simulations by garfield (2 GEMs).
  Presented at the last RD51 meeting on Oct. 2012.
Make spatial profiles of ions created by avalanches for 10usec (100kHz) and 100usec (10kHz) separated seeds.
Ion profile per one seed (Ar/CO2=70/30, Gain~1000)
Et = 3kV/cm
Ed = 0.4kV/cm
Et = 3kV/cm
Ed = 0.4kV/cm
electron
electron
10usec spacing
for avalanches
100usec spacing
for avalanches

## Slide 17: Space charge and IBF
Seed/hole=3
Seed/hole=10
Seed/hole=25
IBF vs. rate (time separation between 2 coming seeds)
  Rate/hole=10-50kHz in the lab. and less than 1kHz for LHC Pb-Pb 50kHz collisions
IBF strongly depends on rate, gain, and # of seeds/hole in case of high rate operations.
  Qualitatively consistent with the measurements.

## Slide 18: IBF from TUM (lower rate)
Reading currents from all electrodes.
3mm as drift space. X-ray tube with 2mm collimator.
Pad current ~ 5uA(<< current at CERN measurements.)
No strong rate dependence (may be due
to low rate and less space charge). IBF = 2-4%
Ar(70)/CO2(30)
A, Honle
K. Eckstein
M. Ball
S. Dorheim
B. Ketzer

## Slide 19: IBF from TUM (lower rate)
Reading currents from all electrodes.
3mm as drift space. No collimator.
No strong rate dependence (might be due to lower rate, much less space-charge).
IBF = 7%
Ar(70)/CO2(30)

## Slide 20: Different field configurations
Lowering ET2 and high ET1
  0.8% of IBF in Ar/CO2 with ET2=0.16kV/cm and ET1 = 6kV/cm (Ed=0.25kV/cm)
IBF=3-5% for Ne/CO2
  ET1 cannot be so high.
    Adding N2 to achieve
    high ET1?
gain=600~2000
gain=2000~6000
Scale=1.0
Scale=1.07

## Slide 21: IBF for various GEM configurations
Search for optimal solutions for IBF by simulations
T. Gunji
2GEM standard (same GEMs)
3GEM standard
2GEM, low Et (50V/cm)
3GEM, low Et2 (50V/cm)&VGEM2
(for various VGEM1/VGEM3)
Large pitch GEM1 + standard
GEM2 (Flower GEM structure)
Large pitch GEM1 + standard
GEM2 & GEM3
2 layers of cobra GEMs
3 layers of cobra GEMs
More studies are on going.
(higher gain, combination
of different geometry, etc…)

## Slide 22: Summary and Outlook
R&D of gain stability and ion back flow are on-going.
Gain stability <3% (2 GEMs, higher gain and rate)
  Stability of 3 GEMs will be studied under Ne/CO2.
IBF depends on rate of X-ray, spread of seed electrons (diffusion), and gain of GEM under high rate conditions.
  Space charge plays an important role for IBF under high rate.
  This is (partially) confirmed by garfield simulations.
Under low rate, IBF is 2-5% with 3 standard GEMs.
Try to reduce IBF further:
  More study on asymmetric field configurations
  Use standard GEMs with different geometry (hole size, pitch)
  Use exotic GEMs (Thick COBRA GEM, Flower GEM)
Simulation studies are on-going.

## Slide 23: Backup slides

## Slide 24: R&D of gain stability