## Slide 1: Status of the Micromegas industrialization activity with CIREA
December 2012 – RD51 MINIWEEK @ cern
Alan PEYAUD

## Slide 2: Elvia / cea partnership
Common fundings since one year (ANR)
develop large area detectors protected against sparks (SPLAM)
Main Bulk Tech. Transfer developments:
COMPASS: large pixel MM detector with buried resistors
CLAS12: large bulk on thin PCB (resistive solution also studied for the forward)
          …..few slides at the end of this talk
HL-LHC NSW program @ CEA Saclay

## Slide 3: Industrialisation for the HL LHC - nsw
GOAL
end 2012/ early 2013 - Large resistive prototype detectors (0.5x1 m2 ?)
2013 - Technical R&D (investment in machines) to achieve areas of 1 x 2 m2
2014/2016 - Production
ROAD MAP (NSW milestones)
Nov 2012: delivery of 3 BXY 10x10 cm2 detectors (CERN/Saclay)
End 2012: towards large areas (0,5x1 m2)
   Geoazur project (P. Salin) – to be redefined
2013/2014: towards nominal areas (1 x 2 m2) …… TO BE DISCUSSED
INDUSTRIAL PARTNERS
ELVIA / CIREA (France), ELTOS (Italy), Triangle Labs (Nevada, US)

## Slide 4: first step: small bulk XY
Use of an existing layout
Resistive bulk (50 to 100 MW along the strips)
X-Y readout
Active area: 10 x 10 cm2
Geometry:
      strip pitch 250 µm
      strip width 150 µm
Y readout
X readout
Coverlay
FR-4
128 um
Bulk Mesh

## Slide 5: Elvia / cirea detector CROSS section
PCB multi-layer
Resistive bulk XY cross section: pillar
449 µm
201 µm
187µm
Nominal figures
Coverlay: 64 µm
Amp gap: 128 µm
Pillar above mesh: 64 µm
Total pillar height: 192 µm
Resist. thickness: 35 µm
Mesh : 45/18 µm (30 µm)
97µm
38 µm
355 µm
resistive strip
wires

## Slide 6: BXY1 : First batch (july 2012)
3 detectors are delivered beginning of July
Resistivity meas. ~50 MW per strip (decent homogeneity)
…But visual inspection show waves on the mesh
Diagnostic: the mesh has not been stretched correctly
Measurements in gas give 55Fe spectrum with low energy resolution and evidence of an inhomogeneous amplification gap (lack of correct stretching)

## Slide 7: Bxy2 : Second batch (september 2012)
Grid manufactured by SERITEC (Switzerland) – stretch @ 15 N.cm
 3 detectors delivered at Saclay at the beginning of september
X
X
X
X
X
X
X
X
X

## Slide 8: Bxy2 : 55Fe calibration
Setup
Gas: Ar+10% CO2
Preamp: Ortec 142B
Amp: Ortec 472A
MCA
HV mesh: 520 V
HV drift: 750 V

## Slide 9: F4 after washing
Washed in demineralized water
Nitrogen drying
Oven curing (50°C) overnight
HVm: 550
HVd: 650
Gain: 7500
dE/E ~ 18%
Transparency

## Slide 10: The other detectors…
Same washing program than for F4 but no improvement - the issues:
the polymerisation was not long enough and unsufficient mesh bonding through the pillars
Visually the mesh seems not to be attached to the pillars except in the corners
F1
F3

## Slide 11: BXY2 3rd Batch – November 2012
2 new detectors were delivered early November
What has been done:
          Proper stretch from SERITEC
          Doubled the insulation time
          Proper curing time
Measurements in Ar + 7% CO2 with the 55Fe gives some very good result

## Slide 12: Synthesis of CIREA measurements on Bxy 3rd batch – received 24/10/12
Measurement made from the « comb » side to the extremity of each strip (1 meas. every 10 strips) with a Megger MIT 420 - calibre M 500V
MW
MW

## Slide 13: meas. On BXY 3rd batch
Chart: BXY F2.1 & BXY F2.2 behavior in AIR meas. @ CIREA
| 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900
BXY F2.1 | 21 | 40 | 68 | 90 | 120 | 148 | 180 | 265 | 320
BXY F2.2 | 12 | 25 | 32 | 45 | 56 | 68 | 88 | 128 | 191
Chart: Behavior in AIR meas. @ CEA
| 500 | 600 | 700 | 800 | 900 | 1000
F2.1 | 0.9 | 2 | 3 | 4 | 12 | 100
F2.2 | 0.6 | 0.6 | 1 | 1 | 1 | 25

## Slide 14: Transparency F2.1 & F2.2
Chart: Transparency @ HVmesh -540V
| 550 | 580 | 600 | 650 | 750 | 1000 | 1500 | 2000
F2.1 | 0.93367 | 1 | 0.9949 | 0.9949 | 0.9949 | 0.96939 | 0.76531 | 0.62755
F2.2 | 0.95408 | 1.0102 | 1.0306 | 1.0255 | 1.0357 | 1.0306 | 1.0255 | 0.82143 | 0.67347

## Slide 15: GAIN F2.1 & F2.2
Chart: Gain F2.1
| 480 | 500 | 520 | 540 | 560 | 580
Top left | 1350 | 2430 | 4275 | 7695 | 15075 | 31500
Top center | 1260 | 2250 | 4140 | 7785 | 15525 | 40500
Top right | 1372.05 | 2430 | 4455 | 8325 | 16155 | 35730
Bottom left | 1575 | 2115 | 3735 | 7200 | 13770 | 33300
Bottom center | 1125 | 2160 | 3780 | 7065 | 14625 | 40140
Bottom right | 1260 | 2295 | 4320 | 8865 | 15480 | 31500
Center left | 1260 | 2340 | 4095 | 7515 | 14535 | 29160
Center right | 1
CENTER | 1350 | 1755 | 2385 | 3420 | 4500 | 6075 | 8370 | 11520 | 16470 | 23850 | 36000
Chart: Gain F2.2
| 480 | 500 | 520 | 540 | 560 | 580
Top left | 1530 | 2790 | 5040 | 9360 | 18405 | 36090
Top center | 1350 | 2475 | 4365 | 8280 | 16830 | 45000
Top right | 1440 | 2655 | 4590 | 8640 | 15975 | 35550
Bottom left | 1
Bottom center | 1620 | 3015 | 5490 | 10575 | 20880 | 43200
Bottom right | 1350 | 2475 | 4500 | 8550 | 17370 | 44190
Center left | 1485 | 2700 | 4860 | 9180 | 18315 | 38700
Center right | 1260 | 2340 | 4320 | 8100 | 16110 | 43380
CENTER | 1350 | 1800 | 2565 | 3420 | 4590 | 6255 | 8730 | 12915 | 17280 | 26280 | 39600

## Slide 16: Homogeneity of BXY 2nd batch F2.1 @ HVmesh 540 V & HVdrift 790 V
<DE/E> ≈ 20% but some spread in peak position due to the position of the source on the edge of the drift

## Slide 17: Homogeneity of BXY 2nd batch F2.2 @ HVmesh 540 V & HVdrift 790 V
<DE/E> ≈ 20% and peak position more uniform
peak offset
1
3
2
4
6
5
7
9
8

## Slide 18: SUMMARY for BXY & NSW studies
Success!
Detectors F4, F2.1 & F2.2 are fully functioning goal reached 3 detectors to RD51 collaboration by the end of 2012
Mesh stretching problem is solved
Resistivity is homogenous
Fine tuning to improve the process (insulation time and coverlay curing)
What’s next?
Large PCB (1 and 2 x 0.6 m2) have been ordered for strips precision meas.
REDISCUSS the roadmap for the industrialisation process
              Now a few slides from Damien Neyret & David Attie

## Slide 19
 480 mm
Layout design made by CERN/Saclay:
  scale 1 of the final design
  960 resistive strips of 500 µm pitch
  100 µm double sided epoxy
  glued on a 4 mm-thick Rohacell
  copper layers: 35 µm
  « classical » manufacture of the resistive strips
Status of the prototype production:
  ELTOS (PCB) + CERN (resistive strips + bulk) × 2 (+1)
  CIREA (PCB + resistive strips + bulk) × 3
    quotation of CIREA: 11 k€ in 5 weeks delay
    order sent 2 weeks ago in ELTOS and CIREA
Final design production (end of 2013):
  8 detectors: 6 + 2 spares
    layout adjustments
    copper layers: 5 µm
  manufacturing of resistive strips using serigraphy?
  produced by CIREA if prototypes are qualified
Bulk on active area: 410 mm
Prototype of the CLAS12-Forward Tracker

## Slide 20: R&D on production at CIREA company For COMPASS
Activities on bulk production
A few small prototypes (6x10cm²) produced during 2011
1 with good gain performances
Several larger PLV3 prototypes (12x50cm²) batches produced since end of 2011 on thick boards
Also PLV3 on thin boards glued on Rohacell sandwich
Several problems at the beginning: bubbles in coverlay, pillars unstuck off board, low pillar thickness, waves on mesh
Quality has been improved with the time, almost good now
First large size bulk Micromegas produced at CIREA
Damien Neyret - CEA Saclay IRFU/SPhN
RD51 3-5/12/2012
| PAGE 20

## Slide 21: PLV3 boards from CIREA
Gain measured on April batch
Ar+5%iC4H10 gas
Very inhomogeneous gain
Some areas with nominal gain
Explained by bad pillar thickness
Finally solved by higher UV insulation
Batch produced in June with nominal thickness
First large size bulk Micromegas produced at CIREA early 2012

## Slide 22: PLV3 boards from CIREA

## Slide 23: PLV3 boards from CIREA
Present status of R&D on PLV3 prototypes
Last batch sparks at >750V in air
Pillar thickness nominal
Remaining problem: waves on mesh
New batch under production with more stretched mesh to get rid of waves
Part of next batch with thin board on Rohacell sandwich

## Slide 24: Production of large size prototypes at CIREA
59cm
73cm
59cm
73cm
Large size COMPASS prototypes
New rectangular design with active area 40x40cm²
Simplified design for first prototypes
Thin board on Rohacell sandwich to lower material budget
2 proto produced end of October
Installed in beam at Compass Thursday 29/11
In parallel discussions on production of buried resistors MM…

## Slide 25: Compass pixel MM board built at CIREA
Rohacell sandwich
16 large pillars to support
additional GEM foil

## Slide 26
CIREA pixelMM installed at COMPASS
1st prototype almost ready to be installed at COMPASS
First profile taken with beam
Bad APV card