## Slide 1: Six modules Micromegas TPC beam test & Test bench
D. Attié, P. Colas, M. Dixit, P. Hayman, W. Wang

## Slide 2
7 Micromegas modules with integration of electronics
2008-2011
2012-

## Slide 3
6 Modules Test
Cosmic event
In July 2012, 6 modules were installed on the endplate of LP-TPC at DESY and tested at the same time. Data were taken with cosmic rays and 5 GeV e- beam.
Notes: 33

## Slide 4
Event display (the first 100 triggers of a good run)
Each module has 1726 pads
(24 rows x72 columns)

## Slide 5
Outside crack
In crack
Upper crack
Lower crack
Padrow number (3x24=72)
5 GeV
e- Beam
Crack scan
x

## Slide 6
Number of hits vs x (72 pad-rows)
Study the hit reconstruction efficiency at or near the crack: take data with beam, moving by steps of 2mm
Crack scan

## Slide 7
Gain study
Using the fitted MPV of the (Landau) distribution of the measured charge
Effective gain vs mesh voltage
(compared with the direct gain measurements using 55Fe source on a bulk detector without resistive foil)
Effective gain vs drift distance
(Vmesh=380V, Vdrift=140V/cm, peaking time 100 ns and 400 ns)

## Slide 8
Track fit

## Slide 9
(b): After alignment
df= 2.9 , 2.1 and 2.2 mrad
dx = 47 µm, -43 µm and -52 µm
Row Number: 3 x 24 = 72 Displacement in mm
(a): Before alignment
Residuals by padrow
One module has 1726 pads
(24 rows x72 columns)

## Slide 10
Test Bench
The 55Fe source is put in an aluminum collimation tube and fixed on a set of two mechanical arms (X-Y).

## Slide 11
Test Bench
Module S1D - ~30 disconnected pads
Module S8G - ~30 disconnected pads
Module S7H - ~5 disconnected pads + some at the boarder
Module S4E - ~100 disconnected pads

## Slide 12
Evt with uncollimated source (each cluster is an X-ray photon)

## Slide 13
Energy Resolution (preliminary) :
~15% rms
Test Bench

## Slide 14
Successful test of 6 modules simultaneously with full integration of the electronics flat behind the panels. The analysis work is on-going.
Test bench is on-going to study the modules one by one: energy resolution, calibration, cross-talk…
Next step: 7 modules high quality test at the end of January 2013
Summary

## Slide 15
Thank you!