## Slide 1: Optimization of timing performance of large-area FBK SiPMs in the scintillation light readout.
A. Gola, A. Ferri, C. Piemonte, A. Picciotto, T. Pro, N. Serra, A. Tarolli, N. Zorzi
http://srs.fbk.eu

## Slide 2: Overview of the SiPM technology at FBK

## Slide 3
3
Original technology
FBK technology evolution
2006
RGB-SiPM
(Red-Green-Blue SiPM)
excellent breakdown voltage uniformity
low breakdown voltage temperature
dependence (gain variation < 1%/C)
higher efficiency
lower noise
2010-11
NUV-SiPM
(Near-UV SiPM)
excellent breakdown voltage uniformity
low breakdown voltage temperature
dependence (gain variation < 1%/C)
high efficiency in the near-ultraviolet
very low dark noise
2012
2012
RGB-SiPM_HD
(Red-Green-Blue SiPM – high density)
small cell size with high fill factor:
- high dynamic range
- low excess noise factor

## Slide 4
4
RGB-SiPM HD
SiPM:
size: 4x4mm2
cell size: 30x30um2
# cells: ~17000
SiPM:
size: 2.2x2.2mm2
cell size: 15x15um2
# cells: 21316
Fill factor = 74%
Fill factor = 48%
Redesigned cell border, for obtaining small cells with high Fill Factor (FF).
The 15 um cell RGB-HD has the same FF of the 50 um cell RGB technology.

## Slide 5
5
RGB-SiPM HD
t = 9ns
2.2x2.2mm2 15um
Response to fast light
pulse from LED
very short decay!!
Photo-detection efficiency
measurement by Musienko @ CERN

## Slide 6
6
NUV-SiPM
The NUV SiPM is based on a p-on-n junction, for increased PDE at short wavelengths.
NUV-SiPM: 1x1mm2 50x50um2.Total and primary dark count rate at 0.5 phe.
Very low primary DCR
PDE vs. wavelength for a NUV-SiPM and RGB-SiPM with 50x50um2 cell, 42% fill factor.
Increased PDE at low λ

## Slide 7: Effects of the SiPM noise: Dark Count Rate

## Slide 8
high LED threshold
 worse photon stat.
Effect of dark noise in LED
baseline fluctuation
 time jitter
In large area SiPMs the dark rate can be quite high and consequently also the effect described above.
Leading Edge Discriminator (LED) is commonly used for time pick-off with PMTs.
high dark rate
long tail of the
dark events

## Slide 9
SiPM Signal filtering: DLED
Original signal
Delayed replica
Differential signal
We exploit the difference
between rise time and
decay time to obtain a
signal:
“free” from baseline
fluctuations
identical initial part
of the gamma signal
Important: electronic noise is ~ √2 higher in differential
signal so its effect must be negligible for DLED to be effective
Then, we use the LED on the differential signal s2(t).

## Slide 10
Experimental set-up
Na22 source
SiPM + LYSO
current
feedback
amplifier
Digital scope (1GHz, 10GSa/s, 8 bit ADC)
1MW
50W
Energy channels
Timing channels
Ch1
Ch2
Ch3
Ch4
for energy eval
same signals
in a very expanded
scale to reduce
quantization noise
dark events
current
feedback
amplifier
4ns delay

## Slide 11
Detector “cube”
Detector “PET”
Detectors tested
LYSO crystal 3x3x5mm3
height ~ side
to test ultimate SiPM
performance
LYSO crystal 3x3x15mm3
height ~ 5 x side
real PET configuration
(timing affected by light
propagation in crystal)
SiPMs used
3x3mm2
4x4mm2
67um cell-size
produced by FBK, Trento

## Slide 12
DLED vs. LED performance
Detector cube,
3x3mm2 SiPM
DLED
LED
Low over-voltage:
low gain, low dark rate
 electronic noise dominates
 DLED slightly worse than LED
T=20C
High over-voltage:
high dark noise/rate
LED starts deteriorating
DLED still improves
for high PDE and
good noise compensation
Medium over-voltage:
gain increases
dark noise ampl. > elect. noise
LED is flat (increase of PDE is compensated by increase of noise)
DLED improves following PDE
DLED DT=500ps
good up to 1ns

## Slide 13
CRT vs. Temperature
0C
LED
DLED
20C
-20C
0C
20C
-20C
LED strongly improves with
temperature because of noise
(DCR) reduction.
Detector cube
3x3mm2 SiPM
DLED improves less with temperature
and only at high over-voltages.
LED@-20C is almost equivalent to DLED@20C

## Slide 14
Hardware Implementation
The DLED is difficult to integrate in an ASIC  Pole-Zero Compensation.
Non Compensated
Compensated
The passive recharge of the cells of the SiPM corresponds to a single real pole in the pulse response of the detector in the frequency domain.
 Pole cancellation through a zero at the same frequency in the front end.

## Slide 15
The PZ in the Time Domain
The tails have equal amplitudes and cancel each other if:
Simple circuit implementation of the PZ method.

## Slide 16
CRT with PZ
DLED sw
PZ hw
20C
-20C
The results are comparable or slightly better with the PZ method than with DLED  less noise without numerical differentiation.
Detector cube
3x3mm2 SiPM

## Slide 17: Effects of the SiPM noise: Optical Cross-talk Amplification

## Slide 18
CRT limit at high over-voltage
With Pole-Zero noise compensation and smaller detectors we observe almost no dependence of the measured CRT on the DCR of the detector  timing resolution should be limited by PDE of the device only.
PZ method
1.8x1.8x2 mm3 LYSO
2x2 mm2 SiPM
CRT limited by the front-end electronic noise.
CRT limited by the SiPM PDE (light collection statistics).
CRT divergence at high over-voltages.
Very good value!

## Slide 19
CRT limit at high over-voltage
The optimum CRT value is not obtained at maximum PDE.
There is some other effect that prevents to operate the device at higher over-voltage and PDE.

## Slide 20
Optical Cross-talk Amplification
The scintillator reflects the photons emitted by the hot carriers during the avalanche.
Increase in the collection efficiency of the optical cross-talk photons.

## Slide 21
Reverse SiPM Current
The phenomenon can be easily observed from the reverse current of the device, measured with and without the scintillator.
The current increase is larger for the shorter scintillator.
3 cases:
No scintillator
2x2x2 mm3 LYSO
2x2x10 mm3 LYSO
3 temperatures: 20°C, 0°C, -20°C

## Slide 22
DCR amplification
It is possible to define a DCR amplification coefficient due to the scintillator enhanced cross-talk.
There is no amplification if the LYSO top face is covered with black paint.
Extremely high values!
The process rapidly diverges.

## Slide 23
Correlation with Timing
The γ coefficient shows an almost perfect correlation with the CRT.
The PZ filter can no longer compensate for the detector noise.
Reduction of the effective PDE, due to cells busy with DCR.
Basically, the SiPM stops working.

## Slide 24
Possible Solutions
For reaching higher over-voltage and PDE, it is necessary to reduce the amount of cross-talk:
Color filters in the scintillator
  Suppression of external cross-talk
Optical Trenches and Double Junction
  Suppression of internal cross-talk (total CT is determined by both internal and external components)
Smaller cells with smaller gain but same fill-factor
  Updated fabrication technology and cell-edge layout

## Slide 25
Conclusion
Once the effects of the DCR are removed, we can observe a different phenomenon, limiting the timing resolution of the SiPM, related to the optical cross-talk, increased by the scintillator.
We have demonstrated that the effect of dark noise on timing measurements with SiPM coupled to LYSO crystal can be largely compensated.
The baseline compensation can be implemented with a simple, ASIC compatible, analog circuit.
Possible solutions are under investigation, however the HD technology seems a very promising option.

## Slide 26
Acknowledgments
HyperImage project
SUBLIMA project
FBK-INFN MEMS2 agreement

## Slide 27
Back-up slides

## Slide 28
In the next slides:
CRT vs Temperature
Threshold level
crystal height
SiPM size

## Slide 29
Optimum Threshold level
0C
LED
DLED
20C
-20C
0C
20C
-20C
Detector cube
3x3mm2 SiPM
0C
20C
-20C
0C
20C
-20C

## Slide 30
Leading Edge Discriminator
signal
distortion
baseline
fluctuation
forced to use a
higher threshold
What is the effect of noise on timing?
jitter
jitter
jitter for worse
photon statistics
HF noise
LF noise
furthermore
LED is widely used in PMT-based systems

## Slide 31
First approach: SiPM signal shape
In this case we increase both
the quenching resistor and the
quenching capacitor in order
to enhance the fast component
and decrease the slow one.
slow tail
fast component
SCR of a 4x4mm2 SiPM
Steeper rising edge of the
gamma pulse and lower
baseline fluctuation
- SiPM 4x4mm2
- LYSO 3x3x5mm3
- LED

## Slide 32
original
gamma
signals
differential
gamma
signals
+ initial part of gamma ray
signal preserved
+ baseline compensated;
+ electronic noise increased;
Example
expanded y scale

## Slide 33
Detector PET
Detector cube
SiPM 3x3mm2
SiPM 4x4mm2
DLED – comparison
20C
-20C
20C
-20C
20C
-20C
20C
-20C