## Slide 1: Summary of experience with LHC and SPS IPM monitors in 2011 & 2012 (& 2013)
Mariusz Sapinski, CERN BE/BI

## Slide 2: Outlook
High-brightness beams
Short description of SPS and LHC IPM
  HV
  Gas injection
  Magnets
  Imaging, signal distortion (along the path)
Hardware failures
Calibration methods

## Slide 3: Beams
Beams with emittance as small as 0.7 μm∙rad (pPb run)
Typical emittance for high-intensity protons: 1.5 μm∙rad
(designed emittance 3.5 μm∙rad)
Up to 1.7∙1011 protons/bunch
Beam size: at 7 TeV, β=100 m: 100 μm
2800 bunches

## Slide 4: System: HV cage

## Slide 5: System: HV cage + PS + control
EGPin - 1800 V
EGPout - 1900 V
Cgrid - 2000 V
MCPin +2000V
MCPout +2600V
Phosphor +5000V
HVctrl
PS
VME
~ 180 m
beam
anti-SEM wires
Field-shaping electrodes
Accelerating potential

## Slide 6
VME
Negative electrode

## Slide 7: System: gas injection
VME
needed in LHC (normal vacuum 10-11 mbar)
gas injection up to 10-8 mbar
PVSS application
manual control (3-4 steps)
timeout after 12 h
typical LHC fill
beam intensity
beam energy
pressure
max after ~2h on flat top

## Slide 8: System: magnetic field
VME
Magnetic field needed to keep minimize beam space charge effect
0.2 T magnets originally from ISR, yoke modified to extract light
Need to exchange power converters on SPS magnets to allow cycling.
magnets are compensated (ie. 2 magnets/detector/plane in the same circuit)
length 43 cm
20 cm space between poles
Field quality:

## Slide 9: System: imaging
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
EGP with anti-SEM wires
(calibration of MCP ageing)

## Slide 10: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Anti-SEM wires glowing:

## Slide 11: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Solution: masking at level of image processing
Interestingly it is mainly observed in
vertical IPM

## Slide 12: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Beam space charge
And electron gyroradius
See Marcin’s presentation

## Slide 13: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Electron cloud (suspected)
Appears only in specific conditions (scrubbing).
electrodes and MCPs are NEG-coated to suppress e-cloud
scrubbing

## Slide 14: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Nonuniform MCP ageing
Phosphor screen burn-in?
During winter 2011/12 MCPs in 2 BGIs have been exchanged. Signals improved a lot, but both MCPs died soon because of HV/high signal problems.

## Slide 15: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Solution: use EGP to measure nonuniformity and correct for it.
But EGPs did not work in a very reliable way.
They stopped working at the end of run (ageing?)
Orbital bump possible, but small at injection.

## Slide 16: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
MCP hole pitch: 32 μm
+
Signal broadening between MCP and Phosphor screen due to emission cone.

## Slide 17: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Optical system PSF is
estimated to be 22 μm (ZMAX)
D. Kramer et al., CERN-AB-2005-072

## Slide 18: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Camera tilt, measured and corrected.

## Slide 19: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Transmission using analog video signal:
line-to-line noise and other artefacts

## Slide 20: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
Transmission using analog video signal:
line-to-line noise and other artefacts
Solution: image filtering, eg. 2D FFT filters (still being tested)

## Slide 21: Signal distortion
beam
E
B
MCP
Phosphor
Prism
Optical system View port
CID camera
(intensified)
Thermo Scientific
CID8712D1M-XD4
Video amplifier
~180 m
Frame grabber
SPS: picking up a huge noise on video signal correlated with beam type/intensity
Cable shielding and exchange of video signal amplifier were helpful.

## Slide 22: Hardware failures (other than already mentioned)
HV ctrl card – problem with compatibility with VME (linux CPU)
5 CID cameras stopped working, in most cases we suspect that intensifier reached MTTF (tbc by ThermoFischer).
failures of MCPs
  “conditioning effect” for MCPs
  too high input electron current might kill MCP
  abrupt HV change might kill MCP (and dump the beam!)
Killed MCP: creation a conducting channel through the plate: cannot set HV anymore, cannot amplify the signal.

## Slide 23: Fit and amplitude dependence
ROOT fitting library (the same as for BSRT)
there is a dependence of measured beam sigma on the signal amplitude
amplitude feedback is crucial (we do it through camera intensifier gain)
(changing HV on the MCP in the vacuum is risky)

## Slide 24: Calibration from specifications
Camera pixel size: 11.5 μm * 1.6 (taper)
Optical system magnification: 0.2
Calibration = 92 μm/pixel
Calibration in the lab
The value found:
115 ±3 μm/pixel
by Marcin

## Slide 25: Beam-based calibration
Interpolate beam position from neighbor BPMs.
Make an orbital bump
Error on orbit position in BGI location: 100 μm (interpolation error)
BPM1 75m BGI 30m BPM2
Calibration = 95-99 μm/pixel
Installation of BPM close to
BGI is foreseen during LS1

## Slide 26: Cross-calibration (WS, BSRT)
Because of old MCPs BGI sensitivity starts where WS cannot measure
But for ions there was an overlap
BSRT uses cross-calibration with WS, so calibration with BSRT is of “second order”.
Additional broadening in BGI – point spread function?
Pb82+
Pb82+

## Slide 27: Ramp-based calibration
σPSF=√p1 = 0.5 mm
It is possible to optimize the fitting procedure, but is the quadratic correction enough to correct for all errors, especially the beam space charge?

## Slide 28: Conclusions
LHC IPM deals with beams of unprecedented brightness.
System was in commissioning for the whole period. Almost not used for operation during Run 1 because of difficulties to calibrate it.
But a lot of studies were done, a lot of data collected, significant upgrades.
We suspect that there are strong physical reasons behind calibration issue.
We think that we need a stronger magnetic field.
SPS IPM was renovated but noise on video signal was too large for most beams to use it.
There are some hints on how to reduce it.
A series of upgrades during LS1 is planned but we are happy to hear your advices.