## Slide 1: Summary of experience with LHC synchrotron light monitors
9th DITANET Topical Workshop on Non-Invasive Beam Size Measurement for High Brightness Proton and Heavy Ion Accelerators
CERN – 16-Apr-2013
F.Roncarolo for the BI-BSRT team
W.Andreazza, A.Boccardi, E.Bravin, F.Caspers, A.Fisher, A.Goldblatt, J-J. Gras, A.Guerrero, M.Kuhn, L.Jensen, R.Jones, T.Lefevre, A.Nosych, G.Schneider, G.Trad, R.Veness, M.Wendt … and many others.

## Slide 2: Scope and Contents
SCOPE
Give an overview of the LHC sync. Light systems (BSRT), outlining reached performances and issues.
More details and numbers in G.Trad’s presentation
CONTENTS
LHC sync. Light detector (BSRT) introduction
  Light source(s)
  Imaging system
  Image acquisition
Performances and issues
  Heating
  Calibration
  Accuracy and resolution

## Slide 3
Introduction

## Slide 4: LHC BSRT
Synchrotron radiation is generated by the SC undulator up to proton energies of ~ 1.5 TeV and by the D3 dipole for higher energies
~ 30 m
Extraction mirror
View Port

## Slide 5: Extraction Mirror

## Slide 6: Optical Telescope Layout
Gated camera
(BSRTS)
DC Camera
Abort Gap Monitor
(AGM)
Long. Density Monitor
(LDM)
Optical delay line
Neutral filters
Color filters
Proton/Ion
beam
90 %
10 %
60 %
40 %
40 %
60 %
Light delivered (after extraction) among transverse and longitudinal diagnostics

## Slide 7: Image Acquisition
Proxitronics gated intensified (MCP) camera
  Intensifier max trigger rate : 200 Hz (~55 LHC turns)
  Intensifier min gating : 25ns (1 LHC bucket)
Present max acq. rate (including low level SW overheads)
  10 Hz  on paper 10 bunches per second, then slower to get statistics on bunches
Overall system sensitivity
  We have enough light to see
    single proton pilot bunch (5e9p) at injection (450GeV), single turn (not consecutive turns due to 200Hz trigger limit)
    ~20 Ion Pb bunches at injection, averaged over 4 turns
Photocathode response
(our cameras are equipped with N type for the moment)

## Slide 8: BSRTS - Proton Image Example
Single bunch ~1.1e11p @ 3.5 TeV
Accumulated over 4 turns separated of 55 turns (200 Hz trigger rate limit on image intensifier)

## Slide 9: Telescope Optics
Imaging system needs to
  Propagate light in the wavelength accepted by the camera photocathode
  Adjust focusing according to the dominant source (undulator, dipole edge, dipole core)
Version 1 : 2 focusing mirrors and adjustable delay-line
  Mirrors minimize aberrations over a large wavelength range (need infrared for ions)
Version 2 (middle 2012): 2 focusing lenses and adjustable Lens1 position
  Lenses optimized in the visible
  Allow a much simpler telescope layout

## Slide 10: BSRT Optics
Old Optics
New Optics
Entrance (steering) mirror
Mirror 1
Mirror 2
Lens 1
Lens 2
Delay Line

## Slide 11: Old Optics
Pictures showing only part of the optical path
LIGHT EXTRACTION WINDOW
FIRST MIRROR ON TELESCOPE BELOW THE BEAM PIPE
MIRRORS ON MOVABLE STAGE “Trombone”
 VARIABLE OPTICAL PATH LENGTH

## Slide 12: BSRTS -Effect of optical path length and color filter
[mm]

## Slide 13
BSRT Heating

## Slide 14: BSRT heating
From end 2011: observed hints of extraction system deformation
  beam spot @ camera displacing
  correlated to beam intensity
  not compatible with beam orbit changes
Beginning of 2012:
  spot displacement correlated to temperature outside the tank

## Slide 15: Temperature Probes (outside vacuum)
Flange 1
Flange 2
View Port 1
View Port 2
Bellow

## Slide 16: BSRT - Heating
15
B1 heating ~= B2 heating
Bellow heats more and faster than flange (== is radiation higher than conduction?)
Mirror OUT
Mirror IN
Mirror Falling
Before TS#3 (‘bellow’ probe B1 not at same location as B2)
After TS#3
‘bellow’ probe B1 displaced to be = B2
Added 2 probes on flange
BEAM 1
BEAM 2
1380 b
1380 b
Bellow
Bellow
Flange
Flange
View Port
View Port
1380 b
1380 b
1380 b
1380 b
LAMP ON

## Slide 17: BSRT – Heating – B2 Mirror 28-Aug-2012
1- Beam (spot not in the ‘right’ place vertically)
2- Beam dump
3- Close Vacuum sector and retract mirror to OUT
4-No beam, no motors movement, mirror moved
Access to remove mirror

## Slide 18: BSRT – Heating – Findings after B2 mirror removal
Mirror coating blistered
Mirror clamps deformed
Later Removal of B1 mirror evidenced similar effects
  Both mirrors were: silicon bulk + dielectric coating
TS#3: replaced both mirrors
  B1: glass bulk + metallic coating – OK only at low beam intensities
  B2: silicon bulk + polishing (no coating) – not usable for imaging

## Slide 19: Broken Ferrite

## Slide 20: Temperatures in vacuum
Xmass 2012: installed probes in vacuum
Mirror 1
Shaft 1
Shaft 2
Ferrite 1
Ferrite 2
Mirror 2
Ferrite heating likely > Curie Temp
Clear correlation to beam intensity and bunch length
Many RF simulations and lab meas. ongoing  aim at new Tank design for after LS1

## Slide 21: Mirrors
Silicon Bulk – Dielectric coated mirror: found damaged
After this, we tested different solutions (LAB + beam tests)
  A silicon bulk, uncoated mirror showed a reduced heating (as measured with temperature probes outside the BSRT tank), but resulted to be unusable for imaging, given the distorted recorded images.
  A glass bulk, metallic coated mirror resulted in a reduced heating effect at low beam intensities, but suffered coating deformation (evidenced by the beam spot image deformation) at high intensities.
  A glass bulk, dielectric coated mirror resulted in a reduced heating (w.r.t. the original silicon bulk, dielectric coated mirror) and did not show any coating deformation according to the recorded images, also at high intensity.

## Slide 22
BSRT Calibration

## Slide 23: BSRT - Calibration
Strategy is simple:
  Measure beam size with WS and BSRT simultaneously
  Machine development periods with large emittance difference among bunches
sm [pixels] = Measured Sigma
M [mm/pixels] = System magnification (optical line + pixel size)
  determined by optical line adjustment with energy (UND  D3)
  reference target vs closed orbit bumps give up to 10% differences
spsf [pixels] = measurement error due to
  aberration, diffraction
  extraction mirror aging/deformation
During 2012 heavily affected by extraction mirror heating (deformation + coating damages)
Difficult to find a stable calibration w.r.t. WS
Normally assuming
BSRT magnification
BEAM optics from WS to BSRT
as known
See G.Trad presentation
From WS

## Slide 24: LHC Beams – BSRT magnification
Need to measure beams as small as 300um (150-200um at 7 TeV)
Optical magnification ~0.3 (Simulation, Zemax)
Old Optics : ~0.11 mm/pix
New Optics: ~0.05 mm/pix (2 pixels per sigma criterion respected)
Magnification measurements done with calibration target and beam orbit bumps
Calibration target @ same distance as undulator
Conclusion (if we believe target): we can’t say we know magnification better than 10%
(again: error on beam optics == error on target)

## Slide 25: BSRT – WS calibration example (2011)
Bunches with different emittances from the SPS

## Slide 26: BSRT – WS calibration Example (2012)
Injected 12 bunches per beam with emittance as small as possible
Use ADT to blow-up single bunches
Different ADT strength, excitation time on different bunches
 circulating bunches with emittance spread in the 1.5 – 5 um range
Measure with WS and BSRT
Calculate BSRT magnification and point spread function to match WS
Calibration BSRT - WS
Emittances after calibration
MAY 2012 – 450GeV

## Slide 27: BSRT – Calibration Example (2012, new optics)
B1 with new optical line (focusing lenses instead of mirrors), 450 GeV during MD period
Excellent agreement BSRT – WS over a wide emittance range, after applying
  Magnification within 10% w.r.t. nominal
  PSF ~20% smaller than typical values with old optics
Scraping
B1 Vertical
Similar examples exist at 4 TeV
More numbers on corr. factors in G.Trad’s presentation
1.3 mm
1 mm
0.8 mm

## Slide 28: BSRT calibration - Outlook
Combining
  magnification estimate from target and bumps
  Cross-calibration w.r.t. WS
  We managed to find calibration factors providing excellent agreement w.r.t. WS, over a wide emittance range
BUT:
  Calibration is not stable
    Surely depending on extraction mirror ageing
    Dependence on light traversing the optics off-axis?
      Hope to improve stability with new optics (less folding mirrors, didn’t have enough statistics with new optics yet)
  Correction in quadrature (psf form aberration, diffraction etc) higher than expected from theory

## Slide 29: Issues Snapshot
Overall reliability
auto-steering, auto-gain
New FESA server
Improved reliability
Fast bunch per bunch scans (~10Hz)
Problems encountered
Improvements – Upgrades during 2012
Electro-magnetic coupling - heating
  Mirror support failures, coating blistering  needed to replace both mirrors after August
Absolute calibration
  Drifting due to mirror coating damages
  Noisy due to numerous folding mirrors (?)
   Image blurring
New optics
Simplified imaging system – lower PSF corrections
New mirrors
From Si to Fused Silica bulk – reduced EM coupling

## Slide 30
BSRT useful data examples for LHC operation

## Slide 31: Bunch per bunch from BSRT
Bunch per bunch emittance vs time
Bunch per bunch emittance @ given time
Allowed diagnosing systematic emittance pattern from injectors (PSB rings)

## Slide 32: Bunch per bunch from BSRT
During 25ns bunch spacing scrubbing run:
E-cloud makes bunches in the second part of each train to blow-up

## Slide 33: Outlook
Upgrades for LS1
  New light extraction system to cope with RF coupling
    Option 1: Ceramic holder + support
    Option 2: Longer mirror through a slit (metallic parts “hidden”)
  Optics compatible with higher energy / higher diffraction contribution
Further (not decided, not scheduled) upgrades
  We have a good optical calibration system: long (multi folding mirrors) line on the optical table) for alignment but not for imaging a target
    Should we envisage a target system in vacuum?
  New (digital) intensified cameras?
  Thank you for the attention and ….
  Ideas are very welcome !! ;-)

## Slide 34
SPARES

## Slide 35: BSRT – RF simulations
Longit. wake impedance of BSRT
with and without Ferrite damping
Measured LHC bunch power spectrum. A 650 MHz resonance is very dangerous.
LHC bunch power spectrum
F = 650 MHz
P_loss = 10-50W
Credits: T.Mastoridis,
P.Baudrenghien/CERN
B-filed of the beam in Time Domain.
Red = Hot (bigger current density)
Blue = Cold
Mirror Back
Mirror front
E-field of a dominant resonating mode at 650 MHz.
(Q = 1263 / Rsh = 25841 Ohm)
From January:
RF laboratory measurements on spare tank

## Slide 36: Long. Wake Impedance (magnitude)
BSRT w/o Ferrite not shown!
M.Wendt, A.Nosych

## Slide 37: Calibration with target

## Slide 39: MACOR version

## Slide 40
Slit is 1 mm wider than the mirror in each direction (so the mirror can be adjusted)
Ferrite TT2 ring to damp any field entering through the slit. However no visible effect in simulations.
Introducing: elongated Fused Silica mirror in an obstacle-free copper holder with a slit.

## Slide 43: AGM and BSRT - Simulated vs Measured light intensity
Intensity per charge (equiv. AGM counts at 3200V
Visible Photons per Charge
Intensity per charge (equiv. AGM counts at 3200V
Visible Photons per Charge
Beam Energy [GeV]
Beam Energy [GeV]
- - Simulation p+
LEAD IONS, Abort Gap Monitor (AGM) and BSRTS
- - Simulation Pb ions
⎯ Measured AGM p+
⎯ Measured AGM Pb Ions
Abort Gap Monitor (AGM)
⎯ BSRTS HOR.
⎯ BSRTS VER.
- - AGM
- - Simulation Pb Ions
At least a factor 1e4 difference between protons and ions at injection energy
Nevertheless: we managed to see light at injection with ions

## Slide 44: Fast Camera Tests
Single Nominal bunch 450 GeV
Mode Normal
Gain 4.99V
Exposure 3ms
Gate size 25ns
Scan of 3 consecutive 25ns slots
Prove that gating can be done

## Slide 45
Exposure 3ms
Exposure 1ms
Single Nominal bunch 450 GeV
Mode Normal
Gain 4.99V
Gate size 25ns
Exposure
Fast Camera Tests
Need to integrate > 10 turns to measure a single bunch
Camera and gating control via FESA not fully proved