## Slide 1: Overview of the control and acquisition systems
CERN wire-scanner development review 18.04.2013
J.Emery for the BWS team

## Slide 2: Plan
Initial system layout consideration
Overview
Acquisition part
Control and monitoring part
Details on the wire measurement
Connection to external systems

## Slide 3: System layout consideration
Minimisation of complexity - As less as possible electrical connections - Avoid multiplexing channel - Settings and parameters (control and acquisition)
Separate critical and non-critical functions - Instrument safety in the low level hardware - High level software configure, analyse and record the measurements
Early failure detection - Constant hardware Condition Monitoring (CM) - Programmed maintenance
Continuous connection between mechanics and electronics => no multiplexing scanner control
Position measurement and motor driver on the same board
Self contained actuator control disconnected from the crate => usable without CPU
Monitoring at the same location as the actuator to act in case of failure
Acquisition (non-safety critical) don’t share the same firmware with the actuator (safety critical)

## Slide 4: Future Scanner Control and Acquisition System
Scanner control, monitoring and supplies BWS_SCMS
High Precision Bunch by Bunch Profiling BWS_HPBBP
CPU
T IMING
BST receiver
Optical encoder interface
Fast integration secondary shower
SFP Ethernet
SFP Optical link
SFP Optical link
Long range resolver interface
3-phases PWM
Magnetic stopper
External temperature and vibration
SFP Optical link
SFP Optical link
Tunnel Electronics and GBT Link
Optical encoder interface
Wire resistivity & Strain gauge
Secondary particles shower acquisition
See Jose’s talk
HV
CTRL

## Slide 5: Acquisition system overview High Precision Bunch by Bunch Profiling (HPBBP)
Programmable device (FPGA)
Link to BWS_SCMS
Bunch synchro
CPU interface
laser
Pdiode
Circulator
laser
Pdiode
Circulator
Analog frontend
Secondary particules detector
Fork
Vacuum barrel
Shaft
Motor
Break
Resolver
Res stator
Break stator
Optical disk
Vacuum Focuser
Optical feedthrough
Electrical feedthrough
Wire
Bearings
Optical feedthrough
Motor stator
Vacuum Focuser
Bearings
Combines the diamond detection and optical encoder => See Jose’s talk
Digital platform (VME, SFP and FPGA)
Specific interfaces to detectors
Custom firmware

## Slide 6: Control and monitoring overview Scanner Control Monitoring and Supplies (SCMS)
Fork
Vacuum barrel
Shaft
Motor
Break
Resolver
Res stator
Break stator
Optical disk
Vacuum Focuser
Optical feedthrough
Electrical feedthrough
Wire
Bearings
Optical feedthrough
Motor stator
Vacuum Focuser
Bearings
Programmable device (FPGA)
Digitalisations (A/D, D/A)
DC bus 300V
DC low voltages
Main supply input
Ethernet monitoring
Link to HPBBP & CPU
Machine synchro
GBT interface
Stopper Driver
Resolver Driver
3 phases Motor PWM
Wire resistivity
Filters OTS
Temperature vibrations
Optical position sensor interface
Strain gauge

## Slide 7: Interface to the wire resistivity, thermionic and RF induced current
PSB Ring 1 H at 15 m/s
Wire resistivity variation during a movement due to mechanical constraints (See Juan’s talk)
SPS 416 H at 6 m/s
Wire current variation when interacts with beam (thermionic effect)
SPS 416 H at 6 m/s
Wire current variation induced by the beam exitating RF modes inside a cavity (BWS tank)

## Slide 8: Wire interface: Functions and principle
Survey nominal wire resistivity
Record eventual resistivity variation over time
Visualise during movement: - Resistivity variation (vibration) - Thermionic current - RF induced current
Speculate about the temperature during bakeout (~0.5[Ohm/deg])
Slide from C. Morais Pereira

## Slide 9: Wire interface: Implementation
Injected current measurement
Voltage across the wire
Returning current measurement
Chart: Frequency Response of VITH, VI and VWire
Slide from C. Morais Pereira

## Slide 10: Wire interface: Operating point
Chart: Wire Current (mA)
| 0 | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1 | 1.1000000000000001 | 1.2 | 1.3 | 1.4 | 1.5 | 1.6 | 1.7 | 1.8 | 1.9 | 2 | 2.1 | 2.2000000000000002 | 2.2999999999999998 | 2.4 | 2.5 | 2.6 | 2.7 | 2.8 | 2.9 | 3 | 3.1 | 3.2 | 3.3 | 3.4 | 3.5 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.0999999999999996 | 4.2 | 4.3 | 4.4000000000000004 | 4.5 | 4.5999999999999996 | 4.7 | 4.8 | 4.9000000000000004 | 5
RW = 1kΩ | 0 | 0.01 | 0.04 | 0.09 | 0.16 | 0.25 | 0.36 | 0.49 | 0.64 | 0.81 | 1 | 1.21 | 1.44 | 1.69 | 1.96 | 2.25 | 2.56 | 2.89 | 3.24 | 3.61 | 4 | 4.41 | 4.84 | 5.29 | 5.76 | 6.25 | 6.76 | 7.29 | 7.84 | 8.41 | 9 | 9.61 | 10.24 | 10.89 | 11.56 | 12.25 | 12.96 | 13.69 | 14.44 | 15.21 | 16 | 16.81 | 17.64 | 18.49 | 19.36 | 20.25 | 21.16 | 22.09 | 23.04 | 24.01 | 25
RW = 2kΩ | 0 | 0.02 | 0.08 | 0.18 | 0.32 | 0.5 | 0.72 | 0.98 | 1.28 | 1.62 | 2 | 2.42 | 2.88 | 3.38 | 3.92 | 4.5 | 5.12 | 5.78 | 6.48 | 7.22 | 8 | 8.82 | 9.68 | 10.58 | 11.52 | 12.5 | 13.52 | 14.58 | 15.68 | 16.82 | 18 | 19.22 | 20.48 | 21.78 | 23.12 | 24.5 | 25.92 | 27.38 | 28.88 | 30.42 | 32 | 33.62 | 35.28 | 36.98 | 38.72 | 40.5 | 42.32 | 44.18 | 46.08 | 48.02 | 50
RW = 3kΩ | 0 | 0.03 | 0.12 | 0.27 | 0.48 | 0.75 | 1.08 | 1.47 | 1.92 | 2.43 | 3 | 3.63 | 4.32 | 5.07 | 5.88 | 6.75 | 7.68 | 8.67 | 9.72 | 10.83 | 12 | 13.23 | 14.52 | 15.87 | 17.28 | 18.75 | 20.28 | 21.87 | 23.52 | 25.23 | 27 | 28.83 | 30.72 | 32.67 | 34.68 | 36.75 | 38.88 | 41.07 | 43.32 | 45.63 | 48 | 50.43 | 52.92 | 55.47 | 58.08 | 60.75 | 63.48 | 66.27 | 69.12 | 72.03 | 75
RW = 4kΩ | 0 | 0.04 | 0.16 | 0.36 | 0.64 | 1 | 1.44 | 1.96 | 2.56 | 3.24 | 4 | 4.84 | 5.76 | 6.76 | 7.84 | 9 | 10.24 | 11.56 | 12.96 | 14.44 | 16 | 17.64 | 19.36 | 21.16 | 23.04 | 25 | 27.04 | 29.16 | 31.36 | 33.64 | 36 | 38.44 | 40.96 | 43.56 | 46.24 | 49 | 51.84 | 54.76 | 57.76 | 60.84 | 64 | 67.24 | 70.56 | 73.96 | 77.44 | 81 | 84.64 | 88.36 | 92.16 | 96.04 | 100
RW = 5kΩ | 0 | 0.05 | 0.2 | 0.45 | 0.8 | 1.25 | 1.8 | 2.45 | 3.2 | 4.05 | 5 | 6.05 | 7.2 | 8.45 | 9.8 | 11.25 | 12.8 | 14.45 | 16.2 | 18.05 | 20 | 22.05 | 24.2 | 26.45 | 28.8 | 31.25 | 33.8 | 36.45 | 39.2 | 42.05 | 45 | 48.05 | 51.2 | 54.45 | 57.8 | 61.25 | 64.8 | 68.45 | 72.2 | 76.05 | 80 | 84.05 | 88.2 | 92.45 | 96.8 | 101.25 | 105.8 | 110.45 | 115.2 | 120.05 | 125
RW = 6kΩ | 0 | 0.06 | 0.24 | 0.54 | 0.96 | 1.5 | 2.16 | 2.94 | 3.84 | 4.86 | 6 | 7.26 | 8.64 | 10.14 | 11.76 | 13.5 | 15.36 | 17.34 | 19.44 | 21.66 | 24 | 26.46 | 29.04 | 31.74 | 34.56 | 37.5 | 40.56 | 43.74 | 47.04 | 50.46 | 54 | 57.66 | 61.44 | 65.34 | 69.36 | 73.5 | 77.76 | 82.14 | 86.64 | 91.26 | 96 | 100.86 | 105.84 | 110.94 | 116.16 | 121.5 | 126.96 | 132.54 | 138.24 | 144.06 | 150
RW = 15kΩ | 0 | 0.15 | 0.6 | 1.35 | 2.4 | 3.75 | 5.4 | 7.35 | 9.6 | 12.15 | 15 | 18.15 | 21.6 | 25.35 | 29.4 | 33.75 | 38.4 | 43.35 | 48.6 | 54.15 | 60 | 66.15 | 72.6 | 79.35 | 86.4 | 93.75 | 101.4 | 109.35 | 117.6 | 126.15 | 135 | 144.15 | 153.6 | 163.35 | 173.4 | 183.75 | 194.4 | 205.35 | 216.6 | 228.15 | 240 | 252.15 | 264.6 | 277.35 | 290.4 | 303.75 | 317.4 | 331.35 | 345.6 | 360.15 | 375
 | Wire resistivity on the operational system today (PSB, SPS and PS and LHC) |  |  |  |  |  | 7 um wire | 
RW | 1kΩ | 2kΩ | 3kΩ | 4kΩ | 5kΩ | 6kΩ | 15kΩ | 
PWMax | 10mW |  |  |  |  |  |  | 
PWNOM | 5mW |  |  |  |  |  |  | 
IWire | 2.2mA | 1.5mA | 1.3mA | 1.1mA | 1mA | 0.9mA |  | 0.55mA
VWIRE | 2.2V | 3V | 3.9V | 4.4V | 5V | 5.4V |  | 8.25V
RRESOLUTION | 1Ω |  |  |  |  |  |  | 
VRESOLUTION | 2.2mV | 1.6mV | 1.3mV | 1.1mV | 1mV | 0.9mV |  | 0.5mV
ADC Min Resolution | min. 14 bits |  |  |  |  |  |  | 
PWNOM = 5mW
considering
ITHMAX = 0.5mA
Slide from C. Morais Pereira

## Slide 11: Interfaces to external systems for synchronisation, data visualisation and storage
SFP Ethernet
SFP Optical link
SFP Optical link
Long range resolver interface
Wire resistivity & Strain gauge
3-phases PWM
Magnetic stopper
External temperature and vibration
Optical encoder interface
Scanner control, monitoring and supplies BWS_SCMS
SFP Optical link
High Precision Bunch by Bunch Profiling BWS_HPBBP
CPU
Trigger input
T IMING
General Machine Timing (GMT)
Low jitter < 1ns Granularity: 1ms
Control room (CCC)
Expert monitoring
Databases (Logging) Long term storage offline analysis
Trigger input
Trigger output
BST receiver
Beam Synchronous timing (BST): - Bunch synchronisation => 25 ns accurate clock - Revolution frequency synchronisation => for LHC 89us pulse period
Triggers: Wire-scan trigger, post-mortem, acquisition
- Granularity: 89us, low jitter < 1ns
- Scan trigger at ms scale use serial data link
- If we need ns scale, use of dedicated line

## Slide 12: Summary
Somme consideration about the system layout
Acquisition part => More to come next talk
Control and monitoring part => More to come on talk on the actuator
Details on the wire measurements
Connection to external systems
Phase 1 (2013-2014) Build the Scanner Control Monitoring and Supplies (SCMS)
Phase 2 (2015-) Build the High Precision Bunch by Bunch Profiling (HPBBP)
Questions?