## Slide 1
Silicon Photomultipliers
Principle, Performance and Some Applications
Christian Joram
CERN / PH Department
CLIC ECAL lab meeting
25 January 2013
This talk is largely based on a seminar I gave in Dec 2012 at …
I will skip many slides !!!!

## Slide 2
Photodetection
The “classical” application domains of photodetectors in particle physics
Calorimetry
  Readout of organic and inorganic scintillators, lead glass, scint. or quartz fibres  Blue/VIS, usually 10s – 10000s of photons
Particle Identification
  Detection of Cherenkov light  UV/blue  single photons
  Time Of Flight  Usually readout of organic scintillators (not competitive at high momenta)
  Transition radiation (X-rays, not covered in this talk)
Tracking
  Readout of scintillating fibres, blue-green, few photons
… and beyond particle physics
Medical imaging
  Nuclear Medicine: PET & SPECT (g-camera)  Readout of inorganic scintillators
Analytical, industrial, …
158,000 photodetectors in the LHC experiments
4,45 M readout channels
35.4 m2 photosensitive area
C.J., Photodetection in the LHC experiments, NIM A 695 ( 2012) 13-22

## Slide 3
A novel type of photodetector has emerged over the last ~15 years (too late for LHC!)
by now reached a reasonable level of maturity
a number of attractive features, but also some serious limitations.
It has many names like Silicon PM (SiPM), Geiger APD (G-APD), Metal Resistive layer Semiconductor (MRS-APD), Multi Photon Pixel Counter (MPPC), SiPM, SPM,…
SiPM is by now the most used name (I don’t like it).
Outline
Requirements on photodetectors
Principle of operation of the Silicon Photomultiplier
A few words on AX-PET – a demonstrator of an axial PET scanner
A few words on “The Digital SiPM” by Philips
SiPM R&D for a SciFi tracker for LHCb and Calorimetry for ILC

## Slide 4
Requirements on photodetectors
”Typical” requirements on a photodetector
High sensitivity or
What really counts: Photon detection Efficiency
High gain M > 105  we have a chance to see single photons
Small signal fluctuations, low Excess Noise Factor F
ENF impacts energy resolution
Low dark noise / current
Fast timing properties (rise time, fall time, dead time)
Large dynamic range with good linearity
For certain applications: Radiation hardness, magnetic field compatibility, compactness, cost
general definition
(M = 1)
(M ≠ 1)

## Slide 5
(Photonis)
Eg
h
e-
semiconductor  vacuum
Basics of photodetection
Nγ
photodet
Npe
P [W]
I [A]
photodet
Many photosensitive materials are semiconductors, but photoeffect can also be observed from gases and liquids.
Internal photoeffect
External photoeffect
EA = electron
affinity
Internal photoeffect: Eg > Eg
External photoeffect: Eg > Eg + EA

## Slide 6
Family tree of photo-detectors
Photodetectors
“Vacuum”
External photoeffect
Gas
External photoeffect
Solid state
Internal photoeffect
Avalanche gain
Process
Dynodes  PMT
Continuous dynode
Channeltron, MCP
Multi-Anode devices
Other gain process
= Hybrid tubes
TMAE MWPC
TEA + GEM
CsI …
PIN-diode
APD
SiPM
Digital SiPM
CMOS imager
CCD
Doesn’t exist as ‘real’ device
Proposed by G. Barbarino et al., NIM A 594 (2008) 326–331
Proof of principle by CJ et al., NIM A 621 (2010) 171-176
Silicon
anode
HPD
HAPD
(G-APD-HPD)
Lumines-
cent anodes
SMART/Quasar
X-HPD

## Slide 7
(Si) – Photodiodes (PIN diode)
P(I)N type
p layer very thin (<1 mm), as visible light is rapidly absorbed by silicon
High QE (80% @ l  700nm)
Gain = 1
Avalanche photodiode (APD)
High reverse bias voltage: typ. few 100 V
Special doping profile  high internal field (>105 V/cm)  avalanche multiplication
Avalanche stops due to statistical fluctuations.
Gain: typ. O(100)
Rel. high gain fluctuations (excess noise from the avalanche). CMS ECAL APD: ENF = 2 @G=50.
Very high sensitivity on temp. and bias voltage DG = 3.1%/V and -2.4 %/K
Solid-state photon detectors
e h
p+ i(n) n+
g
Hamamatsu S8148.
(140.000 pieces used in CMS barrel ECAL).

## Slide 8
PIN  APD  Geiger mode Avalanche Photodiode (GM-APD)
How to obtain higher gain (= single photon detection) without suffering from excessive noise ?
Operate APD cell in Geiger mode (= full discharge),
however with (passive/active) quenching.
Photon conversion + avalanche short circuit the diode. A single photon (or anything else) is sufficient!
J. Haba, RICH2007
J. Haba, RICH2007
J. Haba, RICH2007
A single-cell GM-APD is just a binary device (=switch).
Info on Ng is lost by the Geiger avalanche.
It will become more interesting when we combine many cells in one device …
Ng

## Slide 9
Signal characteristics and Gain of a GM-APD
ID
  tf = RQCD
RQ ~ >300 kW
 tf ~ 10ns
tr = RSCD
RS ~ kW
CD~ 10 fF
 tr < ns
Imax~(VBIAS-VBD)/RQ
Gain = Q / e
= Imax·t / e
= (VBIAS-VBD)CD / e
J. Haba, RICH2007
J. Haba, RICH2007
Sample of 3 G-APDs
Sample of 3 G-APDs
DV (overvoltage)
G ~ 105 -106 at rel. low bias voltage (<100 V)
dG/dT and dG/dV similarly critical as for APD.
  very fast spike
  ts = RLoadCtotal
The avalanche formation is intrinsically very fast, because confined to a small space.

## Slide 10
GM-APD
Quench
resistor
1mm
100 – several 1000 pix / mm2
Bias bus
Multi pixel GM-APD, called SiPM, G-APD, MPPC, …
The operation of many binary devices in parallel leads to a quasi-analog detector with very
nice properties (photon counting).
20 x 20 pix
-Vbias


Many GM-APDs
g
g
Q
Q
2Q
Quench
resistor
Sizes up to 6×6 mm2 now standard.
Only part of surface is photosensitive!
1 pixel fired
2 pixels fired
3 pixels fired

## Slide 11
Source: http://www.ketek.net/products/sipm-technology/microcell-construction/
Schematic representation of a blue sensitive pn structrue

## Slide 12
The slope of the linear fit of G(Vbias) is the pixel capacitance
Cpixel and hence G increase with the pixel geometrical dimensions and the depletion depth.
  Cpixel ~ 0rS/d
    Cpixel = O(10-100 fF)
N. Dinu & al, NIM A 610 (2009) 423–426
Gain of 3 Hamamatsu devices with 25, 50 and 100 mm cell size.
Gain = DVBD· CD / e

## Slide 13
Thermal/tunneling : thermal/ tunneling carrier generation in the bulk or in the surface depleted region around the junction
After-pulses : carriers trapped during the avalanche discharging and then released triggering a new avalanche during a period of several 100 ns after the breakdown
Optical cross-talk: 105 carriers in an avalanche plasma emit on average 3 photons with an energy higher than 1.14 eV (A. Lacaita et al. IEEE TED 1993). These photons can trigger an avalanche in an adjacent µcell.
Limit gain, increase threshold
add trenches btw mcells
N. Dinu & al, NIM A 572 (2007) 422–426
FBK-irst device
2007 production
th=0.5pe
V=0.5V
V=1V
V=1.5V
V=2V
V=2.5 V
V=3 V
N. Dinu & al, NIM A 572 (2007) 422–426
FBK-irst device
2007 production
  Dark counts due to …

## Slide 14
… and/or cool
N. Dinu et al., NSS Conf Record (NSS/MIC), 2010 IEEE , vol., no., pp.215-219,
Hamamatsu 1mm2 device
Overvoltage must take into account the
T-dependence of VBD
Reduce noise by a factor ~2 every 8 K

## Slide 15
Limited linearity
is simply a consequence of the limited of available pixels. Some pixels get ≥2 hits.
B. Dolgoschein (Light06)
Ntrue/Navailable ~ 0.5
 Nfired/Ntrue ~ 0.8

## Slide 16
SiPM designs (just examples)
Hamamatsu HPK (http://jp.hamamatsu.com/)
25x25µm2, 50x50µm2, 100x100µm2 pixel size
1x1mm2
3x3mm2
Arrays
1x4mm2
1x4 channels
6x6 mm2
2x2 channels
FBK-IRST
50x50µm2 pixel size
1x1mm2
2x2mm2
3x3mm2
4x4mm2
4x4mm2
2x2 channels
3x3 cm2
8x8 channels
SensL (http://sensl.com/)
20x20µm2, 35x35µm2, 50x50µm2, 100x100µm2 pixel size
3.16x3.16mm2
4x4 channels
3.16x3.16mm2
4x4 channels
6 x 6 cm2
16x16 channels

## Slide 17
 | PMT | SiPM
QE (VIS) | 0.2-0.4 | 0.2-0.7
Gain | 106 @ O(kV) | 106 @ O(50V)
Timing | Tr ~ O(1ns) TTS ~ O(100 ps) | Tr ~ O(1ns) TTS O(100 ps)
Dynamic range | O(106) | O(103)
ENF | 1.1-1.5 | ~1
Dark noise rate | O(kHz/cm2) | O(MHz/mm2)
Single photon sensitivity/ counting | / | /
Magnetic field immunity |  | 
Robustness & compactness |  | 
PM vs SiPM

## Slide 18
A new geometrical concept for a PET scanner
with SiPM readout.
P. Beltrame et al., The AX-PET demonstrator—Design, construction and characterization, NIM A 654 (2011) 546-559
P. Beltrame et al., The AX-PET Concept: New Developments And Tomographic Imaging, 2011 IEEE NSS Conf. Rec. MIC22-5

## Slide 19
Positron Emission : p  n + e+ ne
β+ decay of various radionuclides
Positron Annihilation : e+e−  gg (Eg = 511 keV ). 2 photons emitted almost “back-to-back”
The above two processes define the fundamental limits to the spatial resolution of PET
‣ Finite positron range (ρ): annihilation position ≠ emission point, ρ depends on the energy of the positron (i.e. on the radioisotope). r ~ mm
‣ Non-collinearity of the 2 photons: residual momentum of the e+e- at the annihilation => the 2 photons are emitted with a small deviation from 180° (Δθ ~ 0.5°) => blurring of the spatial resolution RFWHM ~ 0.0022 x D [mm]
PET basics
Radionuclide | Half life | Emax_e+(MeV)
11C | 20.4 min | 0,96
15O | 122 sec | 1,73
18F | 109,8 min | 0,63
22Na | 2.6 years | 0,55

## Slide 20
In practice…
Inject the radiotracer into the body (0.1-0.2 mCi per kg body weight)
Wait for up-taking period (~h)
Acquire data (~0.5 h for full body)
Feed the data into the reconstruction algorithms
(2/3D) image of the activity concentration
PET is an “in-vivo” functional imaging technique
It produces a (quantitative) image of the radio-tracer concentration
Use of F-18 labeled glucose  image of the bodies metabolism
PET detector principle :
coincidence of two 511 keV photons define a line of record.
Take projections under all angle
(2/3D) Tomographic reconstruct of data

## Slide 21
Quality of a PET image depends on
Detector resolution  crystal size and readout type)
Counting statistics  efficiency, detector thickness, scintillator material
Scatter fraction (Compton scattering in body and detector)  energy resolution
Accidental coincidences and pile-up  coincidence window
…
A.Del Guerra, CERN Academic Training 2009
50000 counts 100000 counts 500000 counts 2M counts
Statistics matters ….
The challenges in PET are today (and since a long time)
Improve resolution (approach physical limits)
Improve efficiency
Reduce background
Combine it with other imaging modalities, in particular MRI

## Slide 22
The Axial PET concept was conceived to simultaneously optimize resolution and efficiency.
“Standard” radial PET geometry
The parallax dilemma
L
Crystals need to have a minimum thickness L
Efficiency for pair detection
la = photon attenuation length of crystal
L = la  e2 ~ 40%, L = 2 la  e2 ~ 75%, …
la = 1-2 cm (depending on material, see below)
A standard PET does not measure the depth of interaction (DOI) in the crystal.
a
dp
This introduces a parallax error
The resolution in the off-center region degrades significantly
Solution: reduce L (bad e2) OR measure DOI OR invent a different geometry

## Slide 23
The AX-PET concept
From short radially oriented,
block readout crystals …
… to long, axially oriented, individually readout crystals
PMT
1
PMT
2
PMT
3
PMT
4

## Slide 24
Our implementation of the AX-PET concept
x
y
z
scintillation light
wavelength
shifted light
Individual readout of crystals. Crystal address gives transverse coordinates (x,y)
Small crystal cross section  parallax error practically negligible !
Additional layers allow to increase sensitivity
 Spatial resolution and sensitivity are decoupled and can be optimized independently
How to read crystals ?
How to measure axial coordinate ?
x
y
z
SiPM
SiPM

## Slide 25
It consists only of two camera modules à
48 crystals (6 layers x 8 crystals)
156 WLS strips (6 layers x 26 strips)
7
3.5
Crystals are staggered by 2 mm. Crystals and WLS strips are read out on alternate sides to allow maximum packing density.
The layers are optically separated from each other.
Since 2008, the AX-PET collaboration (Bari, Cagliari, CERN, Michigan, Ohio, Oslo, Tampere, Valencia, Zurich, < 10 FTE ) has built and tested a fully operational PET demonstrator scanner.

## Slide 26
AX-PET components
Density [g/cm3] | 7.1
Attenuation length for 511 keV [cm] | 1.2
Wavelength of maximum emission [nm] | 420
Refractive index at W.L. of max. emission | 1.81
Light yield [photons/keV] | 32
Average temperature coefficient [%/K] | -0.28
Decay time [ns] | 41
Intrinsic energy resolution [%, FWHM] | ~8
Natural radioactivity [Bq/cm3] | ~300
Effective optical absorption length [mm] | ~ 420
The scintillator crystals are Ce doped LYSO (Lu1.8Y.2SiO5:Ce) single crystals, fabricated by Saint Gobain and commercialized under the trade name PreLude 420.
The main characteristics are:
Dimensions: 3 x 3 x 100 mm3
PMT (with borosilicate window)
Hamamatsu SiPM (MPPC)
Emission / detection efficiency (a.u.)
One end is read out, the other end is mirror-coated (evaporated Al-film).
This determines the transverse resolution (1 module).
Expect: sx = sy = 3/sqrt(12) ~ 0.86 mm
(2 mm FWHM)

## Slide 27
The WLS strips are of type EJ-280-10x from Eljen Technologies
Shift light from blue to green
Density: 1.023 g/cm3
Absorption length for blue light: 0.4mm (10 x standard concentration)
Index of reflection: 1.58
Decay time: 8.5ns
Size: 0.9×3×40mm3
sz = 3/sqrt(12) ~ 0.86 mm
(2 mm FWHM).
Use COG algorithm on top.
One end is read out, the other end is mirror-coated (evaporated Al-film).

## Slide 28
Two different types are used:
MPPC S10362-33-050C  Crystal readout
sensitive area: 3 × 3 mm2.
pixel size: 50 × 50 mm2 (60 × 60 = 3600 pixels)
typical operational voltage: 70 V.
Typical gain: 7·105.
Photon detection efficiency: ~40 % (400 nm)
MPPC 3.22×1.19OCTAGON-SMD  WLS readout
custom specific device
sensitive area. 3.22 × 1.19 mm2
The pixel size is 70 × 70 mm2,(46 × 17 = 782 pixels).
The photodetectors are SiPMs of type MPPC from Hamamatsu.
The electrical connections to the MPPCs are made
via special Kapton cables soldered to the MPPCs.
Some MPPC ‘features’
Gain is temperature dependent (~5%/°K)  T monitoring
Limited dynamic range  saturation effects  correction
Thermal noise, O(100 kHz) at 1 photon level.  set threshold

## Slide 29
Fully assembled module (48 crystals, 156 WLS stips)
+ lots of cables: 204 x (bias + signal out)

## Slide 30
1 channel.
1 module.
Mod 1: Energy sum of all LYSO crystal of 1 module. 22Na source
LL (50 keV)
HL (~450 keV)
HHL (~550 keV)
photopeak
Compton
1275 keV
Good hit in 1 module = LL × HL × HHL
GH1
GH2
Tcoinc = 20 ns
Trigger
Front-end electronics
Fully analog, but relatively slow …
Max Readout rate ~ 20 kHz

## Slide 31
module
Na-22
Source
0.25 mm Ø
2-D moving station
PMT
tagger
I. Single module characterization
II. Coincidence of 2 modules
Module in coincidence with a tagging scintillator
Use of different tagging crystals
module 2
module 1
source
Distance between modules = 15 cm
Test set-ups
III. Gantry for tomographic measurements
Rotating support

## Slide 32
Single module characterization
Typical charge spectrum from 1 LYSO crystal
Energy resolution of all 96 LYSO crystals
Separate calibration experiment:
511 keV ~ 1100 pe in SiPM
intr.
 Expect RE ~ 10.7 %
<RE> = 11.7 % (FWHM)

## Slide 33
Energy calibration
First correct for possible temperature changes
SiPM gain drops by ~5% per degree K.

## Slide 34
Energy calibration
176Lu  176Hf* + ß-
306 keV
203 keV
306 keV
203 keV
55 keV
89 keV (IC)
Intrinsic radioactivity
511 keV
Small non-linearity due to limited pixel number
of SiPM (3600)
5% effect
The formula would predict a 15% effect
LYSO photons arrive over ~100 ns. Some hit
pixels will have recovered when the 2nd hit arrives.
Lutetium in LYSO is radioactive.
~ 250 Bq per crystal.

## Slide 35
The axial coordinate …
511 keV
Cluster
noise
centre-of-gravity
algorithm
Calibration (separate set-up):
511 keV  100 pe in WLS cluster
sz ~ 0.75 mm
(average over whole module, after correction for e+ annihilation physics).

## Slide 36
First coincidence measurements
Point-like 22Na source
Photoelectric events only (1 hit crystal per module)
Draw “LOR” (pure geometrical, no tomographic reconstruction)

## Slide 37
AXIAL PLANE
TRANS-AXIAL PLANE

## Slide 38
Look at the YZ distribution at x=0 (“confocal reconstruction”)
sz = 0.63 mm (coincidence) is compatible with the single detector resolution.
After correction of the e+ annihilation physics:
sz (coinc.) ~ sz (detector) / sqrt(2)
y coordinate is quasi-discrete (crystal positions).
RMS =
0.89 mm

## Slide 39
Experimental check: Are there parallax errors ?
Face-to-Face
30○
30○
Oblique
22Na point-like source
There is however an unavoidable degradation of the resolution, when the source is very close to one detector. The coincidence resolution looses the 1/sqrt(2) ‘bonus’
As expected from the design, there is no noticeable parallax error, neither in the axial nor transaxial plane.

## Slide 40
Ready for tomographic imaging
Basics of tomographic imaging (in AX-PET this is the speciality of the team from Valencia (M. Rafecas et al.)
A PET scanner measures projections of the activity distribution. Reconstruction means to calculate from these projections the actual 3D distribution of the activity. There are two groups of methods:
1. Analytical algorithms
(usually based on sinograms)
LOR
PET ring
“sinogram”
a 2D histogram
of all projections g(s,q)
s
q
angle q
(Filtered) Back Projection
(2D case)
One doesn’t know where on the LOR the count occurred. Therefore, “back projection” means re-distributing the counts uniformly to all pixels which fall within the projection path.
1 angle
4 angles
64 angles
The more angles,
the better!
2. Iterative algorithms
(mainly used in AX-PET, see back-up slides)

## Slide 41
20○
FOV
‘’Step and shoot”
Rotate object in 9 steps of 20 deg.
We mimic a 18-module scanner, however with a very limited FOV (width and length of 1 module). We allow only coincidences between opposite modules.
FOV
‘’Step and shoot”
1.) Rotate object in 9 steps of 20 deg.
2.) Rotate 1 module by 20 deg.
3.) Rotate object in 18 steps of 20 deg.
Now we allow coincidences between opposite ±1 modules.  Larger FOV.
How do we get all these angles with just 2 modules ?
Price to pay:
Longer acquisition times (several h).
Decaying activity
Changing conditions (pile-up, dead time)

## Slide 42
The AX-PET set-up has exactly the size of a EURO palette. Very useful for measurement campaigns at other places
Measurements with phantoms or animals can’t be made at CERN!
ETH Zurich, PET lab, 2010
AAA (Saint GEnis), F) 2010
AAA (Saint GEnis), F) 2011
ETH Zurich, PET lab 2012
phantoms
animals

## Slide 43
Phantoms are just containers which can be filled with a radiotracer. We usually work with F18 in aqueous solution. Some ink is added to see air bubbles.

## Slide 44
capillaries :
L = 3 cm
∅ = 1.4 mm
pitch = 5 mm
Profiling
Resolution
counts [a.u.]
FWHM [mm]
ETH 2010

## Slide 45
Hot & Cold rods for contrast
Homogeneous cylinder for assessing the ability to reconstruct homogeneous distributions
Series of small rods for resolution
Three regions in the same phantom to address three different aspects
34 mm Ø
 63 mm 
1 mm (i.e. < Resolution)
2 mm
3 mm
4 mm
5 mm
Reconstructed 1 mm rod
=> FWHM ~ 1.6 mm
NEMA mouse phantom

## Slide 46
(A)
(B)
(A)
Rods oriented parallel to Z axis
Mini Deluxe phantom
Rods oriented perpendicular to Z axis
High density resolution phantom
Diameter (75 mm) is larger than the extended FOV

## Slide 47
June 2012, small animals tests campaign at ETH Zurich, Radiopharmaceutical Inst.
Possibility to use their GE Explore Vista PET/CT scanner as reference. This is a dedicated small animal scanner with 1.45 x 1.45 mm crystals. Block readout. DOI via Phoswich approach.
one mouse, FDG
one (young) rat, FDG
one (young) rat, 18-F
=> organs structure
=> bones, skeleton structure
Rat, FDG - Details of the heart
AXPET
Explore Vista
Small Animal Tests

## Slide 48
Rat, F-18 - Details of the skeleton
AXPET reconstructed image
CT image from Explore Vista
(not enough activity left for the PET acquisition with Explore Vista)
axial field of view

## Slide 49
Rendered F-18 image of rat
(click to animate)

## Slide 50
And finally … Digital SiPM
With the GM-APD being a binary device…
one could also conceive a fully digital SiPM, where the output is just the digital count of the number of fired cells.
Philips

## Slide 51
Different from analog SiPMs: Upon the detection of a photon, the avalanche is actively quenched using a dedicated transistor, and a different transistor is used to quickly recharge the diode back to its sensitive state.
v
1 digital cell (area depends on type, e.g. 30 x 50 mm)
Digital SiPM
Philips

## Slide 52
Digital SiPM
Compared to the analog technology, the digital one (so far only offered by Philips) has a number of
advantages
Integration of bias supply, amp, TDC, counter…
Fast active quenching  no afterpulses
Possibility to de-activate noisy cells  potentially lower dark noise
Reduced sensitivity to voltage and temperature variations
Compactness
Possibility to add local intelligence
… problems shared with analog
High dark noise (a discharging cell doesn’t know whether it is digital or analog)
Signal saturation (limited number of cells)
… and also has some drawbacks
The local electronics is a source of heat  cooling advisable
The readout functionality is designed into the sensor. In case of mismatch with the needs, relatively expensive modifications of the sensor/FPGA may be required.
Front and back sides of a 64 channel digital tile (DPC6400-22-44) or (DPC3200-22-44)
Philips
Packaged module, as delivered to clients (DPC3400-22-44). Includes a 100 mm thick protective glass layer.
38.5 x 38.5 mm2

## Slide 53
S < 1 ms Expect readout rate up to O(1 MHz)
trigger
Digital SiPM Readout
Every non-validated trigger leads in principle to the recharging of all cells. Without cooling, the device can loose efficiency/availability.
Philips

## Slide 54
We got fascinated by the idea of a digital AX-PET module, mainly in view of a possible TOF-PET extension.
Aim for dual sided readout of long crystals to get rid of propagation delays
Luckily, crystals and WLS would just fit (without big compromises) on the existing Philips 64 channel DSiPM modules DLS-6400
In reality, we have 2 x 8 WLS strips

## Slide 55
Principle of TOF-PET
Idea: use time information to localize annihilation along the line of response. LOR does no longer contribute uniformly to all pixels but according to the measured time and its resolution.
Images contain less noise and show better contrast (particularly noticeable for large scanners and big patients)
Dt = 2·Dx/c
Need dt = 66 ps
for dx = 1 cm
Gain of a TOF system over a non-TOF system is
G = 2D / c· dt
D = dimension of object being imaged.

## Slide 56
2 x 8 WLS strips
4 LYSO crystals
dSiPM
Kapton
cable
Water cooling
Verification of normal AX-PET characteristics
Two modules setup. Single sided readout plus WLS strips.

## Slide 57
~25% effect
Saturation effect is much larger:
More detected p.e.
Fewer pixels (3200 vs 3600)
Recharging only after readout!
Energy calibration and resolution
(single sided readout)
In the end we get a comparable
Energy resolution.
(analog AX-PET was 11.7 %).

## Slide 58
22Na source of 250 microns diameter, 150 mm between the two modules
Slightly better than with the analog AX-PET set-up (1.48 mm)
Confocal reconstruction in axial direction (WLS array)

## Slide 59
readout
readout
Aluminium coating
t1
>
Δz
t2
This is only possible after correcting for light paths using WLS information
Coincidence time resolution
(single sided readout)

## Slide 60
Energy calibration and resolution
(Double sided readout)
raw
saturation
corrected
2250  2800
(25%)
DE = 10.45% (FWHM)

## Slide 61
Coincidence time resolution by mean timing
(Double sided readout)
300 mm
100 mm
TDC bins a 19.5 ps
= 96 ps
FWHM = 225 ps
-
Small differences in timing of pixels (offsets) can be eliminated.
Mean Timing corrects automatically for light propagation times in the crystals. The resolution becomes fully independent of the source position in the FOV.
Vertical scan
(1 module +
small tagger)
Horizontal scan

## Slide 62
Brief look at applications in particle physics
The SiPM technology wasn’t mature enough when the choices were made for LHC (>10 years ago).
But we start to see SiPMs in R&D for the LHC upgrade.
The obvious applications are readout of scintillators and scintillating fibres.
Also demonstrated: Detection of Cherenkov light (even with digital SiPM).
First major project was T2K (scintillator tracker + ECAL). ~65k SiPMs.
The CALICE collaboration builds HCAL and ECAL prototypes based on small scint. plastic tiles (ILC). ~8k SiPMs.
CMS is working on an upgrade of HCAL barrel (SiPMs replacing HPDs).
LHCb is working on a study of SciFi tracker to replace Si strips and gas straw tubes.
~ 300k SiPMs
CLIC detector study for a W/scint ECAL. Many 100k SiPMs.

## Slide 63
photo-detector requirements:
insensitive to magnetic field (~ 5T)
coupling with a scintillator (blue emission)
stability (T-control or T-monitoring)
rel. large dynamic range
low cost
High granularity hadronic calorimeter optimised for the Particle Flow measurement of multi-jets final state at the ILC
e+ e- collider (1 TeV)
SiPM
1 mm²
216 tiles/layer (38 layers in total) ~8000 channels
CALICE tests with MePHI/PULSAR, HAMAMATSU, KETEK SiPMs
Readout of SiPMs by the SPIROC ASIC (LAL)
3 x 3 cm² plastic scintillator tile
with embedded WLS fiber + SiPM
PROTOTYPE
SiPM in the ILC HCAL and ECAL (R&D)

## Slide 64
DESY 3 GeV electron testbeam
Channel energy calibration by ~3 GeV electrons (MIPs).
Multi-photon peaks allow simple absolute calibration.
Different SiPM couplings to scintillator tile
with WLS fiber (for MEPHI SiPM)
direct coupling (for MPPC)
hole (“dimple”) for more uniform light yield close to photodetector
Calibration & monitoring
test beam (absolute)
LED monitoring system (stability)
SiPM in the ILC HCAL (R&D)

## Slide 65
R&D for LHCb SciFi tracker
Npe
5 m
5 m
3 x ( X-U-V-X )
= 12 det. layers
1 det. layer = 5 fibre layers
250 mm Ø scint fibres.
~20 pe, sx ~ 50 mm
Detector concept follows largely the work of the Aachen group
(S. Schael et al.)
NIM A 622 (2010) 542–554
~ 300 k SiPM channels

## Slide 66
R&D for LHCb SciFi tracker
Main challenges:
Rad hardness of SiPM (10 years ~ 6·1011 n/cm2)
(Rad hardness of scint. Fibres (10 years ~ 30 kGy))
Main observation: dark current / dark rate increase linearly with the NIEL
In analogy to a standard pn detector…
Id~α·Φ(1 MeV n eq.) ·Veff ·G (α = 3·1017 A·cm)
~10 MHz/mm2
~100 MHz/mm2
~1 GHz/mm2
I. Musienko, SiPM workshop, CERN 2011.
But also gain starts to drop
Devices with smaller cell size and fast cell recovery survive longer

## Slide 67
increase due to scintillation of the glass window
time (ns)

## Slide 68
Cooling helps a lot. Every 8K reduce the dark noise by a factor 2
+20°C
0°C
-20°C
-40°C
Dark current (Hamamatsu SiPM module of 128 ch à 0.25 × 1.1 mm2 ~ 35 mm2)
But also reducing overvoltage, increase threshold, clustering, short shaping time, …
can save you 
Nominal operation point
After 1·1011 n/cm2
2.5·1010 n/cm2
Non-irrad (20°C)
G. Häfeli, EPFL
(LHCb SciFi)
Current [uA]
Bias Voltage [V] (corrected for 25oC)

## Slide 69
SiPMs are photodetectors with great properties (sensitivity, gain, compactness, simplicity) and a few limitations (dark noise, rad. tolerance, dynamic range, rel. high cost).
In the past decade they have stimulated huge interest (by research teams, industry) and their properties improved a lot. They enabled new applications and will continue to do so.
Digital SiPM technology is a consequent next step. Their acceptance by the market depends on many other aspects.
AX-PET is an example of a detector concept which would be very difficult to implement without SiPMs. It overcomes the sensitivity-resolution dilemma of standard PET detectors. Its acceptance by the market …
Summary / Conclusions
Many thanks to my AX-PET colleagues Chiara Casella (ETH Zurich) and Matthieu Heller (CERN) for many plots and schematics used in this talk. Thanks also to Felix Sefkow for CALICE material.

## Slide 70
Many thanks for your interest !

## Slide 71
At long l, temperature effects dominate
Light absorption in Silicon

## Slide 72
! not to scale !
Extrapolation to d=0 reduces beam spot size to the true source size (250 mm) and also eliminates the effect of non-collinearity
d
small
tagger
… and its resolution
sz ~ 0.75 mm (1 module)
(after subtraction of contribution of finite positron range (0.23 mm)).

## Slide 73
Initial estimate of the activity distribution f(x,y)
Calculate Forward Projection
Compare
Converged ?
Mij
The Forward projection
requires the knowledge of the so-called system matrix Mij
Its elements describe the probability of the activity in the j-voxel to be detected by the i-LOR.
Geometry specific. Needs to be calculated only once, but ca be VERY large.
2. Iterative algorithms

## Slide 74
Initial estimate of the activity distribution f(x,y)
Calculate Forward Projection
Compare
Converged ?
Converged ?
Correction matrix

## Slide 75
Updated estimates of the activity distribution f(x,y)
Calculate Forward Projection
Compare
Converged ?
Converged ?
Correction matrix
Iterative approaches are more accurate than FBP but very CPU consuming.
All AX-PET data shown in the following, is reconstructed by an iterative method.