## Slide 1: SPS kick response measurements
H. Bartosik, S. Cettour Cave, B. Salvant
Acknowledgements: J. Wenninger
MSWG, 26. March 2013

## Slide 2: Introduction
Measurement of the orbit response to individual corrector kicks
  ±25μrad for each orbit corrector (COD)
    Two consecutive cycles
    The orbit difference between positive and negative kick is the orbit response
  In practice
    Injection on well corrected orbit
    Then all CODs set to zero and finally the COD of interest is powered
    This is done using YASP

## Slide 3: Orbit response analysis
SVD analysis of orbit response data
  Relation between beam position ui for corrector kicks θj is given by Response matrix Rij
  Analysis principle
    Difference vector between measurement and MADX model to minimize
    Sensitivity matrix S for BPM calibration factors and COD calibration factors ci
    Solving linear equation
    Perform Singular Value Decomposition on S and compute correction for calibration factors
  Reference: J. Wenninger, Orbit Response Measurements at the SPS, CERN-AB-2004-009
μ: phase advance
β: beta function
Q: tune
… only Rmod depends on calibration factors!

## Slide 4: Calibration results - Horizontal

## Slide 5: Orbit response: MDHB.61804
COD of special type …
Problem with magnet or calibration issue?

## Slide 6: Orbit responses at: BPH.13008
Largely wrong calibration!
Error too large to be caused by beta-beat …

## Slide 7: Orbit responses at: BPH.61408
The sensitivity of this BPM is less than 10% of the expected value!

## Slide 8: Calibrated measurement vs. model - Horizontal
Standard deviation of difference between model and calibrated measurements
  Typical standard deviation of about 0.3mm
  A few BPMs with large standard deviations, due to
    Small calibration value and thus reduced resolution
    Large measurement noise
    Spurious readings for some measurements
  Acceptable limit arbitrarily set to ~0.5mm for horizontal BPMs

## Slide 9: Orbit responses at: BPH.20608
Orbit response cannot be fitted consistently at this BPM  large difference between model and measurements even after “calibration”

## Slide 10: Results - Horizontal
  This COD is having significantly “reduced strength”
    MDHB.61804 (gain=0.55)
  These BPMs are not working:
    BPH.12408, BPA.21706, BPA.21805, BPH.22208, BPH.23008, BPCE.41705, BPCE.41931, BPH.52008, BPH.60808, BPH.62808;
  These BPMs have large calibration errors:
    BPD.11906 (gain=1.24)
    BPH.13008 (gain=1.48)
    BPH.13608 (gain=0.16)
    BPA.21605 (gain=1.24)
    BPH.31808 (gain=1.35)
    BPH.41608 (gain=1.28)
    BPH.60408 (gain=1.20)
    BPH.61008 (gain=1.24)
    BPH.61408 (gain=0.08)
    BPH.62008 (gain=1.43)
  These BPMs have large standard deviation after “calibration” with respect to the model:
    BPH.20408, BPH.20608, BPH.21008, BPH.23408, BPH.62408;

## Slide 11: Calibration results - Vertical

## Slide 12: Orbit responses at: BPV.11108
This BPM is simply inverted …

## Slide 13: Orbit responses at: BPCE.61805
Largely wrong calibration!
Error too large to be caused by beta-beat …

## Slide 14: Calibrated measurement vs. model - Vertical
Standard deviation of difference between model and calibrated measurements
  Typical standard deviation of less than 0.2mm
  A few BPMs with large standard deviations, due to
    Small calibration value and thus reduced resolution
    Large measurement noise
    Spurious readings for some measurements
  Acceptable limit arbitrarily set to ~0.3mm for horizontal BPMs
BPV.63508

## Slide 15: Orbit responses at: BPV.63508
Looks like BPM has very little sensitivity – but in 4 cases it gives even larger values than expected ???

## Slide 16: Results - Vertical
  These BPMs are not working:
    BPD.11833, BPD.11906, BPA.21605, BPA.21706, BPA.21805, BPCN.22708, BPV.30908, BPCE.41801, BPV.42108, BPV.62508;
  These BPMs have large calibration errors:
    BPV.11108 (gain=-0.97)
    BPV.11706 (gain=1.39)
    BPCN.12308 (gain=0.08)
    BPCN.12508 (gain=0.09)
    BPA.21931 (gain=0.23)
    BPV.22108 (gain=1.34)
    BPV.31308 (gain=1.28)
    BPV.51908 (gain=0.19)
    BPV.52108 (gain=0.09)
    BPV.60108 (gain=1.23)
    BPV.61508 (gain=1.34)
    BPCE.61805 (gain=2.30)
    BPV.62308 (gain=1.36)
    BPV.63508 (gain=0.23)
  These BPMs have large standard deviation after “calibration” with respect to the model:
    BPV.11308, BPV.20308, BPV.20508, BPV.20908, BPV.23108, BPV.30108, BPV.63308;

## Slide 17: Summary and Conclusions
Horizontal BPMs
  10 BPMs are not working at all (some of them in hot zones)
  10 BPMs seem to have grossly wrong calibration (some of them are special type)
  5 BPMs have large spread around expected readings (after “calibration”)  should be investigated
Vertical BPMs
  10 BPMs are not working at all (some of them in hot zones)
  14 BPMs seem to have grossly wrong calibration (some of them are special type)
  7 BPMs have large spread around expected readings (after “calibration”)  should be investigated
Horizontal closed orbit correctors
  Identified 1 closed orbit corrector with reduced kick strength (is it hardware or simply calibration problem?)
  Average fitted kick strength is 0.98±0.07
Vertical closed orbit correctors
  Work properly
  Average fitted kick strength is 1.00±0.03