## Slide 1: Beam Instrumentation Challenges
Dr. Rhodri Jones
Head of the CERN Beam Instrumentation Group
1st oPAC Topical Workshop
Grand Challenges in Accelerator Optimisation
CERN, Geneva, Switzerland - June 26th – 27th 2013

## Slide 2: Introduction to Beam Instrumentation
Beam instrumentation provides “eyes” for machine operators
  The instruments that observe beam behaviour
  An accelerator can never be better than the instruments measuring its performance!
What does work in beam instrumentation entail?
  Design, construction & operation of instruments to observe particle beams
  R&D to innovate or improve existing techniques to fulfill new requirements
  A combination of many disciplines
    Applied & Accelerator Physics and a lot of Mechanical, Electronic & Software Engineering
  A fascinating field of work!
What are the main beam parameters measured?
  Beam Position to control location of the beam in the accelerator
  Beam Intensity to measure operational efficiency
  Beam Loss to ensure safe operation
  Transverse Beam profiles
  Longitudinal Beam profiles
to optimise operation

## Slide 3: Trends in Accelerators
Bigger, Faster, Better!
  High energy machines for particle physics
    LHC, HL-LHC, HE-LHC (CERN)
    CLIC/ILC (R&D)
  High current machines
    Neutron sources – material science
      SNS (US), ESS (Sweden), CSNS (China), IFMIF (Japan), …
    Neutrino Production
      T2K (JPARC, Japan), NuMI/Nova (FNAL, US), CNGS (CERN), Project-X (FNAL, US), …
  High Brightness machines
    X-ray Free Electron Lasers – probing complex, ultra-small structures
      LCLS (SLAC, US), European XFEL (DESY, DE), …
Compact or Exotic!
  Rare radioactive isotope machines
    FAIR (GSI, DE), HIE-ISOLDE (CERN)
  Anti-matter machines
    ELENA (CERN), FAIR (GSI, DE)
  Medical proton/ion therapy machines

## Slide 4: Challenges for Beam Instrumentation
Unprecedented request for precision
  Positioning down to well below the micron level
Treatment of increasingly more data
  Bunch by bunch measurements for all parameters
Dealing with high beam powers
  Non-invasive measurement techniques
  Robust and reliable machine protection systems
Dealing with the ultra-fast
  Measurements on the femto-second timescale
Dealing with the ultra-low
  Measurement of very small beam currents

## Slide 5: Challenges for Beam Instrumentation
The Quest for Unprecedented Precision

## Slide 6: Beam Position Monitoring – The Challenges
Typically one of the largest BI systems
  LHC has 1100 BPMs – CLIC drive beam requires 40000! This implies:
    Precision engineering at low cost
    Simple, robust, yet very highly performing electronics
      Digitise as soon as possible (now doable with advances in ADCs & digital treatment systems)
      Radiation Tolerant (improves S/N, minimises expensive cabling BUT needs significant testing time)
  XFEL’s and Linear Colliders increasingly asking for very high resolution
    Well below 1mm for single pass

## Slide 7: Improving the Precision for Next Generation Accelerators
Standard BPMs give intensity signals - need to subtract them to obtain a difference which is proportional to position
  Difficult to do electronically without some of the intensity information leaking through
    When looking for small differences this leakage can dominate the measurement
Solution – cavity BPMs allowing sub micron resolution
  Design the detector in such a way as to only collect only the difference signal
    Dipole Mode TM11 proportional to position & shifted in frequency with respect to monopole mode
f / GHz
U / V
Frequency Domain
TM01
TM11
TM02
U~Q
U~Qr
U~Q
Courtesy of D. Lipka,
DESY, Hamburg
TM01
TM11
TM02

## Slide 8: Today’s State of the Art BPMs
Monopole Mode
Dipole Mode
Courtesy of D. Lipka,
DESY, Hamburg
Courtesy of D. Lipka & Y. Honda
Obtain signal using waveguides that only couple to dipole mode
  Further suppression of monopole mode
Prototype BPM for ILC Final Focus
  Required resolution of 2nm (yes nano!) in a 6×12mm diameter beam pipe
  Achieved World Record (so far!) resolution of 8.7nm at ATF2 (KEK, Japan)

## Slide 9: Challenges for Beam Instrumentation
Maximising the Available Data

## Slide 10: Maximising the Available Data - Example
APD
TDC
synchrotron light
LHC turn clock
Electrical pulse
Arrival time
filter
Longitudinal Bunch Shape
LHC Longitudinal Density Monitor
  Aims:
    Profile of the whole LHC ring with 50ps resolution
    High dynamic range for ghost charge measurement
  Method:
    Single photon counting with Synchrotron light
      Avalanche photodiode detector
      50ps resolution TDC
Courtesy of A. Jeff

## Slide 11: LDM On-line Correction
Main bunch
Satellites
Afterpulsing
Deadtime
Ghost bunches
Corrected for deadtime
Corrected for deadtime and afterpulsing

## Slide 12: LDM Results
Results
  Able to profile the whole ring within a matter of minutes
  Critical input for accurate luminosity calibration of the experiments

## Slide 13: Challenges for Beam Instrumentation
Dealing with high beam powers
  Non-invasive measurement techniques

## Slide 14: Dealing with High Beam Power - The Issues
Traditional invasive diagnostics no longer suitable
  Materials do not withstand energy deposition
    Burning of holes or breaking of wires
  Use of superconducting RF structures
    Risk of contamination on breakage or sputtering of the intercepted material

## Slide 15: Measuring the Size of High Power Beams
Synchrotron Light Diagnostics
  Only for electron & very high energy proton/ion machines (LHC)
    Difficult to separate the light from the beam for linear accelerators
  Difficult to get absolute calibration
    Image correction factors typically bigger than the beam size!
  Additional challenges lie in fast cameras & signal treatment chains for bunch by bunch measurement
LHC Synchrotron Light Diagnostics

## Slide 16: Measuring the Size of Intense Bunches
Standard OTR not an option for short intense bunches (main issue for LCLS)
  Signal swamped by Coherent OTR from micro-bunching within the main bunch
One proposal is to go to luminescent Screens
  There will still be COTR at polished surfaces
  Resolution might be a problem
Investigation of temporal suppression of COTR
  Scintillation screen + fast gated camera
  Proof-of-principle at FLASH, DESY
Resolution of scintillation screens needs study
  Synergy with ion beam diagnostics
  Similar studies ongoing for FAIR
Al coated Si OTR screen,
COTR light,
Coherent SR
LuAG screen,
COTR &
scintillation light
LuAG screen
+100ns delay
Only scintillation light
Courtesy of M. Yan
Courtesy of K. Wittenburg

## Slide 17: Non-Invasive Beam Size Measurement
Laser wire scanner
  Good candidate for H- & electrons
    Can measure down to micron level
  Needs conversion to turn-key instrument
    Reliable laser system
    Laser distribution to multiple measurement stations
Courtesy of A. Alexandrov

## Slide 18: Non-Invasive Beam Size Measurement
Beam Induced Fluorescence
Beam
N2 ground state
e-
N2 excited state
Photon
emitted

## Slide 19: Non-Invasive Beam Size Measurement
Beam Induced Fluorescence
  Issues
    Sensitivity to radiation
    Low signal yield  vacuum bump
    Drift of excited gas molecules
     broadening of profile
Ion beam
Blackened walls
Vacuum gauge
Valve
Viewport
150mm flange
Lens, Image-Intensifier
and CCD FireWire-Camera
N2-fluorescent gas
equally distributed
Courtesy of M. Schwickert

## Slide 20: Non-Invasive Beam Size Measurement
Residual Gas Ionisation
  Some 10 times more sensitive than BIF
  More complicated to build
    High voltage network
    Guiding magnetic field
  Image broadening due to space charge
  Ageing of detectors an issue
Courtesy of M. Schwickert

## Slide 21
Electron Beam Scanner
  Derive beam profile from deflection of low energy electron beam scanned across the high energy proton beam
Non-Invasive Beam Size Measurement
Courtesy of W. Blokland

## Slide 22
Electron Beam Scanner – how it works
  Look at the deflected projection of a tilted sheet of electrons
Non-Invasive Beam Size Measurement
dy
dx
Quadrupoles
Deflector
Electrons 
Multiple scans
take the derivative to get the profile
Courtesy of W. Blokland

## Slide 23
Electron Beam Scanner – the challenges
Non-Invasive Beam Size Measurement
Courtesy of W. Blokland

## Slide 24: Challenges for Beam Instrumentation
Dealing with high beam powers
  Robust and Reliable Machine Protection Systems

## Slide 25: Beam Loss Monitors – New Materials
Diamond Detectors
  Fast & sensitive
  Used in LHC to distinguish bunch by bunch losses
  Investigations now ongoing to see if they can work in cryogenic conditions
Courtesy of E. Griesmayer

## Slide 26: Cryogenic Beam Loss Monitors
Silicon
  Successfully used at 1 K at CERN in 1976
Diamond
  Successfully in use as LHC BLM at room temperature
  Better resistance to radiation than Si at room temperature
  Less leakage current than Si at room temperature
  Does it work in liquid helium?
Liquid helium ionisation chamber
  Advantage : no radiation hardness issue
  Disadvantage : slow (charge mobility of 0.02 cm2/V/s)

## Slide 27: Cryogenic Beam Loss Monitors
Diamond results
  Single particle (response averaged from ~5000 pulses)

## Slide 28: Challenges for Beam Instrumentation
Dealing with the ultra-fast
  Measurements on the femto-second timescale

## Slide 29: Ultra-Short Bunch Length Diagnostics
Next Generation FELs & Linear Colliders
  Use ultra short bunches to increase brightness or improve luminosity
How do we currently measure such short bunches?
  Transverse deflecting cavity
p+ @ LHC | 250ps
H- @ SNS | 100ps
e- @ ILC | 500fs
e- @ CLIC | 130fs
e- @ XFEL | 80fs
e- @ LCLS | 75fs
Destructive Measurement

## Slide 30: Non Destructive – Electro-Optic Sampling
Spectral Decoding
Temporal decoding
Limited to >250fs by laser bandwidth
Limited to >30fs by sampling laser pulse &
response of electro-optic materials

## Slide 31: Non Destructive – Electro-Optic Sampling
Shortest measured electron bunch profile
  79.3  7.5 fs at DESY FLASH
  Using 65um thick GaP crystal
Courtesy of A. Gillespie

## Slide 32: Challenges for Beam Instrumentation
  Dealing with the ultra-low
  Measurement of very small beam currents

## Slide 33: Measuring very low Intensity Beams
Measuring rare ions and antiprotons
  Currents well below 1mA resolution of traditional Direct Current Beam Transformers
Superconducting QUantum Interference Devices
  Can overcome the limits of conventional systems
  Both low and high temperature versions have been developed
Challenges
  Shielding the detector from stray fields
    Earth’s magnetic field ~ 50µT
    Field produced by 100 nA beam at 10 cm ~ 0.2 pT !
  Making the system compact
Cryogenic Current Comparator (CCC)
  Charged particle beam induces screening currents
  Toroidal pick-up coil of niobium with ferromagnetic core is used to measure magnetic field
  Signal from coil is fed through a superconducting transformer for impedance matching to the input coil of the readout DC SQUID.

## Slide 34: Conclusion
Particle Beam Diagnostic systems are continually evolving to meet the requirements of
  ever more powerful and sophisticated accelerators
  ever more demanding specifications
Many synergies between the developments required for all types of new accelerator facilities
  Ideal subject for collaboration on an international level
  Was already underway between the major accelerator laboratories worldwide
  Now strengthened and extended to universities and industry via European Networks