## Slide 1: Materials studies and tests at CERN
Stefano Sgobba
EN-MME-MM
CSEM-CERN meeting , 23 May 2013
- Mandate and expertise - Equipment
Examples of material studies and tests
Perspectives

## Slide 3
3/22

## Slide 4
Sectioning
Linear precision Saw. Buehler IsoMet 4000
Precision diamond wire saws. Well 3242 and 3241
Microstructural characterization. Specimen preparation
Grinding & polishing
Manual polishing machines. LamPlan M.M. 8027 S. and Plamopol 2
Automatic polishing machines. Buehler Phoenix 4000 and LamPlan M.M. 8055
Vibratory polishing machine. Buehler vibromet 2
Automatic electrolytic polishing-etching. ATM Kristall 620 with electrolytic cell
Mounting
Hot mounting presses. Buehler Simplimet 1000 and 2000
Cold mounting pressure chamber. LamPlan M.M. 806
4/22

## Slide 5
Optical microscopy
Various stereo microscopes.
Metallographic microscope. Leica DMRM
Objectives for magnifications from x16 to x1500
Image analysis system
Digital portable microscope. Keyence VHX-1000
Resolution max 54 MPixel
Objectives 0 - x50, x20 - x200 and x100 - x1000
Microstructural characterization. Observation and analysis
Electron microscopy
SEM. Leo 430i
Chamber 190 mm x 300 mm x 265 mm
SE and BSE imaging detectors. Max resolution 4 nm
Beam 0.5 to 30 kV, 1pA to 500 nA
EDS analyser Oxford Isis.
FE-SEM. Zeiss Σigma
Chamber ø365 mm x 275 mm
SE, BSE, in-lens SE imaging detectors. Max resolution 1.5 nm
Beam 0.1 to 30 kV, 4 pA to 20 nA
EDS analysis Oxford Inca with 30 mm2 SDD detector
EBSD analysis HKL Chanel 5
Spectroscopy and X ray diffraction
Powder XR diffraction. Siemens D5000
Cu and Cr X-ray sources
Vertical and horizontal goniometers
Portable OES analyser. Oxford Instruments PMI MasterPro
Calibrations for steels, Al, Ni, Cu and Ti alloys
UV probe for S and P
Morphology of a Nb coating on Cu structure, original magnification 50000 x
5/22
New optical microscope being purchased in 2013
New SEM (replacement LEO) in 2015

## Slide 6
Scanning Electron Microscopy and microanalysis support
Carbon wires from SPS beam scan: region degraded by the beam (left) and unused reference (right, same scale)
NEG coatings
Thin films
Nb-coatings,
HIE-Isolde cavities
C-coatings for SPS upgrade
orientation map
inverse pole figure
Microstructural characterisation by EBSD of Bi2212 wire

## Slide 7
Mechanical testing
Tensile testing machines
Two column electromechanical universal testing machine UTS 200
Load cells 1 kN, 20 kN and 200 kN, stroke 800 mm
Knives and clip-on extensometers
Tensile grips, compression plates, bending tools
System for tests at 77 K and 4.2 K, 25 kN load cell
Single column press ZPM 1000-500. Load cell 1 kN, stroke 500 mm
Hardness
Hardness. Wolpert 2RC testor
Load 1 to 250 kg
Brinell, Vickers and Rockwell indenters
Micro hardness automatic tester Wilson Hardness 402MVD
Load 10 g to 2 kg, Vickers indenter
7/22
New dynamic mechanical testing unit in 2014
100 kN cryostat for tensile testing and eventually fracture mechanics at 4.2 K in construction (2013)

## Slide 8
8/22
Mechanical testing at 4.2 K, the past…

## Slide 9
Evolution of the equipment
Carbon potentiometer: range: up to 50 mm resolution: 0.1 mm
core
coil
LVDT:
range: up to 3 mm
resolution: 0.5 µm
Temperature probe, CERNOX CX-1050-SD-30 (Lakeshore Cryotronics):
thermal response time: 15 ms at 4 K
resolution: ± 15 mK
Internal load cell:
immersed in liquid He, based on a deformable 316LN ring equipped by strain gauges
conditioner: MVD2555 of HMB
9/22

## Slide 10
10/22
Evolution of the equipment
D. J. Marcinek, Experimental Study of Discontinuous Plastic Flow, Phase Transformation and Micro-damage Evolution in Ductile Materials at Cryogenic Temperatures, Master of Science Thesis, CERN and CUT, 2009
B. Skoczen, J. Bielski , S. Sgobba , D. Marcinek , Constitutive model of discontinuous plastic flow at cryogenic temperatures, International journal of plasticity, 26 (2010) 1659-1679
Sampling frequency up to 20 kHz!

## Slide 11
D. J. Marcinek, Experimental Study of Discontinuous Plastic Flow, Phase Transformation and Micro-damage Evolution in Ductile Materials at Cryogenic Temperatures, Master of Science Thesis, CERN and CUT, 2009
OFE Cu, serrated yielding recorded during necking (engineering stress-engineering strain plot)
11/22

## Slide 12
Major Tool and Machines, 11/04/2012, Indianapolis, IN (USA). ITER CS magnets,, precompression structure
Mechanical testing at 4.2 K, the present and future
CERN
CERN + CEME

## Slide 13
Present involvements
13/22
CS Nitronic 50 single piece, 15 m forged precompression tie plate, G & G Steel Inc., Russellville, AL /US
Courtesy of Kind

## Slide 14
14/22
Radiographic testing, RT
Mobile X-ray computed tomography system. RayScan Mobile
Mini-focus x-ray source focal spot size 0.4/1.0 mm, 225 kV, 3.6 – 8 mA,
Flat panel detector, active area 410x410 mm, 4 MPixel
Software for 3D reconstruction and visualization
X-ray sources. Philips 160 kV, 0. 10 – 160 kV, 4 and 10 mA
Sieffert isovolt 160 kV, 1 to 10 mA (mobile device)
Computer radiography system
High definition scanner Duerr HD-CR35NDT, spatial resolution 40 µm
High resolution monitor 3 MP
Controlled area (bunker) of 30 m2
Level 2 certified operators
Non destructive testing
Ultrasonic testing, UT
Portable flaw detectors. Krautkramer USN 60
    Gain 0 dB to 110 dB, probe from 0.25 MHz to 25 MHz,
Portable flaw detector with Phased Array mode. GE Phasor XS
    Gain 0 dB to 40 dB, probe from 0.5 MHz to 10 MHz,
    Various sectorial or linear transducers
Immersion tank scanner
    Scan-surface: 500 mm x 500 mm, 0.1 MHz to 150 MHz
Level 3 and 2 certified operators
LHC warm modules, spring of the module VMTSA-4L2-3259-left side
Visual testing, VT, two level 2 certified
Penetrant testing, PT, two level 2 certified

## Slide 15
Non destructive testing
 Computed Radiography Testing (RT) of LSS warm modules. Completion of the radiography campaigns during LHC technical stops. 1767 modules examined, 107 non-conformities identified to be repaired during LS_1, 46 nights.
 Computed RT of CMS RF modules
 Computed RT of ATLAS RF modules
set-up in the tunnel
X-Ray tomography of ITER CC He-inlets: thin defects precisely identified and confirmed by metallography
7 5
He-inlets
courtesy of RX-Solutions
15/22

## Slide 16
Delivery and
test description
ITER CC He Inlets samples delivered to CERN by ASIPP in March 2013
Two welded with buttering (B-L and B-S) and two without buttering (WB-L and WB-S)
Laminography (planar tomography) of all four samples performed by RX Solutions /FR, allowing to access defects in the whole weld volume with higher resolution than through computed RT
Classification following ISO 6520-1 and judged with respect to EN ISO 5817 - level B
detector
X-ray source
He inlet sample
B-S B-L
WB-S WB-L
Laminography of WB-S sample

## Slide 18
3D tomography
Technique description
Specimens placed on a mobile and rotatable stage
Motorized zoom
Full inspection of samples up to 240 mm length
Sealed microfocus source, max. tension 150 kV
HR area image sensor, 1920 pixels x 1536 pixels
Detection surface: 200 mm x 250 mm
14 bits – 16000 gray levels
1-30 images /s
2nd sample of the VPI
11 separate 360° scans
were combined to compile a final single 3D object reconstruction while avoiding shadows due to steel conduits

## Slide 19
3D tomography
View in myVGL software

## Slide 20
20/22
DFBAA-1L
 Complex failure analysis of DFBA flexible metal hoses,
TE-TM, 01/09/2009
G. Stachowiak, A.W. Batchelor, Engineering Tribology, 3rd ed., Elsevier Butterworth-Heinemann, Amsterdam (2005)
Progression of stress corrosion cracking
Pitting
Final separation by overload (microvoid coalescence)
Divisional seminar, 01/02/2006
QRL: corrosion leaks on the tube F close to the fix point 

## Slide 21
21/22
21/22
100 °C
RT
92 days, Pb 99.99%+, erosion-corrosion, impingement
erosion and impingement quantified by 3D metrology
2008-2009, material selection and corrosion studies for n_TOF
S. Sgobba, L. Marques Antunes Ferreira, Material Selection and Corrosion Studies for a High Energy Physics Experiment: the Neutron Time-of-Flight Facility at CERN, NACE Conference Papers, March 01, 2011, pages: 1 – 15
b
a
117 days
23 days
28 days

## Slide 22
Conclusions
CERN-wide (and beyond) activity
Limited but selected and highly specialized equipment
Covers a wide range of materials, welds and components
Provides support to CERN stores (specifications, identification of suppliers, audit…)
Several decades of experience, including audit and follow-up of industrial production
In house expertise completed by an extensive network
Prone to intervene in or pilot activities with a strong emphasis on materials or controls (quality control for LS_1 weld activities; RIAC; ITER…)
Section very exposed to international exchanges and contacts:
  Industrial partners - National labs, academy
  CERN-ITER cooperation agreement (IO, DAs via IO)
- Training of students and fellows, hosted a large (3 digit) number of trainees since the creation of the section in the early 90ies
22/22