## Slide 1: Micro-strip based detection systems: advances and new technological developments
Yoshinobu Unno
KEK

## Slide 2: Your Lecturer
Yoshinobu UNNO

## Slide 3: Your Lecturer
Mt. Tsukuba
..is from KEK, Japan
Yoshinobu UNNO
Professor of KEK
.. his office
Home of Belle, K2K, … experiments

## Slide 4: Your Lecturer
Mt. Tsukuba
..is from KEK, Japan
.. his office
… working for an experiment at CERN
Yoshinobu UNNO
Professor of KEK
ATLAS experiment at CERN
Home of Belle, K2K, … experiments

## Slide 5: Your Lecturer
Mt. Tsukuba
..is from KEK, Japan
.. his office
… working for an experiment at CERN
Home of Belle, K2K, … experiments
Yoshinobu UNNO
Professor of KEK
ATLAS experiment at CERN
Design, construction, … and upgrade, of the silicon tracker
- together with many colleagues
- for the great discovery

## Slide 6: Japan - Wonderland
One of the advantage is the industry, specially electronics
AKIhabara, Tokyo

## Slide 7: Japan - Wonderland
Big electronics giants like HITACHI, …
AKIhabara, Tokyo

## Slide 8: Japan - Wonderland
Big electronics giants like HITACHI, SONY, …
AKIhabara, Tokyo

## Slide 9: Japan - Wonderland
AKIhabara, Tokyo
But, we have benefitted from a smaller company …

## Slide 10: Japan - Wonderland
AKIhabara, Tokyo
who has a long history of collaboration with our fields …
Nobel prize for opening the Neutrino astronomy

## Slide 11: Japan - Wonderland
AKIhabara, Tokyo
Now, we are benefitting from the company…
Nobel prize for opening the Neutrio astoronomy
ATLAS Supplier Award

## Slide 12: Japan - Wonderland
AKIhabara, Tokyo
Nobel prize for opening the Neutrio astoronomy
ATLAS Supplier Award
for the silicon detectors: ATLAS strips ~92%, CMS strips ~97%

## Slide 13: Contributions by Japanese teams in the ATLAS construction
In addition, many Japanese industries provided high quality detector components: Hamamatsu Phonics, Kawasaki Heavy Industries, Toshiba, Kuraray, Arisawa, Fujikura, etc
13
1200 TGC chambers and
320K ch. L1 Electronics of
endcap muon trigger system
KEK, Tokyo, Kobe, Nagoya...
Superconducting Solenoid, KEK
6000 sensors and 980 modules of barrel SCT system, KEK, Tsukuba, Okayama, Hiroshima …
400k ch. of TDC chips for MDT system, KEK
22 m

## Slide 14: This lecture
Micro-strip based detection systems
  advances and new technological development
It is not a monopoly of ATLAS nor HPK nor Japan. It is an example of our fields.
  Other experiments: CMS, LHCb, …
  Other industries: Microns, CiS, …
Content
  Brief overview of the current ATLAS silicon microstrip tracker (SCT)
  Issues and achievement of the LHC tracker
  New technological development for the high-luminosity LHC tracker
  Understanding the underlying physics – TCAD simulation

## Slide 15: ATLAS Inner Detector
ID ¼ volume
Installation into ATLAS
ID=Silicon (Pixel + Strip) + TRT
  Radius~ 1 m, L=5.4 m
Position (f) resolutions
  Pixel: ~15 µm (50 µm pitch)
  Strip (axial-stereo pair): ~17 µm (80 µm pitch)
  TRT: ~ 22 µm (130 µm drift reso., 36 sampling)
Why Silicon?
Notes: p: momentum (GeV/c)
B: magnetic field (tesla)
L: track length in bending plane

## Slide 16: ATLAS Inner Detector
Barrel modules
  Double-side, stereo readout (40 mrad)
  Sensors: 2 x (6.4 x 6.4 cm2) /side x 2 side (top and bottom）
    4 in. FZ crystal wafer, <111> and some <100>
    80 µm pitch, ~12.6 cm strip lenth, 768 strips
  Hybrids: 6 ABCD chips (128 ch)/side
    Cu/Polyimide flex circuit + Carbon-carbon substrate
Red letters: Contributions of Japan
  Japan: ~980/2112+spares in barrel modules
ID ¼ volume
SCT barrel module
Hybrid
Microstrip
sensor

## Slide 17: Silicon Detector (LHC)
Wire bonding of hybrid-sensor and sensor-sensor: ~3200 bonds
Assembly station: aligning two sensors in <5 µm
Completed modules
Mounted on the barrel cylinder
Assembled ID barrel section
Then, installed into ATLAS
Placement robot
at Oxford
at CERN
in Japan, …
in Japan, …

## Slide 18: Silicon Detector (LHC)
Silicon area
ATLAS ~60 m2
LHC silicon microstrip detectors are the largest silicon trackers ever built.
  A scale: cost of the sensor ~1 million Euro/m2
Silicon microstrip detector is the “must” for large area coverage.
CMS ~200 m2

## Slide 19: Particle fluence
Yearly fluence of particles
  in the unit of 1-MeV neutrons equivalent per cm2
  Luminosity: 1034 cm-2s-1, integrated:100 fb-1/yr
End of life fluence at SCT (r=30 cm)
  2×1014 1-MeV neq/cm2 (including 50% uncertainty in pp cross section)
TRT
SCT
PIXEL
(1-MeV neq/cm2)
(from ATLAS TDR)

## Slide 20: Non Ionizing Energy Loss (NIEL)
1-MeV neutron equivalent

## Slide 21: Radiation Damage Studies
Radiation damage of silicon
  the 1st, in our field… 30 yrs ago
    T. Kondo et al, Radiation Damage Test of Silicon Microstrip Detectors
    Proc. of the 1984 Summer Study on the Design and Utilization of SSC, June 23-July 13, 1984, Snowmass, Colorado, pp. 612-614
The messages were
  The prevailing opinion was that silicon vertex detectors were not possible at 1033 luminosity, but…
  It was shown that silicon is rad-hard, little pulse-height change, cooling needed.

## Slide 22: Radiation Damage Studies
Since then, radiation damage studies are continued in Japan, Europe, US., and elsewhere
  Two papers were already published in 1988, (25 yrs ago)
  T. Ohsugi, ... T. Kondo, … K. Yamamoto .., “Radiation Damage in Silicon Microstrip DetectorsT”, Nucl. Instr. Meth. A265(1988)105
  M. Nakamura,…T. Kondo, "Radiation Damage Test of Silicon Multistrip Detectors", Nucl. Instr. Meth. A270(1988)42, using the irradiated sensor by 800 GeV protons

## Slide 23: Increase of leakage current
Radiation Damage in Silicon Microstrip Detectors
  T. Ohsugi, ... T. Kondo, … K. Yamamoto .., Nucl. Instr. Meth. A265(1988)105
Also, temperature dependence of bulk leakage current

## Slide 24: Type inversion of the silicon
M. Nakamura,…T. Kondo, "Radiation Damage Test of Silicon Multistrip Detectors", Nucl. Instr. Meth. A270(1988)42, using the irradiated sensor by 800 GeV protons
Abstract:
……. The effective impurity density decreases with fluence up to ~4x1013/cm2, but for greater fluences, it increases. This may indicate the type conversion of the bulk silicon ………

## Slide 25: Evolution of depletion voltage
  A thorough study of the radiation damages has been made by RD50 collaboration. But, also done elsewhere…
  E.g. Michael Moll, Ph.D Thesis, 1999.
24 GeV/c proton irradiation (n-type silicon)
K. Hara et al.,
IEEE Trans. Nucl. Scie. 56 (2009) 468
70 MeV proton irradiation
(p-type silicon)
4 in.
6 in.
6 in. FZ is as same as DOFZ or MCZ

## Slide 26: Signal degradation in LHC Silicon Sensors
Charge trapping effect
  ← Most of signals at around the strips (see Appendix)
Depleted region
  p-in-n → p-in-p (requires “full depletion”)
  n-in-n → n-in-p (works under “partial depletion”)
Strip sensors: max. cumulated fluence for LHC
Pixel sensors: max. cumulated fluence for LHC
(p-in-”p”)
(n-in-”p”)
RD50 collaboration

## Slide 27: Choice of LHC Experiments
Cost consideration and compromises
p-in-n:
  single-side process (lower cost)
  requires full depletion, high voltage operation
n-in-n
  double-side process (higher cost)
  works under partial depletion, less requirement for high voltage
Experiment | Type | Wafer
ALICE pixel | p-in-n | standard FZ
ATLAS pixel | n-in-n | oxygenated
ATLAS strips | p-in-n | standard FZ <111> (some <100>)
CMS pixel | n-in-n | standard FZ
CMS strips | p-in-n | standard FZ <100>
LHCb VELO | n-in-n | standard FZ

## Slide 28: Silicon Detector (LHC ATLAS)
LHC: p-readout in n-bulk (p-in-n)
Silicon sensor principle
  Deplete the bulk by holding the bias voltage at p-n junction
LHC ATLAS: p-in-n strip sensor
  N-bulk: conventional
  Cheaper than other options
  Need full depletion
    500 V max operation/specification
after type-inversion
n-bulk
strip region
edge region
p+
n+
p-n junction
Depletion
“p”-bulk
p+
n+
p-n junction
Depletion

## Slide 29: Radiation damage – Surface effect
The interfacial region is a single-crystal silicon followed by a monolayer of SiOx, incompletely oxidized silicon, then a strained region of SiO2 roughly 10-40 A deep.
Interface traps (Qif) and fixed oxide charges (Qf) exist, (as a consequence of thermal oxidation)
Oxide trapped charges (Qot) can be created by radiation.
Qf and Qot are “positve” and attract electrons in the Si-SiO2 interface.
⊖ ⊖ ⊖ ⊖
Electrons attracted to the (Positive) “Fixed Oxide charges” (and interface trapped charges
SiO2 passivation
Incident particles
Vacancies
Interstitial silicons
+ + + + + +
- - - - - -
+ + + + + +
- - - - - -
S.M. Sze, Physics of Semiconductor Devices, 2nd ed., p380
Surface effect
Bulk effect

## Slide 30: 1st Visualization of Microdischarge
High bias voltage → High electric field → avalanche breakdown
  Breakdown field ~ 30 V/µm in silicon
Visualization with an infra-red sensitive camera
T. Ohsugi et al., Nucl. Instr. Meth. A432 (1994) 22

## Slide 31: Other examples of hot spots
  Y. Takahasi et al., http://dx.doi.org/10.1016/j.nima.2012.04.031
Y. Unno et al., Nucl. Instr. Meth. A Supplement 636 (2011) S24

## Slide 32: Microdischarge after Irradiation
Hot electron images confirm that
  hot spots were observed first at the edge of the bias ring, and then at the inside of the edge metal.
  the highest electric field is at the bias ring (n+ implant), not at the edge ring (p+ implant).
CYRIC proton irradiated
1x1014 neq/cm2
10 uA at 2000 V
-15 °C
S. Mitsui et al.,
Nucl. Instr. Meth. A699 (2013) 36-40

## Slide 33: Design of the sensor to high bias voltage
SiO2 passivation
Bias ring
+ + + + + +
- - - - - -
+ + + + + +
- - - - - -
Surface effect
Move the highest electric field from Si (to SiO2)
→ Extended electronde
Under the condition of
Same potential in implant and metal
①
②
Reduce the oxide fixed/trap charges
→ <100> crystal orientation
Less “dangling” bonds in interface
<100>: ~1010 ions/cm2
<111>: ~1011 ions/cm2
Guard ring
Edge ring
③
Optimize the edge structure
→ Width, guard ring
between the bias and the dicing edge
HPK design:
width ~ 1 mm
minimum # guard rings, i.e. 1
⓪
Clean and high quality process
→ little irregularity at high field
where HPK is appreciated.

## Slide 34: ATLAS SCT in operation
Radiation damage monitoring
  Leakage currents are well consistent with the expectation
99.3% modules are working
  30/4088 modules were disabled due to LV, HV, Cooling problems, …

## Slide 35: LHC Upgrade (HL-LHC)
Higgs – Great discovery of the century
Need high statistics
  Study of properties
  Search for “something” in TeV mass region
Phys. Lett. B 716 (2012) 1-29 Fig.7c
Higgs search
1 TeV
Higgs coupling
Spin-Parity 0+ favoured

## Slide 36: Schedule for HL-LHC
~2022
2018
2013/14
2009
Start of LHC
Run 1: 7 and 8 TeV centre of mass energy, luminosity ramping up to few 1033 cm-2 s-1, few fb-1 delivered
2030
Phase-II: High-luminosity LHC. New focussing magnets for very high luminosity with levelling
Injector and LHC Phase-I upgrades to go to ultimate luminosity
LHC shut-down to prepare machine for design energy and nominal luminosity
Run 4: Collect data until > 3000 fb-1
Run 3: Ramp up luminosity to 2.2 x nominal, reaching ~100 fb-1 / year accumulate few hundred fb-1
Run 2: Ramp up luminosity to nominal (1034 cm-2 s-1), ~50 to 100 fb-1
HL-LHC
LHC
IBL installation

## Slide 37: Inner Detector Upgrade (HL-LHC)
2022 – Full tracker replacement
  Area: ~200 m2
  Silicon strips: ~1×1015 neq/cm2
Silicon sensors
  Max. 1000 V operation
  Full depletion might not be possible…
Full silicon tracker
TRT → Strip
Strip
Pixel
~1x1016
~1x1015
Strip sensor module, example

## Slide 38: Cost-effective n-in-p planar sensor
for heavy radiation environments
Bulk radiation damage
  one way to be “p” type
n+ readout
  p-n junction to allow getting signals from “partially” depleted sensor
Special in n+ readout
  conductive layer in the surface
    ~MΩ/square
    due to the electrons attracted to the oxide trap/fixed charges
    no junction effect at the n+ implant
  the electron layer must be
    interrupted (p-stop), or
    cancelled (p-spray)
Single-side mask process
n-in-n
Radiation damage
n-bulk → “p”-bulk
Double-side mask process
n+
p+
p+
p+
n+
+ + + +
- - - -
+ + +
- - -
Radiation damage
p-bulk → “p”-bulk
n-in-p
n+
p+
p+
p+
+ + + +
- - - -
+ + +
- - -
Diffusion process
ATLAS choice for strip sensors for HL-LHC
(LHC pixel sensors of ATLAS, CMS)
+ + +
- - -

## Slide 39: n-in-p sensors for HL-LHC
Sensors with the p-stop isolation
Operable to 1000 V bias voltage.
  Equivalently, suppressing “microdischarge” breakdown up to ~1000 V
How?
  Those 0, 1, 2, backed by 3
  In addition, protection against beam splash: punch-through-protection (PTP) structure
n-in-p
Radiation damage
p-bulk → “p”-bulk
n+
p+
p+
p+
+ + + +
- - - -
+ + +
- - -
① Optimization of p-stop structure
② Optimization of edge width
③ Understanding the physics
← Technology CAD (TCAD) simulation
⓪ Hardening the strip edges

## Slide 40: Study of required edge width
Results are from N-type wafer
Thickness (as is, thinned)
  320 (W5), 200 (W7,13) µm
Edge implantation
  N+ or P+
N-sub/N-edge
N-sub/P-edge
N-sub
P+
P+
N+
GR
Edge
N-sub
P+
P+
P+
GR
Edge
Width varied at one edge

## Slide 41: Underlying physics of the edge width
SiO2 passivation
Bias ring
+ + + + + +
- - - - - -
+ + + + + +
- - - - - -
Surface effect
Guard ring
Edge ring
Square root of V_bias is linearly dependent on the edge distance
  Reflecting the depletion along the surface
Distance can be ≤500 µm for the bias voltage up to 1 kV
… Different story if the side wall is implanted e.g., - active edge
Lateral depletion along the surface

## Slide 42: Required width after Irradiation
Required width is ~450 µm to hold 1000 V.
  At around 1x1013, the required edge space is more than 450 µm, but also the depletion voltage is decreased less than that of non-irrad. and anyway it is much less than 1000 V.
  At higher fluences, the required width is less than that of the non-irrad.
S. Mitsui et al, NIMA 699 (2013) 36-40

## Slide 43: P-stops between N-implants
Problems - Hot spots
  IR image overlaid on visual image
  Microdischarge = Onset of leakage current
How to optimize the structures to reduce the electric fields?

## Slide 44: P-stop Structures Optimization
TCAD simulations
Y. Unno et al., Nucl. Instr. Meth. A636 (2011) S118–S124
… and comparison with test structures

## Slide 45: Optimization of the p-stops
P-stop
  away from the n-implant
  symmetric location
N-implant
  pitch not too narrow nor not too wide
Once known, simple.
Stereo strip section
1000 V
10 cm x 10 cm strip sensor in 6 in. wafer

## Slide 46: Technology CAD (TCAD)
TCAD started to build the links between the
  semiconductor physics and electrical behavior
  to support circuit design
Modern TCAD consists of
  Process simulation, and
  Device simulation
Originated from the work of
  Prof. Robert W. Dutton and his group at Stanford Univ.
Widely used in semiconductor industry
  to reduce the development cost and time
  to understand the physics behind
    that is even impossible to measure
TCAD: Computer Aided Design for Semiconductor Technology
The core is the “Finite Element Analysis”.
  The numerical analysis method with modern computer.
Process simulation
Device simulation
MOS transistor

## Slide 47: Brief History
1977: Prof. Dutton, Stanford
Process/Device simulator
SUPREM-I (1D)/PISCES
1979: Technology Modeling Associates (TMA/Synopsys)
TSUPREM4 (2D)/MEDICI
1989: Silvaco International
ATHENA (2D)/ATLAS
1989: Integrated Systems Engineering AG (ISE)/Synopsys)
DIOS (2D)/DESSIS
1992: TMA
TAURUS (3D TSUPREM4/DEDICI)
1993: Prof. Law, Florida
Process sim: FLOOPS (3D)
2002: ISE
FLOOPS (3D)
2005: Synopsys
Sentaurus (3D TAURUS)
In Japan,
1996: 3D HyENEXSS (Selete/TCAD Int.)
Selete: Consortium of 10 semiconductor co.
2011: 3D HyENEXSS (Selete)
Project ends
Prof. Robert W. Dutton
(from Stanford TCAD Home page)
TMA⇒AVANT!/1998⇒Synopsys/2001
ISE⇒Synopsys/2004

## Slide 48: Process Simulator Device Simulator
Process steps
  Oxidation
  Deposition
  Etching
  Ion implantation
  Annealing
Mostly for process experts
  Unless you know the process parameters, you have no way to simulate.
Solving equations
  Poisson eq. (ψ, n, p)
  Current continuity eq. Jn, Jp (ψ, n, p)
  Heat conduction eq. (“Drift Diffusion model) (TL)
  …
Four equations and four variables
  potential ψ, electron-density n, hole-density p, and lattice-temperature TL
ion-implantation process (M.C.-model)

## Slide 49: Caveat
Jungle of semiconductor physics models and parameters
  Device simulator e.g.,
  Transport models
  Mobility models
  Generation-recombination models (SRH, Auger, II, trap, surface…)
    SRH: Shockley-Read-Hall model
    II: Impact Ionization model
Finite Element method
  3D vs. 2D
    3D: Usually “very” time consuming
    2D: Most of the cases, good enough
  Meshing: resolution vs. time
  Convergence of calculations
    Try and error for finding best procedures (method, physics model)
The real caveat is
  “What you get is what you put.”
  Although semiconductor industry is trying to simulate perfectly, we may still miss models, e.g., radiation damages

## Slide 50: TCAD Simulations
Semiconductor Technology Computer-Aided Design (TCAD) tool
  ENEXSS 5.5, developed by SELETE in Japan
  Device simulation part: HyDeLEOS
N-in-p strip sensor
  75 µm pitch, p-stop 4x1012 cm-2
  150 µm thickness
  p-type bulk, Neff=4.7×1012 cm-3, VFDV=80 V at 150 µm
Radiation damage approximation:
  Increase of acceptor-like state → Effective doping concentration
  Increase of leakage current → SRH model
  Increase of interface charge → Fixed oxide charge

## Slide 51: Bulk leakage current
Community has a view that
  the leakage current increases with an introduction of levels near the middle of the forbidden band,
  with the energy of band gap being half (of the full gap), the leakage current flows order of magnitude larger…
Unfortunately, we have no freedom to change/add a program to the ENEXSS, but
  we can simulate the leakage current by modifying the model parameters to an unrealistic world…
After irradiation, the current increases as a function of fluence
∆I /V ~ α × f (neq/cm2)
α ~ 4×10-17 (A/cm) : damage constant
E.g.,
Volume = 75 µm x 1 µm x 150 µm = 1.13 x 10-8 cm3
f=1x1015 neq/cm2
∆I ~45 nA
xxxxx
Donor level
Acceptor level
Damage created levels

## Slide 52: Shockley-Reed-Hall (SRH) Model
Leakage current: SRH model
  Generation-recombination of carriers (electrons and holes) by thermal effect
  An, Ap : model parameters
    Decrease them as though increasing temperature
ni: intrinsic carrier density,
n, p: electron, hole carrier density

## Slide 53: Radiation Damage Approximation
Black: non-irrad.
  Neff=4.7×1012 cm-3, An, Ap = 1.0
Green: Irrad.
  Increase of full depletion voltage, Neff=1.5×1013 cm-3
  Increase of leakage current, An, Ap = 1×10-8
Potential in bulk
Leakage current
  Backplane at 200 V

## Slide 54: Interstrip Resistance, Rint
Bulk - depleted by bias condition
Decrease of interstrip resistance after irradiation
  is quantitatively explained by the increase of leakage current.
  Other factors, the effective doping concentration nor the oxide interface charge, do not change the interstrip resistance.
  In retrospect, it is natural that the current is the other manifestation of the resistance.

## Slide 55: Electric potential of p-stop - Introduction of Si-SiO2 interface charge -
Non-irrad:
  Neff=4.7×1012 cm-3,
  SRH An, Ap=1.0,
  Fixed Oxide Charge =1×1010 cm-2
Irrad:
  Neff=1.5×1013 cm-3,
  SRH An, Ap=1×10-8 ,
  Fixed Oxide Charge =1×1012 cm-2
Potential in bulk, 2D display
n+-Implant
n+-Implant
p-stop
Backplane (-200 V)

## Slide 56: Electric Potential between Strips
Electric potential of p-stop
  decreases as the interface charge increases positively,
  increases as the interface charge increases negatively.
Measurement confirms that the interface charge is positive.
Positive
charge
Negative
charge
Measurement
Y. Unno et al., 10.1016/j.nima.2013.04.075

## Slide 57: Breakdown at High Voltages
Under the “Irradiated” condition
Breakdown occurs at high voltage at the n+ edge, although the p-stop edge was the higher electric field initially.
The rate to increase of the electric field at the p-stop edge is saturating at higher voltage.
The p-n junction eventually overtakes the highest electric field by the time of breakdown.
Why?
Silicon avalanche breakdown
voltage ~300 kV/cm

## Slide 58: Insight into the physics
Electron inversion layer is diminishing
  as the bias voltage is being increased.
  This also explains that in p-bulk the bias voltage helps to isolate the n+ implants.
Understanding the underlying physics is only possible with TCAD simulation, eventually …

## Slide 59: Summary
Brief overview of the current ATLAS silicon microstrip tracker (SCT)
  ATLAS SCT strip detector is working well (so far).
Issues and achievement of the LHC tracker
  Radiation level, 2×1014 1-MeV neq/cm2
  Radiation damage effects were identified
  High voltage operation was designed up to 500 V
  Strip edge hardening against high electric field was applied
New technological achievement for the high-luminosity LHC tracker
  Radiation level, ~1×1015 1-MeV neq/cm2
  High voltage operation up to 1000 V
  Minimum dead area in the edge is evaluated
TCAD simulation
  Simulating radiation damage effects with approximation
  Very effective in understanding the underlying physics
This is all about of the conventional planar silicon microstrip sensor...
  You still have a lot of challenges ahead in different world/requirements.

## Slide 60: Appendix

## Slide 61: Signal from n+ or p+ strips
Signal ~ half of the carriers.
High field around the strips
Weighting field, mobility
n-bulk:
Low field toward n+ strip

## Slide 62: Punch-Through Protection (PTP) Structure
“Full gate” induced PTP onset in lower voltages than “No gate”.
Onset voltage went down first and then started to increase.
  What causes the transitions?
BZ4D-3
(No gate)
BZ4D-5
(Full gate)
Flunece dependence of PTP onset voltage
0
Non-irrad.

## Slide 63: PTP Simulations
TCAD
  no bias resister in parallel
  NPTP:“No gate”
  Others: “Full gate”
Parameters:
  NB/DB: non/damaged bulk
  LT/HT: lo/hi interface charge
  LC/HC: lo/hi current
  Non irrad: NB*LT*LC
  Irrad: DB*HT*HC
Irrad. simulation
  Damaged bulk,
  hi interface charge,
  hi leakage current
TCAD simulation of
“Full gate” PTP, irradiated
Electric field at onset
  when the backplane bias voltage at -200 V
  Vtest (left implant) at -50 V

## Slide 64: PTP Simulations
Onset voltage decreased as
  No gate (black) → Full gate (colored)
  Interface charge increased
Increased as
  accepter-like state increased
The fluence dependence can be understood as the effect of
  Build-up of the Interface charge and
  Increase of acceptor-like levels.
The systematic “offset”
  difference between the 2D simulation and the 3D real.
No gate
→Full gate
Oxide charge
increase
Leakage current
increase
Doping conc.
increase
No gate
→Full gate
Oxide charge
increase
Doping conc.
increase