## Slide 1
THE BASIS of VACUUM
G.Vandoni, CERN
Physics of gases
Flow regimes
Definitions
The pumpdown process
Conductance
Conductance calculation

## Slide 2: Physics of gases

## Slide 3: Units and conversions
Flow or Throughput
To From | Pa m3/s | mbar l/s | torr l/s
Pa m3/s | 1 | 10 | 7.5
mbar l/s | 0.1 | 1 | 0.75
torr l/s | 0.133 | 1.33 | 1
To From | l /s | cm3/s | m3/h
l /s | 1 | 1000 | 3.60
cm3/s | 0.001 | 1 | 0.0036
m3/h | 0.278 | 278 | 1
Pumping speed Volumetric flow
Conductance
To From | Pa | mbar | torr
Pa | 1 | 0.01 | 7.5/1000
mbar | 100 | 1 | 0.75
torr | 133 | 1.33 | 1
Pressure
[Volume/time]
[Pressure ×Volume/time]

## Slide 4: The physics of Gases
p = pressure
V = volume
nm = amount of gas (number of moles)
T = temperature
R = general gas constant [8,314 J/(mol K)]
The ideal gas law
An ideal gas is composed of randomly-moving, non-interacting point particles.
1 dozen =
12 units
6.022x1023 units
n = amount of gas (number of atoms or molecules)
kB = Boltzmann constant = R/6.022x1023
1 mole =
Notes: The state of an amount of gas I determined by its pressure, volume and temperature.
kB: relates temperature and energy
Approximation

## Slide 5: The physics of Gases
Useful forms of the ideal gas law
In a closed volume, increasing temperature from T1 to T2 , pressure increases proportionally from p1 to p2
At constant temperature, the same number of molecules distribute in 2 volumes V1 and V2 at pressures p1 and p2 such that:
V1
V2
Notes: The state of an amount of gas I determined by its pressure, volume and temperature.
kB: relates temperature and energy
Approximation

## Slide 6: Gas Mixtures and Partial Pressures
Dalton law
The total pressure exerted by a mixture of (non-reactive) gases is equal to the sum of the partial pressures of individual gases
DEFINITION
Partial pressure is the pressure which a gas would exert if it occupied the volume of the mixture on its own
n3
n2
n1

## Slide 7: Flow regimes

## Slide 8: Mean free path  and density
l
molecular chaos
trajectory of a molecule
Vacuum range | Pressure [mbar] | n [molecules/cm3] | Mean free path
Ambient pressure | 1013 | 2.5 × 1019 | 68 nm
Low vacuum | 300 – 1 | 1019 – 1016 | 0.1 – 100 μm
Medium vacuum | 1 – 10−3 | 1016 – 1013 | 0.1 – 100 mm
High vacuum | 10−3 – 10−7 | 1013 – 109 | 10 cm – 1 km
Ultra high vacuum | 10−7 – 10−12 | 109 – 104 | 1 km – 105 km
Extremely high vacuum | <10-12 | <104 | >105 km
The average distance traveled by a molecule between collisions with other molecules
At 296K=23C
rarefaction

## Slide 9: Mean Free Path 
In the hard sphere molecules approximation
(d = diameter)
Ideal gas law
Gas |  p | Gas |  p
H2 | 12x10-3 | CO2 | 4 x10-3
N2 | 6.4x10-3 | Ar | 7x10-3
He | 19x10-3 | Ne | 14x10-3
CO | 7x10-3 | Kr | 5x10-3
mean free path × pressure, room temperature, in [m×Pa] = [cm×mbar]
The average distance traveled by a molecule between collisions with other molecules
n = density of molecules
σ = collisional cross section
For N2:
l inversely proportional to p
Notes: Lambda not depending on the temperature and therefore not on the velocity of the molecules, just by the spatial distribution of targets and their size

## Slide 10: Mean free path and flow regime
The flow dynamics is characterized by the comparison of the mean free path  to the dimension D of the vacuum vessel.
Knudsen number
Free Molecular Flow | l >D
Transitional (or intermediate) flow | D/100< l< D
Viscous (continuum) flow | l< D/100
Applying the previous slide, we have a useful relation between pressure and dimension of the vessel to distinguish flow regimes:
Molecular flow: pD<0.064 [mbar . mm]
Viscous flow: pD>6.4 [mbar . mm]
Room temperature, for N2
l
molecular chaos
trajectory of a molecule
Kn >1
0.01< Kn< 1
Kn< 0.01
trajectory in a rarefied gas, l>D
Collisions with wall (much) more frequent than with molecules
Collisions with molecules dominate
Notes: Chambers p.42
- Kn number: criterion for distinguish between the regimes
- Explain each regime
- No sharp transition between the regimes

## Slide 11: Definitions Throughput Pumping speed

## Slide 12: Gas Flow: throughput …
French: débit
When flow doesn’t change with time (we say steady state), Q has the same value at every position along the pipework: mass is conserved.
Throughput=gas flow rate
Quantity of gas d(pV) crossing a plane along a duct in unit time dt.
From the ideal gas law written per unit time, we see that this is a energy flow rate.
Pam3s-1= Nms-1 = Js-1= W
particle flow rate
The same flow subsists in different locations of a continuous unbranched isothermal pipework
packet of energy
count money=count dwarfs
This is equivalent to a particle flow rate only at constant temperature.
Notes: 10: Qin=Qout meaning everything which arrives at the entrance of a pipe get to the other extremity. We will see a generalization of this later on, discussing the continuity equation.

## Slide 13: …and pumping speed
We usually call
the volumetric flow rate (débit volumétrique)
in particular, at the entrance of a pump, we call it pumping speed, S.
Substituting this in the definition of throughput, we obtain:
“The quantity of gas flowing is the product of pressure and the volumetric flow rate at that pressure”
Pumping speed is usually depending on pressure, so S=S(p)
Pumping speed is different for different gas species

## Slide 14: The pumpdown process

## Slide 15: The pumping process
V
p
QL
QG
Qv
S
S*
S* pumping speed of the pump
S pumping speed at the vessel, of volume V
p pressure in the vessel
Qtot = sum of all gas loads entering the vessel: outgassing, leaks and permeation, process gas…
BASIC EQUATION OF PUMPING – or the Continuity Equation
Qp
pump
vessel
Change of amount of free gas in the volume V
Gas pumped away in a time dt at pressure p and speed S
Amount of gas getting into the free volume in a time dt at pressure p and speed S
QTotdt
Spdt
V.dp

## Slide 16: The pumping process
V
p
QL
QG
Qv
S
S*
S* pumping speed of the pump
S pumping speed at the vessel, of volume V
p pressure in the vessel
Qtot = sum of all gas loads entering the vessel: outgassing, leaks and permeation, process gas…
The difference between the quantity of gas entering the volume and the one leaving it in a small interval of time dt is equal to the net change in the quantity of gas in the volume V, d(pV)=V·dp
BASIC EQUATION OF PUMPING – or the Continuity Equation
For conserved quantities:
«Everything which enters a volume minus everything which leaves it, equals the net increase in the quantity in the volume»
NB: T=constant, i.e. isothermal conditions, or pV is not the conserved quantity
QTotdt
Spdt
V.dp
Qp
To go deeper:
pump
vessel

## Slide 17: The continuity equation
Usually, the quantity of gas in a pumped volume decreases with time, so we can turn the equation to better express this reduction:
The rate of change of the amount of gas in a chamber is the difference between the rate of its removal and the influx rate.
What does dp mean? Or dt?
dp is the change in pressure, not the absolute, measured pressure.
dt is a small time interval in which the change occurs.
Notice that it is not always easy to apply the continuity equation to know how pressure reduces with time:
QTot is time dependent, and also depends on the previous history of the system. S is dependent on pressure and also on gas species…

## Slide 18: Example: continuity equation
After 3h pumping, pressure is presently 5.10-7mbar. The pumping speed of the turbomolecular pump, including reduction by the conductance connecting it to the vessel, is 10ls-1.
The vessel is a tube, 1m long and 400mm in diameter. Chronometer in the hand, you notice that 40s later pressure has decreased to 4.10-7mbar.
What is the rate of change of pressure in this moment?
In absence of leaks, can you evaluate the total outgassing rate of the vacuum chamber in this moment?
p=5·10-7 mbar
dt =40s, dp =1·10-7 mbar dp/dt=2.5·10-9mbar/s
Seff =10 l/s
d =400mm, L=1000mm
Qoutg=(5·10-6 +3·10-7 ) mbar ·l·s-1 = 5.3 ·10-7 mbar ·l·s-1

## Slide 19: Pumpdown: initial phase
Initially, the pumpdown process is dominated by evacuation of the free gas in the volume. Let’s write Qtot=0 and let’s call pinitial(t) the pressure decrease curve in the initial phase of pumpdown.
Gas quantity present in the volume (pV) decreases while gas is evacuated by the pump.
Volume depletion
Qpumped
3
Pressure
time
103
102
101
1
2
4
5
6
7
Remember maths! A function which changes with a rate proportional to the function itself is an exponential…
Characteristic time or time constant
po
ln p versus t

## Slide 20: Initial pumpdown time
To make time appear alone, let’s rearrange by taking the natural logarithm on both sides:
We obtain the time to lower the pressure from the initial value p0 to some value p
1000
100
10
1
0.1
0.01
t
p
Example:
A 50l volume is pumped down with
S=1 l.s-1, starting from 1000mbar to 1mbar.
What is the value of the time constant t?
How much time does it take per decade pressure lost?
How much time in total?
t=V/S=50s
Time per decade=t.ln(10)=115s
Total time=3.115s=345s
2.3t
What happens here?

## Slide 21: Pumpdown: effect of outgassing on p(t)
Below 1 Pa=10-2mbar, (~roughing), the curve p(t) starts to deviate from the “free volume” straight line. We cannot neglect outgassing Qout anymore.
n ~1 for metallic unbaked surfaces
~0.5 for elastomers, (for baked metallic surfaces)
with 0.5n1.2
ln p versus t
Volume depletion
3
Pressure
time
10-2
10-3
10-4
1
2
4
5
6
7

## Slide 22: Pumpdown: effect of outgassing on p(t)
Volume depletion
n ~1 for metallic unbaked surfaces
~0.5 for elastomers
with 0.5n1.2
Pressure
time
10-2
10-3
10-4
1
10
100
1000
ln(p) versus ln(t)
Below 1 Pa=10-2mbar, (~roughing), the curve p(t) starts to deviate from the “free volume” straight line. We cannot neglect outgassing Qout anymore.

## Slide 23: Pumpdown: effect of outgassing on p(t)
Volume depletion
n ~1 for metallic unbaked surfaces
~0.5 for elastomers
with 0.5n1.2
3
Pressure
time
10-2
10-3
10-4
1
2
4
5
6
7
ln p versus t
Eventually, the pressure flattens down and becomes “constant” on the time scale of observation. On a long time scale, it continues to decrease slowly.
Below 1 Pa=10-2mbar, (~roughing), the curve p(t) starts to deviate from the “free volume” straight line. We cannot neglect outgassing Qout anymore.

## Slide 24: Pumpdown: when outgassing dominates
For unbaked metals of standard rugosity,
We write pultimate because this pressure won’t decrease on the time scale of observation (ex. 1h)
Actually, it decreases, because Qout decreases, but this process is much slower.
The walls of the vessel get progressively emptied from their initial gas contents and gas release to the free volume decreases.
When pressure ceases to fall and becomes constant on the time scale of observation, dp/dt=0.
The equation becomes:
p
t

## Slide 25: Example: pumpdown with outgassing
From the pressure and applied pumping speed, you can evaluate the outgassing rate and see if it is conform to expectation
Example:
A stainless steel vacuum vessel of 1m length and 250mm diameter is evacuated by one pumping group of an effective pumping speed 50l/s. After 5h pumping, pressure has stabilized at ~1·10-6mbar.
Do you consider that everything is running fine?
p·S=5·10-7mbar·l/s
A=7850 cm2
t=5hr=18,000s j=3·10-9mbar·l/cm2

## Slide 26: Conductance

## Slide 27: Conductance
pU
pD
Q
C
This equation relates throughput (fr. débit) to the difference between upstream and downstream pressure.
It is the DEFINITION of conductance.
The quantity of gas which is flowing across a given pressure difference depends on the ease of flow , described by CONDUCTANCE.
Its reciprocal 1/C is a resistance to flow; i.e., the opposition the system exerts to gas flow.
Useful (but approximate) analogy to electrical circuit
Driving force: voltage drop or pressure difference
Flowing quantity: electrical charge or molecules
Flow: electrical current or throughput
Notes: Across the same pressure difference we have a larger flow of gas when the term C is large. C represents the opposition to flow by the system.

## Slide 28: Combining conductances
pU
pD
Q
pU
pD
C2
C1
In parallel
In series
pU
pD
Q
pU
pD
Q
C1
C2
Approximate: it overestimate resistance, by neglecting beaming and entrance resistance cancelling
Derivation:
Conservation of flux
Definition of C

## Slide 29: Reduction of pumping speed by connecting pipe
Application of the combination of conductances in series.
S* effective pumping speed at the chamber
S real pumping speed
p* pressure in the chamber
p pressure at the entrance of the pump
C connecting conductance
Q throughput through the pipe and into the pump
or
S* is less than S!
Lets plot S*/S reduction of pumping speed, against C/S, ratio between conductance and pumping speed
Ex: If the conductance is equal to the pump’s speed, we only get 1/2 of it at the vessel. To get 80% pumping speed at the vessel, we need a conductance 4x larger than the pumping speed.
Chart: C/S
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
S*/S | 0 | 0.5 | 0.66667 | 0.75 | 0.8 | 0.83333 | 0.85714 | 0.875 | 0.88889 | 0.9 | 0.90909 | 0.91667 | 0.92308 | 0.92857 | 0.93333 | 0.9375 | 0.94118 | 0.94444 | 0.94737 | 0.95 | 0.95238
S
S*
C
p
p*
C=S
80%
C=4S
Notes: What is in ordinate, what is in abscissa. We have C/S, the comparison between C and S. So for example (click), this is a conductance equal to the pumping speed, say 100l/s atached to a 100l/s pump.

## Slide 30: Pressure profile in long tube with localized pumps
It is not difficult to show also that
Q: outgassing flux [mbar.l.s-1]
S: pumping speed [l/s]
C: tube conductance [l/s]
PRESSURE PROFILE
A tube under vacuum will outgas uniformly over its whole internal surface. Pumps are installed along the tube. Due to the limited (i.e., not infinite) conductance of the tube, pressure will have maxima at equal distance from two pumps and minima above the pumps.
From the continuity equation, it is obvious that:
Notice differences in pressure measurement in long tubes, depending on their distance to the pump!

## Slide 31: Conductance calculation Conductance in viscous flow Conductance in molecular flow

## Slide 32: Conductances and flow regimes
Let’s anticipate, that the conductances of pipes will differ in different flow regimes:
In continuum flow, they are proportional to mean pressure p.
In molecular flow they are not a function of pressure.
In Knudsen flow, a transition between the two types of flow, conductances vary with Knudsen number.
At the same diameter, conductance in continuum regime is much larger than in molecular regime
10-2
[mbar]
10-3
10-1
1
1
10-4
10
Above pD>6.4 mbar.mm, flow is viscous
Below pD<0.064 mbar.mm, flow is molecular
106
On the figure: the equal conductivity curves on a p,d graph
- read for example: 106 l/s and 1 l/s
viscous
molecular
100

## Slide 33
Reynolds number Re compares the two effects
Continuum (viscous) flow through pipes
Viscous gas flow may be:
 = density of the fluid
 = viscosity
u = flow velocity
Dh = hydraulic diameter
LAMINAR:
gas flows smoothly in stream lines, parallel to the duct walls.
TURBULENT :
gas flows chaotically and irregularly brakes up into vortexes.
Viscous forces are stabilizing
Inertial forces are destabilizing
Inertia forces compared to viscous forces
Re<2000
Re>2000
A = cross sectional area
B = wetted perimeter
In most cases treated in vacuum technology, we are in laminar conditions
Notes: Lafferty section 2.3

## Slide 34: Viscous laminar flow - Poiseuille
Circular cross section, long tube
Diameter D
Between pressures p1 and p2
With h viscosity of the fluid
This is valid only for long tubes, i.e., rarely in vacuum… But it helps understanding this matter.
NB:
In viscous flow, conductance is proportional to pressure! Already seen 2 slides ago
The effect of diameter D on C is with D4, so doubling D means a factor 16 in conductance
is the average pressure between entrance and exit pressure
Notes: Lafferty pages 110-111

## Slide 35: Viscous regime: Short tubes conductance
The difficulty here is that the velocity profile changes at the aperture, with a minimum cross section at the aperture, then with oscillations of contraction/expansion before reaching fully developed flow.
Nomograms may be used.
Particular case of practical use in vacuum:
choked (or sonic) flow.

## Slide 36: Choked flow
An effect of compressibility of a gas…
pU
pD
In choked flow, throughput is not depending on downstream pressure pD any longer, but only on upstream pressure pU.
This happens when flow velocity in the restriction is equal to sound velocity. For air, it occurs when pD<0.53 pU.
Even if pressure pD is lowered, throughput remains constant.
Q/Qc
pD/pU
1
1
0
0.5
Keep pU constant and lower progressively pD… At some value of pD/pU, you attain choked (or blocked) flow.
Remember…
French: flux sonique
A choked aperture has constant pumping speed.
Perfect flow control! Controlled soft venting! …and more

## Slide 37
Molecular Flow: Introduction
Randomizing effect of molecule wall collisions: there are no favored directions, the probability of emerging in any direction is the same, not related to incident direction
Kn>1: the mean free path of the particles is large compared with the size of the container
The molecules-wall collisions dominate the gas behavior
In molecular flow, a molecule is not “sucked” away, it falls on the pump entrance as the result of random motion.
Important to learn about the rate of molecular impingements on a elementary surface
Notes: Karl Jousten Handbook p.130 figure, lafferty p.85 figure

## Slide 38
Impact Rate J: Number of molecular impacts on a surface per unit time and unit surface
Impact rate
n = density of molecules
The molecules impacting on a surface dS are those having velocity vectors comprised between v and v+dv and direction q. Integrated over the solid angle:
average or mean speed in Maxwell-Boltzmann distribution
Notes: m= molecular mass
Karl Jousten Handbook p.130 figure, lafferty p.85 figure

## Slide 39: Conductance of an aperture
Aperture of area A in a thin wall separating two regions of pressures p1 and p2
A
J1
J2
p1, n1
p2, n2
For air at 20C
Ideal gas law: definition of flow
Net flux= difference between currents in either direction
Definition of conductance
Notes: Chambers p.97, Lafferty p.86

## Slide 40: Conductance of standard flanges
DN | CA Conductance (l/s)
16 | 24
25 | 58
40 | 148
63 | 368
100 | 927
150 | 2085
200 | 3707
Conductance of (infinitely thin) flange of standard opening
How to transform from N2 to other gases?
MEMO
Small mass ↔ large conductance
NB
Usually, we put a ~500l/s pump on a DN150.
The flange’s conductance is >4x the pump’s pumping speed.
Conductance of standard flanges for N2 at 296K
Notes: Karl Jousten Handbook p.133

## Slide 41: Maximal pumping speed
From the preceding, we can see that the maximal pumping speed of a high vacuum pump cannot be larger than the conductance of its aperture.
The maximum throughput a pump can feature is when all molecules crossing its aperture don’t return into the vacuum vessel.
maximal pumping speed for nitrogen, for a pump of aperture diameter D (in cm).

## Slide 42: Maximal pumping speed
maximal pumping speed for nitrogen, for a pump of aperture diameter D (in cm).
High vacuum pumps do not “suck” molecules into them. They rather prevent (a certain number of) molecules having crossed by chance their aperture to travel back to the vacuum vessel.
If the pumping mechanism were 100% efficient, the pumping sped would be equal to the conductance of the entrance.
Pumping speed is highest when backstreaming into the vessel is lowest.
From the preceding, we can see that the maximal pumping speed of a high vacuum pump cannot be larger than the conductance of its aperture.
The maximum throughput a pump can feature is when all molecules crossing its aperture don’t return into the vacuum vessel.

## Slide 43: Conductance of a long tube

## Slide 44: Conductance of a short tube
Very roughly, we can consider that the conductance of a short tube is the conductance of its aperture CA
followed by the conductance of a long tube, CL .
From the combination of conductances formula:
to be used for “back of the envelope” estimations
Not for dimensioning a system, may be wrong up to 12%
Valid with some error, but rather good in the limits L=0 and L>>D, for nitrogen, at 295K=20C.
Elaborated by Dushman, 1922
Aperture
Long tube

## Slide 45: Transmission probability α
If N molecules arrive at the entrance plane of a duct, a number Nα reach the exit plane, with a<1. A number N(1-α) return to the entrance. Remember that N=JA with A area and J impingement rate of molecules at the aperture of the duct.
J1A
J2A
Net flux from entrance to exit
Of the J1A molecules entering the duct, a· J1A reach its exit
Of the J2A molecules entering the duct from downstream, a J2A reach its entrance plane
Net throughput from entrance to exit
Substituting the expression
we make p appear.
The conductance of a component is the conductance of its aperture times a transmission probability

## Slide 46: Transmission probability calculations
The concept of transmission probability was introduced by Clausing (1932). The values he has calculated by integrals (sometimes called Clausing factor) are still widely used for circular tubes and are accurate to ~ percent. Nomograms often use his values.
Cole (1976) refined the same methods, his values are very accurate.
Dushman (1992) has proposed an analytical formula (next slide) easy to apply, but not completely accurate. We use it for quick estimations.
Livesey (1998-2004) collected all transmission probability data, analyzed them and obtained analytical formulas (2004), very useful for computer calculation. For exotic forms, we have proposed his values, which are fairly simple and accurate to <1% over a wide L/D range.
One of the best ways to calculate transmission probabilities is the statistical Monte-Carlo method, which can be applied with success to complicated shapes which escape from a simple geometrical treatment.

## Slide 47: Dushman transmission probability for cylindrical tubes
to be used for “back of the envelope” estimations
Transmission probability of a cylindrical tube, of diameter D and length L, or of form factor L/D.
Dushman 1992
Aperture
Long tube

## Slide 48: Cylindrical ducts
Chart: l/d
| 0.05 | 0.15000000000000024 | 0.25 | 0.35000000000000031 | 0.45 | 0.5 | 0.60000000000000064 | 0.70000000000000062 | 0.8 | 0.9 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 500
 | 0.9524 | 0.86993 | 0.80127 | 0.74341 | 0.69404 | 0.67198 | 0.63223 | 0.59736 | 0.56651 | 0.53897 | 0.51423 | 0.42006 | 0.35657 | 0.31053 | 0.27544 | 0.24774 | 0.22526 | 0.20664 | 0.19094 | 0.1093 | 0.076912 | 0.059422 | 0.048448 | 0.040913 | 0.035415 | 0.031225 | 0.027925 | 0.025258 | 0.002646
Chart: l/d
| 0.05 | 0.15000000000000024 | 0.25 | 0.35000000000000031 | 0.45 | 0.5 | 0.60000000000000064 | 0.70000000000000062 | 0.8 | 0.9 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 500
 | 0.9524 | 0.86993 | 0.80127 | 0.74341 | 0.69404 | 0.67198 | 0.63223 | 0.59736 | 0.56651 | 0.53897 | 0.51423 | 0.42006 | 0.35657 | 0.31053 | 0.27544 | 0.24774 | 0.22526 | 0.20664 | 0.19094 | 0.1093 | 0.076912 | 0.059422 | 0.048448 | 0.040913 | 0.035415 | 0.031225 | 0.027925 | 0.025258 | 0.002646
 | 0.9524 | 0.86993 | 0.80127 | 0.74341 | 0.69404 | 0.67198 | 0.63223 | 0.59736 | 0.56651 | 0.53897 | 0.51423 | 0.42006 | 0.35657 | 0.31053 | 0.27544 | 0.24774 | 0.22526 | 0.20664 | 0.19094 | 0.1093 | 0.076912 | 0.059422 | 0.048448 | 0.040913 | 0.035415 | 0.031225 | 0.027925 | 0.025258 | 0.002646
l/D | a [Cole]
0.05 | 0.952399
0.15 | 0.869928
0.25 | 0.801271
0.35 | 0.74341
0.45 | 0.694044
0.5 | 0.671984
0.6 | 0.632228
0.7 | 0.597364
0.8 | 0.566507
0.9 | 0.538975
1 | 0.514231
1.5 | 0.420055
2 | 0.356572
2.5 | 0.310525
3 | 0.275438
3.5 | 0.247735
4 | 0.225263
4.5 | 0.206641
5 | 0.190941
10 | 0.109304
15 | 0.076912
20 | 0.059422
25 | 0.048448
30 | 0.040913
35 | 0.035415
40 | 0.031225
45 | 0.027925
50 | 0.025258
500 | 0.002646
D
l

## Slide 49: Pumping speed measurement
S
Introduce a known throughput qpV above the pump and measure pressure p at the inlet
Pressure when
Gas flow qpV may be measured by a precise flowmeter or in a double-dome with known conductance between the domes
Measure pressure p in front of the pump is problematic: directional gas flow!
Pressure is defined in an enclosed system, at equilibrium.
qpV
p
To avoid perturbation of the isotropic (i.e. uniform) Maxwellian distribution by the pump inlet, one would need an infinitely large volume.
Solution: Fisher-Mommsen dome (double and “isotropic”)

## Slide 50: Fisher-Mommsen dome
E. Fischer, H. Mommsen, Vacuum 17, 309 (1967)
The Fisher Mommsen dome has geometry, dimensions and position of the gauge such that measured pressures are identical to those measurable in the ideal case.
Thermodynamically, p is defined only in an enclosed system at equilibrium.
To approach isotropy, we need a “infinite volume” dome, and gauges far from the pump. The gauge should then measure a isotropic, Maxwellian gas distribution.
The inlet C and pump S should have a negligible effect on the distribution.
for pumping speed measurement

## Slide 51: Examples
From the table of transmission probabilities, molecular flow transmission probability for a pipe whose length is equal to its diameter is 0.51. This means that only about ½ of the molecules that enter it, pass through. What fraction will get through for a pipe with L/D=5?
a=0.19
The molecular flow transmission probability of a component with entrance area 4cm2 is 0.36. Calculate its conductance for nitrogen at (a) 295K, (b) 600K.
From CA=11.8A l/s (with A in cm2) for nitrogen at 295K, and C=aCA, we get C=17l/s.
The effect of temperature is with √T, conductance being proportional to √T. We must divide by √295 and multiply by √600, to obtain C=24 l/s.

## Slide 52: Examples
A component has a molecular flow conductance of 500l.s-1 for nitrogen. What will its conductance be for (a)hydrogen, (b)carbon dioxide?
We have to multiply by the square root of the molar mass of nitrogen (N2, M=28) and divide by the square root of the molar mass of hydrogen (H2, M=2) or carbon dioxide (CO2, M=44).
In the first case, this gives a factor 3.74, in the second, a factor 0.8.
So the conductance for hydrogen will be 3.74 times larger, the one for carbon dioxide 0.8 times larger. We get 1870l.s-1 for hydrogen and 399l.s-1 for carbon dioxide.
Notice that the square root factor “damps “ the effect of mass, but nevertheless this factor is ~4 for hydrogen if compared to air….
By what factor will the molecular flow conductance of a long pipe be increased, if its diameter is doubled?
With the simplified formula of Dushman for long tubes in molecular flow, we see that the effect of D is with D3.
So doubling D will increase conductance by a factor 8 !!

## Slide 53: Examples
A vessel of volume 3m3 has to be evacuated from 1000 mbar to 1 mbar in 20 min. What pumping speed (in m3 per hour) is required?
How long will it take for a vessel of volume 80 l connected to a pump of speed 5 l/s to be pumped from 1000 mbar to 10 mbar? What is the time per decade?
We apply the formula for the rough vacuum regime (which neglects outgassing), to obtain: t=6.9V/S. So S=6.9V/t=62m3/h
Notice that 3m3 is the typical volume of a Booster sector, of a SPS arc sector. With half this pumping speed, time is doubled...
HINT:
As above: ln(po/p)=4.6, so t=4.6V/S=4.6.16≈80s
Time per decade: 5.2.3s=12s

## Slide 54: Example: pressure profile
In a SPS transfer line, the diameter of the pipe is ~60mm, the distance between pumps ~60m. The effective pumping speed of each ionic pump is ~15 l/s, including a correction for the connecting element. Let’s assume an outgassing rate of 3.10-11mbar.l.s-1.cm-2, typical after 100h pumping time.
Calculate the maximum and minimum pressure measured at a pressure gauge in a segment between 2 pumps.
Formulas to use:

## Slide 55: Example: pressure profile