## Slide 2: PSB and PS DC/Intermediate Current Transformers
P.Odier LIU BI review 3.10.2013
P.Odier, S.Thoulet, R.Ruffieux, J.Longo, L.Jensen, S.Bart Pedersen, J.Tan

## Slide 3: PSB Ring DCCT specifications
Devices: 4 monitors, BRi.TDC.9L1
Derived parameter: total beam intensity evolution in the ring over the cycle
  Intensity range: 5E9 to 2.5E13 charges
  Acquisition sampling: 1 ms
  Accuracy: 1% of measured value + offset/noise
  Resolution: 5E8 charges

## Slide 4: PSB DCCT status and plans
LS1:
  New electronics derived from LHC DCCT (front and back end) ready to be installed by Jan.2014 in replacement of the original electronics installed in 1993
  New 16 bit ADC (replace 12 bit) to be installed by Jan. 2014
After LS1:
  Increase of the full range for the L4 beam
  Modification of the ß normalisation (2 GeV + new “B train”

## Slide 5: PSB Ring intermediate Transformer specifications
Devices: 4 monitors, BRi.TMD.8L1
Common housing with TFA and TSW
Derived parameter: injection efficiency and SW watchdog
  Analog turn by turn acquisition up to 100 turns
  Intensity range
    Min : 5 E9 charges
    Max : 2.5 E13 charges
  Resolution: Detect 1% loss from turn to turn
  Noise~ 109
  Accuracy : 2%
  Watchdog: comparison with BIi.BCT20 after 100 turns

## Slide 6: PSB Ring Intermediate Transformer status and plans
Status:
    Devices present in the PSB since 1971, windings, electronics and cables are to be modified/replaced
    1 partial spare available for a test bench
    4 replacement chambers fabricated (without ceramic gap), requested by TE/VSC
  Plans after LS1:
    Fabrication of spare vacuum chambers (with ceramic gap) since there is currently no spare!
    Pulling of new cables (20 coax + 4 multi-conductors; DIC already sent on May 2012)
    New windings. 3 months required for dismantling, refurbishment in the BI radioactive workshop and reinstallation
    New front and back end electronics to be built
    New acquisition chain (TRIC card) to be installed

## Slide 7: PS Ring DCCT specifications
Device: 1 monitor, PS.TR34
Derived parameter: total beam intensity evolution in the ring over the cycle
  Intensity range: 2E9 to 4E13 charges
  Acquisition sampling: 1ms
  Accuracy: 1% of measured value + offset/noise
  Resolution: 5E8 charges
  confirmation needed

## Slide 8: PS Ring DCCT status and plans
LS1:
  New electronics derived from LHC DCCT (front and back end) ready to be installed in replacement of the original electronics installed in 1992
  16 bit ADC (replace 12 bit) to be installed by Jan.2014
  New Base Line Restoration based on VME modules (higher sampling rate, more base lines memorized, selection by SW)
  For info: monitor temporary dismantling to permit the renovation of bending magnet #34 (Oct.-Nov.2013)
After LS1:
  Increase of the full range for the L4 beam
  Modification of the ß normalisation for the new system of B value distribution (White Rabbit). Waiting for specifications

## Slide 9: Budgetary Requirements
PS DCCT is Funded (CONS 64275)
PSB DCCT/intermediate CT are funded (LIU 64021)
  Spending Profile [kCHF]:
 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | total
PSB DCCT | 7 | 11 | 10 | 10 | 10 | 10 | 0 | 0 | 58
PSB inter. CT | 7 | 10 | 30 | 40 | 50 | 54 | 40 | 20 | 251

## Slide 10: Conclusions
A new electronics, derived from the LHC DCCT, will be installed before end of LS1. Acquisition improved.
We need confirmation of the maximum intensity value in the PS.
The PSB intermediate Current Transformer project is postponed due to cabling issue.
The upgrade of PSB and PS DC/intermediate Current Transformers is funded.
The planning and therefore the spending profile are floating waiting for LS 1.5/2 decision.
We are still waiting for the specification of the new system to replace the B train distribution.