## Slide 1
High-gradient accelerating structures for proton acceleration
Linear collider development applied to cancer therapy

## Slide 2
CLIC accelerating structure development
In the context of the study of a TeV-range linear colliders – NLC/JLC and CLIC - we have developed the science and technology for achieving high gradients, 100 MV/m, using short pulse, normal conducting rf.
We are now actively seeking to expand the technological and industrial base for CLIC by promoting the use of high-gradient technology in other applications such as medical linacs, FELs, Compton sources etc.

## Slide 3
Accelerating gradients achieved in tests.
Status: 4-9-2012
HOM damped

## Slide 4
How we got there: High-power design laws
The functions which, along with surface electric and magnetic field (pulsed surface heating), give the high-gradient performance of the structures are:
global power flow
local complex power flow
New local field quantity describing the high gradient limit of accelerating structures.
A. Grudiev, S. Calatroni, W. Wuensch (CERN). 2009. 9 pp.
Published in Phys.Rev.ST Accel.Beams 12 (2009) 102001
Hs/Ea
Es/Ea
Sc/Ea2

## Slide 5
TD18#3 at SLAC
TD18#2 at KEK
Stacking disks
Temperature treatment for high-gradient developed by NLC/JLC
How we got there: heat treatment and material structure
Attempting to understand why it works.

## Slide 6
Micron precision turning and milling
Accelerating structure tolerances drive transverse wakefields and off-axis rf induced kicks which in turn leads to emittance growth – micron tolerances required.
Multi-bunch trains require higher-order-mode wakefield suppression – cells require milled features.
High-speed diamond machining also seems to be beneficial for high-gradient performance through minimizing induced surface stresses.
Development done “in industry”

## Slide 7
Evolution of machining capability
Up to the 1980’s
1980’s - 1990’s
2000’s - 2010’s
Larger machines
Multiple axis ( X/Y/Z and C)
Future ?
Intelligent machines ?
  Robotisation ?
Pallet machining?
  Robotisation ?
First machines at research institutes and universities
Start of industrialization
Optical recording
contact lenses
Single point diamond turning
Up to the 1990’s
1990’s - 2000’s
2010’s
Future ?
Ultra precision diamond milling (lagging more than a decade behind on turning)
Limited to fly cutting
mirror optics
Laser scanner mirrors
First proto type machines
Micro fluidics
Accelerator parts
Milling as add-on on lathes
Lens arrays
Intra ocular lenses

## Slide 8
Klystron-based test stands for CLIC
NEXTEF at KEK
Xbox-1 at CERN
SLAC XL-5 50 MW
Xbox-2 at CERN based on CPI VKX-8311A 50 MW tube
Xbox-3 at CERN based on Toshiba 6 MW tube

## Slide 9: TERA program: development and construction of « cyclinacs »
pulsed source
CYCLOTRON
100-300 Hz
4 microseconds
many indipendent 3 GHz RF power sources
9-12 Units
to the patient
HIGH REP RATE PULSED SPOT WITH FAST INTENSITY AND ENERGY MODULATION
tumour target

## Slide 10: TULIP at 3 GHz and high gradients (E0 = 30 MV/m)
TR24

## Slide 11: TULIP at 3 GHz and high gradients (E0 = 30 MV/m)
17 m
24 m
TR24
400 m2
24 MeV cyclotron
by
Advanced Cyclotron Systems
(Canada)
1.7 m
TR24

## Slide 12: TULIP at 3 GHz and high gradients (E0 = 30 MV/m)
eleven
10 MW klystrons
patient access
TR24
24 MeV
≤230 MeV
Rotation: +-110° wrt the horizontal plane

## Slide 13: TULIP at 3 GHz and high gradients (E0 = 30 MV/m)
patient access

## Slide 14: TULIP at 3 GHz and high gradients (E0 = 30 MV/m)
scanned area:
20x20 cm2
RF rotatory joints

## Slide 15
This project
The objective of this project is to design and build two accelerating structures for a compact proton accelerator as one might incorporate in TULIP, using the high-gradient technology developed for CLIC.
The proton energies we will design for are 76 and 213 MeV.
The primary challenge comes from the low velocity of the protons, which leads to a very different design rf structure compared to that for relativistic electrons.
We hope to be able to raise the accelerating gradient from the 27 MV/m of LIBO linac tanks up to the range of 50 MV/m.
Our sincere thanks to the KT department to making this project possible.

## Slide 16: Backward wave high-gradient accelerating structure for proton acceleration based on CLIC technology

## Slide 17: gap and angle scan – 120 deg
g4 a25
g4 a55
g5 a25
g6 a25
g6 a55
g5 a55
g4 a75
g5 a75
g6 a75

## Slide 18: Coupling slot optimization
P_in
P_out
Goal: minimize the S11 parameter
Optimization of the accelerating cell diameter for different number of cells (n=1:3), with Cl as parameter
The max S11 among the three computed is considered
Cl
acD
n

## Slide 19: Preliminary design for high gradient bTW linac
Independent rotary joints
-3 dB recirculation
Small RF load compared to TW
T2
T1
load
load
MKs
~15-16 MW
klystron
2
3
T2
T1
load
load
MKs
1
2
3
2 x 7.5 MW klystrons

## Slide 20
SUMMARY | 120 deg |  | 150 deg |  | SCL base | SCL – HG
wall thickness (mm) | 1.5 |  | 1.5 |  | 3.0 | 3.0
gap (mm) | 5.5 |  | 7.0 |  | 5.1 | 9.5
nose cone angle (deg) | 65 |  | 55 |  | 25 | 55
length (mm) | 189.9 |  | 189.9 |  | 189.9 | 189.9
ncell | 15 |  | 12 |  | 10 | 10
Ea_avg (MV/m) | 25 |  | 25 |  | 25 | 25
Sc_nose (MW/mm2) | 0.149 |  | 0.185 |  | 0.486 | 0.188
t_pulse (ns) flat | 2500 |  | 2500 |  | 2500 | 2500
expected BDR (at given Ea and t_pulse) (bpp/m) based on Sc limit | 1.1 E-22 |  | 2.9 E-21 |  | 5.7 E-15 | 3.7 E-21
max Ea (for BDR of 10-6 bpp/m) (MV/m) | 85.2 |  | 76.3 |  | 47.1 | 75.7
Pin (MW) (w/o recirculation) | 2.70 | 5.19 | 2.49 | 5.10 | 1.75 | 2.26
Pout (MW) (w/o recirculation) | - | 2.90 | - | 3.02 | - | -
Q0 (first/last) | 6482/6721 |  | 7088/7545 |  | 8291 | 8250
vg (first/last) [%c] | 0.421/0.226 |  | 0.404/0.236 |  | - | -
R’/Q (first/last) [Ohm/m] | 7872/7847 |  | 7835/7794 |  | 8406 | 6355
time constant (ns) | 320 |  | 340 |  | 440 | 440
field rise time (time to reach 99% field) (ns) (w/o recirculation) | 750 | 204 | 800 | 204 | 1050 | 1050

## Slide 21: Accelerating elements
120° of phase advance
150° of phase advance
4 holes for dimpler tuners

## Slide 22: Evaluation of different cells structural performance
Load: gravitational force
g
g
120° of phase advance
150° of phase advance

## Slide 23: 120°
Field [MV/m] | Espected max T [°C] | Power per cell [W]
100 | 102.8 | 573.9
50 | 43.7 | 143.5
25 | 28.9 | 35.9
Coherent with theoretical values