## Slide 1: Linac4 Ion Source commissioning: Vacuum Aspects
C.Pasquino, N.Thaus, J.Hansen, P.Chiggiato

## Slide 2: Outline
Vacuum aspects and lay-out;
Vacuum simulation benchmark;
Beam transmission measurements;
NEG pump activation procedure and monitoring;
Vacuum sealing: problems faced and solutions;
Main achievements;
Maintenance & resources.

## Slide 3: Vacuum lay-out
TMP
NEG
SIP
SIP
NEG
TMP
TMP
TMP
LEBT
Source
RFQ
The two major pressure constraints are:
  The peak pressure during RF pulsing in the RFQ (5*10-7 mbar);
  Operation of the NEG pumps;
Pulsed injection at 5*10-3 mbar l /pulse
Continuous injection at 1*10-5 mbar
H2 flux at 1*10-3 mbar l/s
Courtesy of D. Steyaert
This value measured during the commissioning of the source by
approximately 5*10-2 mbar l /pulse.
RFQ:
LEBT:
SOURCE:

## Slide 4: Vacuum lay-out
TMP
NEG
SIP
SIP
NEG
TMP
TMP
TMP
LEBT
Source
RFQ
The two major pressure constraints are:
  The peak pressure during RF pulsing in the RFQ (5*10-7 mbar);
  Operation of the NEG pumps;
Pulsed injection at 5*10-3 mbar l /pulse
Continuous injection at 1*10-5 mbar
H2 flux at 1*10-3 mbar l/s
Courtesy of D. Steyaert
This value measured during the commissioning of the source by
approximately 5*10-2 mbar l /pulse.
RFQ:
LEBT:
SOURCE:

## Slide 5: Vacuum Simulation benchmark: plasma chamber
N2 and H2 injections for different pulse lengths.
Hot cathode ionization gauge
Gas injection and fast pressure-rise measurements for the
Linac4 H− source
E. Mahner, P. Chiggiato, J. Lettry, S. Mattei, M. O'Neil, H. Neupert,
C. Pasquino, C. Schmitzer
H2
H2
N2
N2

## Slide 6: Vacuum Simulation benchmark: LEBT
Penning Gauge monitoring: pulsed gas injection from the source + gas density regulation on the LEBT. During this measurement the gas injection in the LEBT was about 3*10-6 mbar in H2 equivalent.
We are working outside of the vacuum limits set in the baseline design:
Plebt max = 4.5*10-5 mbar> 1*10-5mbar, leading to a maximum P in the RFQ of 1*10-6 mbar (9.6*10-7 mbar average measure) .
LEBT: measurement
LEBT: simulation
H2
H2

## Slide 7: Operation of the IS at 2 Hz?
These profiles have been simulated considering the H2 injection from the source and the gas density regulation on the LEBT.
If we increase at 2 Hz the operation of the source there will be less pressure recovery time between two subsequent pulses: the base pressure of the LEBT with no injection would be in the 3*10-6 mbar, while now is in the high 10-7 mbar range.
The actual intensity pulse at 2Hz is not compatible with the vacuum layout of the RFQ as it is designed.
LEBT | 0.8 Hz | 2 Hz
Pmin (mbar) | 3*10-6 | 6*10-6
Pmax (mbar) | 4*10-5 | 4*10-5

## Slide 8: Monitoring the pressure in the RFQ
The pressure is monitored via two cold cathode gauges, positioned on the first and third tank.
The RF is interlocked by two ion pumps located at each end of the RFQ and the cold cathode gauge positioned on the first tank: the valves are interlocked at 1*10-6 mbar, the RF is interlocked at 5*10-7 mbar.
We do not have any fast acquisition on these penning gauges: the data are logged with a time interval which is longer than 2s, which is sufficient for monitoring the vacuum pressure in the RFQ.
In case we need the fast logging, we can access the data the L4 IS team have.
On the RFQ a dedicated residual gas analyzer is mounted: it has been used during the commissioning phase of the source and the RFQ in order to monitor the partial pressures of several gaseous species. Their evolution was monitored during the RF commissioning as well as during the NEG pump activation.

## Slide 9: RGA signals during beam transmission in the RFQ: effect of beam losses.
The residual gas analyzer was monitoring the ionic current of several gaseous species during the transmission measurement.
H2 Injection ( 5*10-7 – 2*10-6 mbar), H2 leading gas;
Sparking activities of the RFQ (beam lost): CO, CO2, CH4, C4H7, C4H9.
Source and LEBT gas load

## Slide 10: NEG pumps activation: monitoring & regeneration
Considering the actual pulsed and the LEBT continuous H2 gas injection, the regeneration of the 4 NEG pumps has to take place every 27 weeks of operation. At 2 Hz or at a higher gas injection rate this value will decrease even more! We need to install a new version of the NEG pumps from SAES, specifically conceived for this application, to achieve the baseline design (1 year of operation with no regeneration). The regeneration time for the NEG pumps is 24 h during technical stops!
By isolating the turbo molecular pumps and switching on and off the sputter ion pumps it was possible to monitor the pumping speed of the NEG pumps.
When | Effective pumping speed for H2 (l/s/pump)
After activation | 800
After 7 weeks of operation | 525
When | Effective pumping speed for H2 (l/s/pump)
After activation | 
After 7 weeks of operation | 
Design | Source Gas load (mbar l/pulse) | Time before activation (months)
2Hz - Baseline | 5*10-3 | ≈ 12
0.8Hz - Actual | 5*10-2 | ≈ 7
2Hz – Future? | 5*10-2 | ≈ 5
Leak of 1*10-8 mbar l/s on a feedthrough!!
H2

## Slide 11: Source Assembly and leak tightness
Several type of sealing are used:
O-ring (elastomeric sealing) on the plasma generator ( )
Single and Double diamond aluminum sealing / aluminum helicoflex sealing on the plasma extraction region
( )
-
c
The source is equipped with metal seals except for the plasma generator.
S
S
D
Both Helicoflex and diamond seals are suited for the source:
With the actual design both types have been used;
Main leak tightness problem on the double diamond seals ( e.g non conformity of the seals);
Main drawback: if a double diamond seal is leaky, the whole assembly has to be dismounted;
The new design will only have single diamond/ helicoflex seals (widely used at CERN).
Conflat gaskets ( )

## Slide 12: Source Assembly and leak tightness
Characteristic | O-Ring | Helicoflex | D/S Diamond | Conflat Gasket
HV | Ѵ | Ѵ | Ѵ | Ѵ
UHV-XHV | - | Ѵ | Ѵ | Ѵ
Roughness | 0.8-3.2 | 0.8-3.2 | 0.8-3.2 | 0.8-3.2
Groove - Limiter | Ѵ | Ѵ | Ѵ | Ѵ
Reusable | Ѵ | - (Ѵ) | - | -
Cost | $ | $$$$ | $$$ | $$$$
Permeation | Ѵ | - | - | -
Low Outgassing | - | Ѵ | Ѵ | Ѵ
High Temperature applications (max 300◦C) | - | - (Ѵ) | - (Ѵ) | Ѵ
Precision in the assembly | - | Ѵ | - | -

## Slide 13: Maintenance and resources (1): Ion sources
Considering the existing sources and 1 or 2 prototypes per year, we will have to take care of the maintenance of up to 6 ion sources, up to the end of the R&D;
Cesiated source:
  Cs is not expected to be spurted inside the area of the vacuum pumps, but time is needed to measure and evaluate eventual problems (learning phase).
  For sure the dismantling of the cesiated source will require more time due to inspection of the inner surfaces of each component.
The maintenance will require:
  Change of the source every 3 months.
  TE-VSC will be required for the preparation, testing, assembly, leak detection, disassembly of all new prototypes and old sources.

## Slide 14: Maintenance and resources (2): Man power and Material costs
  The expected total work load will be 12 weeks per year. In addition, 2 weeks of standard maintenance of the vacuum pumps has to be foreseen (0.3 FTE/y, 1.2FTE for 2014-2017).
  The 1.2 FTE is not in the Work Package and extra man power (industrial support) will be required to fulfill this task, equal to 40KCHF/y or to a total of 160 KCHF.
The total cost to run the two source installations in building 400 and 152 is 25KCHF per year. This cost is not taking into account the maintenance and modification of the installations in building 357.
All pumping units installed on the source will be 5 years or more before Linac 4 will be connected to the PS-Complex in 2018 and a full replacement of all pumps shall be integrated in the maintenance cost for 2018.
Action | Costs (kCHF)
MAINTENANCE (2014-2017) | 25x4
PURCHASE NEW MATERIAL (2018) | ≈ 160
NEW NEG PUMPS FOR LONGER OPERATION | ≈ 50+15 (R&D)
Building | Pumping units and gauge assembly | Maintenance cost of source vacuum /year
400 | 4 + 4 | 12.5KCHF (2014 to 2017 included)
152 | 4 + 4 | 12.5KCHF(2014 to 2017 included)
Man Power + Material ≈ 485 kCHF

## Slide 15: Maintenance and resources (3): Additional tasks
Additional requests, participation in:
  Fast Gauge Test Stand:
    Identify suitable fast injection valves, already existing in the industry (if possible);
    Maintenance and intervention on the test bench in building 357;
    Simulation benchmark on the time dependent pressure profiles with the different type of valves;
  RGA in L4 tunnel:
    The RGA might be monitored locally in the control room with a dedicated PC, following signals of individual masses (1-100 amu);
    No software automatically analyzing these data will be implemented;
    This will be important for the long term optimization of the source;
    RGA has been purchased, but the installation costs of the RGA will have to be covered by the project (PC, Cabling, Pumping system).
  Magnetron source;
    It is not expected any modification in the vacuum system except on the source gas injection that is not under the TE-VSC responsibility;
    New time dependent simulations are required in order to fully evaluate the impact of the new injection on the vacuum system;
  We regret, we do not have the necessary manpower to take part in these activities, unless a GET fellow is paid by the project (210 kCHF in two years).

## Slide 16: Achievements
Simulation of time dependent pressure profiles, for the several design scenarios and for the actual design;
Benchmark of the simulations with fast pressure measurements: good correlation between the two;
Study of high efficiency hydrogen pumping materials (not compatible with Linac4 environment);
PVSS implementation for pressure monitoring for L4 and 3MeV test stand;
Assembly and leak testing of the different type of sources;
LEBT gas density regulation: injection line and control;
Contribution in the design of the sources.

## Slide 17: Conclusions/summary
The Linac4 Source Work Package for vacuum is completed: all the tasks listed in the review of November 2011 are fulfilled.
In order to cope with the maintenance needed for the sources, 1.2 FTE for 2014-2017 are needed: these tasks can be covered by extra resources (industrial support), leading to a total cost of 160 kCHF.
Regarding the hardware maintenance and purchase of new material, the total cost would be around 325 kCHF: 165 kCHF in 2014-2017; 160 by 2018 for the final installation before operation.
If additional tasks are needed (Fast gauges test stand; local monitoring of L4 RGA and design and development of the magnetron source;) no extra resources can be delivered by TE-VSC. A GET fellow paid by the project (210 kCHF per 2 years of contract) could fulfill these activities.
 | Cost (KCHF) | TOTAL Cost (KCHF)
Man power (maintenance) | 160 | 695
Hardware | 325 | 
GET fellow (new activities) | 210 |