## Slide 1
Linac4 ion source review
Beam optics simulations and measurements
Øystein Midttun
European Organization for Nuclear Research
University of Oslo
14.11.2013

## Slide 2
Benchmarking of simulations
Ion beam extraction system
Emittance improvements
This presentations compares H- beam simulations and measurements, and proposes improvements

## Slide 3
Two extraction systems were compared with emphasis on low electron power density and low H- beam divergence
1. e- dump in magnetized Einzel lens
2. e- dump in intermediate electrode

## Slide 4
The new extraction system uses a magnetized Einzel lens to dump co-extracted electrons
Electron beam
H- beam
The electrons are dumped with an energy of 10 keV,
and spread over a large surface to reduce the power density

## Slide 5
Benchmarking of simulations
Ion beam extraction system
Emittance improvements

## Slide 6
Plasma generator
The comparison to simulations has been made
by using the currents measured on four electrodes
Puller
Electron dump
Faraday cup

## Slide 7
Chart: H- density [A/cm2]
| 3.1272379670176764E-2 | 3.2769258633860154E-2 | 3.4459205816281258E-2 | 3.6144830932516393E-2 | 3.7756538636423781E-2 | 3.9201816974075203E-2 | 4.1904188303630081E-2 | 4.2361279384581833E-2 | 4.426031975907932E-2
Plasma generator Meas | 1.0198 | 1.0589 | 1.0778 | 1.1456 | 1.2117 | 1.2409 | 1.3059 | 1.3845 | 1.474
Plasma generator Sim | 0.9783 | 1.0529 | 1.1306 | 1.2095 | 1.2878 | 1.3669
Puller Meas | 0.1238 | 0.126 | 0.12661 | 0.12946 | 0.13467 | 0.13768 | 0.14208 | 0.14939 | 0.1569
Puller Sim | 0.068012 | 0.067828 | 0.069031 | 0.066528 | 0.067407 | 0.073551
Electron dump Meas | 0.8038 | 0.8416 | 0.86301 | 0.92314 | 0.98398 | 1.0097 | 1.0715 | 1.1434 | 1.2193
Electron dump Sim | 0.89785 | 0.97114 | 1.047 | 1.1263 | 1.2027 | 1.2741
Faraday cup Meas | 11.8165 | 11.9245 | 12.447 | 13.023 | 13.8714 | 14.2667 | 14.9305 | 15.6372 | 16.7111
Faraday cup Sim | 10.9 | 12.1 | 13.3 | 14.3 | 15.3 | 16.4
RF-power changes the H- density in the plasma,
thus the extracted beam current
-45
-22
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup
Linear scaling between
RF-power (measurement)
H- density (simulation)

## Slide 8
Chart: H- density [A/cm2]
| 3.1272379670176764E-2 | 3.2769258633860154E-2 | 3.4459205816281258E-2 | 3.6144830932516393E-2 | 3.7756538636423781E-2 | 3.9201816974075203E-2 | 4.1904188303630081E-2 | 4.2361279384581833E-2 | 4.426031975907932E-2
Puller Meas | 0.1238 | 0.126 | 0.12661 | 0.12946 | 0.13467 | 0.13768 | 0.14208 | 0.14939 | 0.1569
Puller Sim | 0.068012 | 0.067828 | 0.069031 | 0.066528 | 0.067407 | 0.073551
Electron dump Meas | 0.8038 | 0.8416 | 0.86301 | 0.92314 | 0.98398 | 1.0097 | 1.0715 | 1.1434 | 1.2193
Electron dump Sim | 0.89785 | 0.97114 | 1.047 | 1.1263 | 1.2027 | 1.2741
The simulations of the puller and electron dump current
do not correspond to the measured values
-45
-22
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup
Secondary electrons emitted from the electron dump
Reduce the current measured on the electron dump
Increase the current on the puller

## Slide 9
The simulations fit the measurement data when secondary electron emission from the electron dump is considered
-45
-22
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup
Chart: H- density [A/cm2]
| 3.1272379670176764E-2 | 3.2769258633860154E-2 | 3.4459205816281258E-2 | 3.6144830932516393E-2 | 3.7756538636423781E-2 | 3.9201816974075203E-2 | 4.1904188303630081E-2 | 4.2361279384581833E-2 | 4.426031975907932E-2
Puller Meas | 0.1238 | 0.126 | 0.12661 | 0.12946 | 0.13467 | 0.13768 | 0.14208 | 0.14939 | 0.1569
Puller Sim | 0.1544 | 0.1553 | 0.1548 | 0.1466 | 0.1426 | 0.1443
Electron dump Meas | 0.8038 | 0.8416 | 0.86301 | 0.92314 | 0.98398 | 1.0097 | 1.0715 | 1.1434 | 1.2193
Electron dump Sim | 0.7959 | 0.8672 | 0.9445 | 1.0259 | 1.1085 | 1.1837
Secondary electrons emitted from the electron dump
Reduce the current measured on the electron dump
Increase the current on the puller

## Slide 10
Chart: H- density [A/cm2]
| 3.1272379670176764E-2 | 3.2769258633860154E-2 | 3.4459205816281258E-2 | 3.6144830932516393E-2 | 3.7756538636423781E-2 | 3.9201816974075203E-2 | 4.1904188303630081E-2 | 4.2361279384581833E-2 | 4.426031975907932E-2
Plasma generator Meas | 1.0198 | 1.0589 | 1.0778 | 1.1456 | 1.2117 | 1.2409 | 1.3059 | 1.3845 | 1.474
Plasma generator Sim | 0.9782 | 1.0531 | 1.1306 | 1.2094 | 1.2877 | 1.3671
Puller Meas | 0.1238 | 0.126 | 0.12661 | 0.12946 | 0.13467 | 0.13768 | 0.14208 | 0.14939 | 0.1569
Puller Sim | 0.1544 | 0.1553 | 0.1548 | 0.1466 | 0.1426 | 0.1443
Electron dump Meas | 0.8038 | 0.8416 | 0.86301 | 0.92314 | 0.98398 | 1.0097 | 1.0715 | 1.1434 | 1.2193
Electron dump Sim | 0.7959 | 0.8672 | 0.9445 | 1.0259 | 1.1085 | 1.1837
Faraday cup Meas | 11.8165 | 11.9245 | 12.447 | 13.023 | 13.8714 | 14.2667 | 14.9305 | 15.6372 | 16.7111
Faraday cup Sim | 10.9 | 12.1 | 13.3 | 14.3 | 15.3 | 16.4
-45
-22
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup
The H- beam extraction has been reproduced for a varying RF-power, with secondary electron emission included in the simulations
Linear scaling between
RF-power (measurement)
H- density (simulation)

## Slide 11
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-18
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup
Measured current increases slightly more than the simulated one
Density varies as a function of the depth of the plasma
Plasma generator current increases with higher extraction voltage
Increased plasma meniscus surface

## Slide 12
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-18
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 13
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
-45
-20
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 14
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-22
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 15
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-24
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 16
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-26
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 17
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-28
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 18
The puller voltage defines the electric extraction field,
and can be used to tune the beam optics
Chart: Vext [kV] (Vpuller - Vsource)
| 15 | 17 | 19 | 21 | 23 | 25 | 27
Plasma generator Meas | 0.76224 | 0.81241 | 0.85243 | 0.88146 | 0.92459 | 0.95465 | 1.007
Plasma generator Sim | 0.9374 | 0.9086 | 0.8784 | 0.846 | 0.8226 | 0.8002 | 0.7773
Puller Meas | 0.052328 | 0.066477 | 0.083587 | 0.10157 | 0.12284 | 0.14445 | 0.16833
Puller Sim | 0.1694 | 0.1521 | 0.1361 | 0.1199 | 0.0959 | 0.0695 | 0.0488
Electron dump Meas | 0.64617 | 0.66984 | 0.69219 | 0.69574 | 0.70651 | 0.71257 | 0.72552
Electron dump Sim | 0.7398 | 0.7281 | 0.7151 | 0.6995 | 0.6997 | 0.7051 | 0.6998
Faraday cup Meas | 11.5208 | 11.2809 | 10.723 | 10.9776 | 10.4083 | 10.2922 | 10.5401
Faraday cup Sim | 10.8 | 10.8 | 10.8 | 11 | 11.2 | 11.2 | 11.5
-45
-30
-36
0
+35
0
V (kV)
Plasma
generator
Einzel
e-dump
Ground
Puller
Einzel
lens
LEBT
Faraday
cup

## Slide 19
The measured beam emittance is similar to what we expect from simulation
Measurement
Simulation
εNORM, 1 RMS = 0.5 µm
εNORM, 1 RMS = 0.5 µm
The emittance should be improved since the RFQ acceptance is 0.25 µm

## Slide 20
Preliminary comparison of the first measured beam from IS02 (12.11.2013) with a simulation
 | Source [A] | Puller [A] | E-dump [A] | Faraday cup [mA] | Emittance [µm]
Measurement | 0.35 | 0.15 | 0.031 | 13 | 
Simulation | 0.38 | 0.065 | 0.030 | 12 | < 0.6
-45
-37
-38
0
+35
0
V [kV]
Simulation input: 12 mA H-, e/H- = 30

## Slide 21
Benchmarking of simulations
Ion beam extraction system
Emittance improvements

## Slide 22
Reducing the number of co-extracted electrons, reduces the emittance growth in the electron dump
Artificial removing of electrons
Simulation input: 30 mA H-, e/H- = 30
-45
-25
-36
0
V [kV]

## Slide 23
The shape of the plasma electrode improves the emittance
DESY
DESY imp
IS02
Simulation input: 40 mA H-, e/H- = 10
-45
-25
-36
V [kV]

## Slide 24
Comparison of different Einzel lens configurations with beam transport through the first solenoid
Acceleration Einzel lens
Deceleration Einzel lens
No Einzel lens (not optimized at the LEBT entry)
Simulation input: 30 mA H-, e/H- = 30

## Slide 25
With the current extraction system, it is feasible
to transport an 80 mA H- beam
DESY
IS02
Simulation input: 80 mA H-, e/H- = 1
H- transmission = 99 %
εNORM, 1 RMS = 0.6 µm
A redesign should be made to not hit the inside of the electrodes with the H- beam, and to collect all electrons in the dump
H- transmission = 79 %
εNORM, 1 RMS = 0.6 µm

## Slide 26
The new extraction system has been commissioned,
and verified with simulations
Stable ion beam extraction with few high voltage breakdowns
Few resources required for operation
  1 expert during start-up
  1 expert on call during operation
Room for improvement of beam current transportation and emittance
Improving the extraction system (simulations, design, production)
  Modifying existing: 6-9 months
  Complete redesign: 9-12 months
I Faraday cup
V e-dump
V puller
V source
Measurements over week-end