1/21
Magnetic Field Mapping
V. Blackmore
CM37
November 7th, 2013

---
2/21
SS Mapping Checklist
Garbage check!
Survey adjustments
Linearity of field with current
  - With and without the Virostek plate
Residual magnetic field
  - With and without the Virostek plate
Hysteresis
  - Should only occur with the Virostek plate
Magnetic axis/coil fits

---
3/21
Mapped Currents
Runs cover the above currents, plus:
0A measurements (residual field)
30A individual coil measurements (superposition)
With and without Virostek plate
A lot of data

---
4/21
Surveys
Mapper position measured as it moves through the SS
  - Does move and is not completely flat
  - Shims under the carriage will improve this for next data set
Surveyed twice, with and without shielding plate
  - SS fiducials did not move when the shielding plate was installed, but the mapper was nudged (so resurvey)
  - No survey when shielding plate was removed (mapper not nudged)

---
5/21
Surveys
Longitudinal motion vs. encoder value
  - Mapper records “local” position
  - Relate to SS co-ordinates
Tilts in the y-z plane (survey co-ordinates)
  - Monitor three points (A, B, C) as a function of longitudinal position
  - Correct Hall probe positions by angle ,
     rotation matrix,
     translation matrix

---
6/21
Surveys

---
7/21
Surveys
Nominal probe position

---
8/21
Surveys
Nominal probe position

---
9/21
What next?
Types of Fit
Minimise all coil parameters and find the best fit
20 parameters for SS!
See the next tale for why this is not preferred
“Mixing/scaling” fit
Take two models and mix/scale for best fit
Preferred method!
Both require Fourier-Bessel
Residual field fit, gets us to desired accuracy
Difficult
Compare zero-field measurements
  - Jitter on Hall probe readings, location of welds
  - Make angular comparison (field at ?)
Check linearity (field at ?)
Check superposition of individual coils at 30A is the same as all coils at 30A
  - Slightly complicated by centre coil, powered in series... However, that is also a check!
Find the magnetic centre, find the best “mixing-scaling” fit to the SS
  - See CM36 talk for details of the fit procedure!
Confident in results? Then do Fourier-Bessel analysis

---
10/21
Once upon a time, in a land very much nearby, there lived a Friendly Creature. Though the Friendly Creature was very, very friendly, it was also very sad. No matter how hard it tried, it always had the nagging dread that it was letting other people down.
One day, after supping on it’s regular meal of cold, noble soup, it resolved to try and achieve its dreams one last time. If it could prove itself other Friendly Creatures would follow! The Friendly Creature’s friends gathered around it and watched it with some trepidation. It had wrapped itself in dials and gauges, determined to prove that it could achieve its lofty goal.
Bit by bit, the dials crept up. Slowly, so slowly, the Friendly Creature stalked forward. For the briefest of moments it achieved enlightenment. It had done it, it was there! But in that lucid moment it realised it had no-one to please other than itself and it no longer cared so much about its goal. What will be will be, it was what it was, and all it really wanted was to sleep.
A Slight Diversion...

---
11/21
Unexpected Behaviour...
Rough field measurements taken whilst ramping
Definite non-linear behaviour
Strange!
(But not unexplainable...)

---
12/21
Unexpected Behaviour...
Thanks to V. Bayliss & S. Watson,
I. Taylor & C. Pidcott
Rough field measurements taken whilst ramping
Definite non-linear behaviour
Strange!
(But not unexplainable...)
Took more data!
I (A)
B (kG)

---
13/21
“Weirdness” Theory #1
Perhaps that field takes time to “settle”
  - Continuous ramping  field has to play catch-up
Recorded field every second and compared to ramp time in Archiver
Field “settles” very quickly
Not the origin of the non-linear field

---
14/21
“Weirdness” Theory #1
Perhaps that field takes time to “settle”
  - Continuous ramping  field has to play catch-up
Recorded field every second and compared to ramp time in Archiver
Field “settles” very quickly
Not the origin of the non-linear field

---
15/21
“Weirdness” Theory #2
Is the Hall probe non-linear?
  - No iron, so field should be linear.
  - “Off-axis” measurements at 50, 100, 150A
  - 2*50A, 0.5*150A should equal the 100A measurements

---
16/21
“Weirdness” Fact #2
Is the Hall probe non-linear?
  - Calibration magnet/doorstop, ~0.5T measured exactly be probe
Compare 50A measurements to 100, 150A
  - Should get a straight line, but it deviates  Hall probe in non-linear region!

---
17/21
“Weirdness” Fact #2
Is the Hall probe non-linear?
  - Calibration magnet/doorstop, ~0.5T measured exactly be probe
Can we trust the 50A measurements?
  - Compare to “as built” calculation  compatible within errors

---
18/21
50A Field Fits
Assume probe is ~10—15mm off-axis
Start with the drawings for coil dimensions and separation

---
19/21
50A Field Fits
Assume probe is ~10—15mm off-axis
Start with the drawings for coil dimensions and separation

---
20/21
50A Field Fits
Assume probe is ~10—15mm off-axis
Start with the drawings for coil dimensions and separation
2D fits show how complicated life is for just two coils!

---
21/21
Summary
SS field mapping analysis initiated, but not yet complete
FC field maps highlight the dangers of 20-parameter fits!
  - Simple set of measurements (though off-axis)
  - Difficult to pin down sources of uncertainty
  - Will do it better for FC2!
Finally, you can do some fairly useful measurements with a “Hall probe on a stick”!