## Slide 1: Introduction to Beam Diagnostics
Ulrich Raich
CERN BE - BI
(Beam Instrumentation)

## Slide 2: A few depicted examples
Introduction
Beam presence: Scintillating screens (LHC)
Intensity measurement, Faraday Cup and Transformer (Linac-4)
Transverse Profile measurement, wire scanner & wire grids (PSB & PS)
Emittance measurement
  Slit and Grid (Linac-4, 3 MeV line)
  Emittance measurement line (Linac-2)
  Longitudinal Emittance measurement (Linac-2)
Trajectory measurement (LHC and PS) using Beam Position Monitors (BPMs)
Longitudinal phase space: Tomoscope (PS) using a wall current monitor
Tune measurement (SPS) using BPMs
Losses: Beam Loss Monitors (BLMs) (LHC)

## Slide 3: Introduction
An accelerator can never be better than the instruments measuring its performance!

## Slide 4: Diagnostic devices and quantity measured
Instrument | Physical Effect | Measured Quantity | Effect on beam
Faraday Cup | Charge collection | Intensity | Destructive
Current Transformer | Magnetic field | Intensity | Non destructive
Wall current monitor | Image Current | Intensity Longitudinal beam shape | Non destructive
Pick-up | Electric/magnetic field | Position, Tune | Non destructive
Secondary emission monitor | Secondary electron emission | Transverse size/shape, emittance | Disturbing, can be destructive at low energies
Wire Scanner | Secondary particle creation | Transverse size/shape | Slightly disturbing
Scintillator screen | Atomic excitation with light emission | Transverse size/shape (position) | Destructive
Residual Gas monitor | Ionization | Transverse size/shape | Non destructive

## Slide 5: Different uses of beam diagnostics
Regular crude checks of accelerator performance
  Beam Intensity
  Radiation levels
Standard regular measurements
  Emittance measurement
  Trajectories
  Tune
Sophisticated measurements e.g. during machine development sessions
  May require offline evaluation
  May be less comfortable

## Slide 6: A beam parameter measurement

## Slide 7: Beam Presence Scintillating Screens
Method already applied in cosmic ray experiments
Very simple
Very convincing
Needed:
Scintillating Material
TV camera
In/out mechanism
Problems:
Radiation resistance of TV camera
Heating of screen (absorption of beam energy)
Evacuation of electric charges

## Slide 8: Screen mechanism
Screen with graticule

## Slide 9: Results from TV Frame grabber
For further evaluation the video signal is digitized, read-out and treated by program
First full turn
as seen by the BTV
10/9/2008
Un-captured beam sweeps through he dump line

## Slide 10: Beam Intensity Layout of a Faraday Cup
Electrode: 1 mm stainless steel
Only low energy particles can be measured
Very low intensities (down to 1 pA) can be measured
Creation of secondary electrons of low energy (below 20 eV)
Repelling electrode with some 100 V polarisation voltage pushes secondary electrons back onto the electrode

## Slide 11: Faraday Cup

## Slide 12: Electro-static Field in Faraday Cup
In order to keep secondary electrons with the cup a repelling voltage is applied to the polarization electrode
Since the electrons have energies of less than 20 eV some 100V repelling voltage is sufficient

## Slide 13: Energy of secondary emission electrons
With increasing repelling voltage the electrons do not escape the Faraday Cup any more and the current measured stays stable.
At 40V and above no decrease in the Cup current is observed any more

## Slide 14: Intensity Principle of a fast current transformer
Diagram by H. Jakob
Image
Current
BEAM
Calibration winding
500MHz Bandwidth
Low droop (< 0.2%/ms)
Ceramic Gap
80nm Ti Coating
20W to improve
impedance

## Slide 15: Magnetic shielding
Shield should extend along the vacuum chamber length > diameter of opening
Shield should be symmetrical to the beam axis
Air gaps must be avoided especially along the beam axis
Shield should have highest μ possible but should not saturate
Permalloy (μ3)
Transformer steel (μ2)
Soft iron (μ1)

## Slide 16: Calibration of AC current transformers
The transformer is calibrated with a very precise current source
The calibration signal is injected into a separate calibration winding
A calibration procedure executed before the running period
A calibration pulse after the beam pulse measured with the beam signal

## Slide 17: Profile measurements
Secondary emission grids (SEMgrids)
The ejected electrons are taken away by polarization voltage
When the beam passes secondary electrons are ejected from the wires
The current flowing back onto the wires is measured

## Slide 18: Profiles from SEMgrids
Projection of charge density
projected to x or y axis is
Measured
One amplifier/ADC per wire
Large dynamic range
Resolution is given by wire
distance
Used only in transfer lines

## Slide 19: Wire Scanners
A thin wire is quickly moved across the beam
Secondary particle shower is detected outside the vacuum chamber
on a scintillator/photo-multiplier assembly
Position and photo-multiplier signal are recorded simultaneously

## Slide 20: Wire scanner profile
High speed needed
because of heating.
Adiabatic damping
Current increase due to
speed increase
Speeds of up to 20m/s
=> 200g acceleration

## Slide 21: Emittance measurements
A beam is made of many, many particles,
each one of these particles is moving with
a given velocity. Most of the velocity
vector of a single particle is parallel to the
direction of the beam as a whole (s).
There is however a smaller component of
the particles velocity which is
perpendicular to it (x or y).
Design by E. Bravin

## Slide 22: Emittance measurements
If for each beam particle we plot its position and its transverse angle we get a particle distribution whose boundary is an usually ellipse.
The projection onto the x axis is the beam size
x’
x
Beam size

## Slide 23: The slit and grid method
If we place a slit into the beam we cut out a small vertical slice of phase space
Converting the angles into position through a drift space allows to reconstruct the angular distribution at the position defined by the slit
x’
x
slit

## Slide 24: Transforming angular distribution to profile
When moving through a drift space the angles don’t change (horizontal move in phase space)
When moving through a quadrupole the position does not change but the angle does (vertical move in phase space)
x’
x
slit
x’
x
slit
Influence of a drift space
Influence of a quadrupole
x’
x
slit

## Slide 25: The Slit Method

## Slide 26: Emittance Meter
SEM grid read-out 250kHz
Stepping motors to allow coarse + fine position tuning

## Slide 27: Transverse Emittance Measurement
Slit and grid phase space scanner
L-shaped 0.1mm slit moves under
45 degrees
Slit and grids move independently
Positioning precision: 50 µm
Movement PLC controlled
Slit and grids mounted in 2 independent vacuum boxes which can be separated
Horizontal and vertical SEMGrid
wire distance .75 mm
30 signal wires
readout with home built 36 channel 250 kHz ADC
time resolved profiles

## Slide 28: Emittance plot Solenoid

## Slide 29: Single pulse emittance measurement
Kickers
slit
SEMgrid
Every 100 ns
a new profile
Quadrupole
Quadrupole

## Slide 30: Transformation in Phase Space
X’
X
X’
X
X’
X
X’
X
X’
X
Quadrupole
Quadrupole
SEMGrid
Kicker
Kicker
at slit
first drift space
first quadrupole
second drift space
Second quadrupol

## Slide 31: Result of single pulse emittance measurement

## Slide 32: Longitudinal Emittance measurement
E
φ
E
E
φ
Kicker
Buncher
Kicker
SEMGrid
Transformer
Spectrometer magnet
E
φ
at slit
first drift space
buncher
second drift space
φ
Spectrometer produces image of slit on second slit
second slit selects energy slice
first kicker sweep phase space over all energies
buncher rotates energy slice in phase space
at second spectrometer the phase distribution is transformed into an energy distribution analyzed by the second spectrometer
second kicker corrects for first kick
Buncher RF

## Slide 33: Trajectory and Orbit Measurements
Measure the particle position
of each bunch as it travels
around the ring

## Slide 34: The PS, a universal machine
All beams pass through the PS
Different particle types
Different beam characteristics
Concept of a super cycle

## Slide 35: The Supercycle

## Slide 36: Position Measurements
BPM signals sampled at 120 MHz
Red: The sum signal
Green: The difference signal
Procedure:
Produce integration gates and
baseline signals
Baseline correct both signals
Integrate sum and difference signals
and store results in memory
Take external timing events into account e.g. harmonic number change, γ-transition etc.

## Slide 37: Baseline restoration
Low pass filter the signal to get an estimate of the base line
Add this to the original signal

## Slide 38: Trajectory measurements in circular machines
Needs integration gate
Can be rather tricky
Distance between bunches
changes with acceleration
Number of bunches
may change

## Slide 39: Beams in the PS

## Slide 40: RF Gymnastics
Bunch splitting or recombination One RF frequency is gradually decreased while the other one is increased
The gate generator must be synchronized

## Slide 41: Trajectory readout electronics

## Slide 42: Changing bunch frequency
Bunch splitting or recombination
One RF frequency is gradually decreased while the other one is increased
Batch compression
For all these cases the gate generator must be synchronized

## Slide 43: Tune measurements
When the beam is displaced (e.g. at injection or with a deliberate kick, it starts to oscillate around its nominal orbit (betatron oscillations)
Measure the trajectory
Fit a sine curve to it
Follow it during one revolution
kicker

## Slide 44: The Sensors
The kicker
Shoebox pick-up
with linear cut

## Slide 45: Kicker + 1 pick-up
Measures only non-integral part of Q
Measure a beam position at each revolution
Fourier transform BPM signal
Search peak in Fourier Spectrum
Fourier transform of pick-up signal

## Slide 46: Q-Measurement Results

## Slide 47: Direct Diode Detection Base Band Q measurement
Diode Detectors convert spikes to saw-tooth waveform
Signal is connected to differential amplifier to cut out DC level
Filter eliminates most of the revolution frequency content
Output amplifier brings the signal level to amplitudes suitable for long distance transmission

## Slide 48: BBQ Results from CERN SPS
Results from Sampling
After Fourier Transform

## Slide 49: Tune feedback at the LHC

## Slide 50: Computed Tomography (CT)
Principle of Tomography:
Take many 2-dimensional Images at different angles
Reconstruct a 3-dimensional picture using mathematical techniques (Algebraic Reconstruction Technique, ART)

## Slide 51: The reconstruction
Produce many projections of the object to be reconstructed
Back project and overlay the “projection rays”
Project the back-projected object and calculate the difference
Iteratively back-project the differences to re-construct the original object

## Slide 52: Some CT resuluts

## Slide 53: Computed Tomography and Accelerators
RF voltage
Restoring force for non-synchronous particle
Longitudinal phase space
Projection onto Φ axis corresponds to bunch profile

## Slide 54: Reconstructed Longitudinal Phase Space
Courtesy S. Hancock

## Slide 55: Bunch Splitting

## Slide 56: Stored Beam Energies
(Based on graph from R. Schmidt)
Quench Levels | Units | Tevatron | RHIC | HERA | LHC
Instant loss (0.01 - 10 ms) | [J/cm3] | 4.5 10-03 | 1.8 10-02 | 2.1 10-03 - 6.6 10-03 | 8.7 10-04
Steady loss (> 100 s) | [W/cm3] | 7.5 10-02 | 7.5 10-02 |  | 5.3 10-03
Notes: LHC extrapolation
Energy in beam 7 times Tevatron
Intensity of beam 30 times Tevatron
Number of LHC of moving elements in the LHC about 200
Energy stored in both beams: 0.7 GJoule
heat up from cryogenic temperature and melt
0.1 m3 of Copper

## Slide 57: Beam power in the LHC
The Linac beam (160 mA, 200μs, 50 MeV, 1Hz) is enough to burn a hole into
the vacuum chamber
What about the LHC beam: 2808 bunches of 15*1011 particles at 7 TeV?
1 bunch corresponds to a 5 kg bullet at 800 km/h

## Slide 58: Beam Loss Monitor Types
Design criteria: Signal speed and robustness
Dynamic range (> 109) limited by leakage current through insulator ceramics (lower) and saturation due to space charge (upper limit).
Ionization chamber:
  N2 gas filling at 100 mbar over-pressure
  Length 50 cm
  Sensitive volume 1.5 l
  Ion collection time 85 s
Both monitors:
  Parallel electrodes (Al, SEM: Ti) separated by 0.5 cm
  Low pass filter at the HV input
  Voltage 1.5 kV
Secondary Emission Monitor (SEM):
  Length 10 cm
  P < 10-7 bar
  ~ 30000 times smaller gain
Notes: Chamber design considerations: Signal speed and robustness against ageing (and loss of chamber gas)
BLMB: Planned to use standard ACEM detectors (or photo-multipliers), and acquisition systems of LHC BCT. Not part of baseline scenario, for refined beam observation.

## Slide 59: Quench levels

## Slide 60: Industrial production of chambers
Beam loss must be
measured all around
the ring
=> 4000 sensors!

## Slide 61: Conclusions
Beam diagnostics is a very wide field where many different competences are needed
  Machine physics
  Electronics
  Computing
  Mechanics
The instruments are the eyes with which we observe the beam
The beam can never be adjusted with higher precision than what can be measured