## Slide 1: Beam Diagnostics Lecture 1
Ulrich Raich
CERN BE - BI
(Beam Instrumentation)

## Slide 2: Overview
First hour:
  Introduction
  Overview of measurement instruments
    Faraday Cup
    Beam Current Transformer
    Beam Position Monitor
    Fluorescence screens
    Profile Detectors
      SEMGrids
      Wire Scanners
    Beam Loss Monitors
Second hour
  Some depicted examples of beam parameter measurements

## Slide 3: Introduction
An accelerator can never be better than the instruments measuring its performance!

## Slide 4: Different uses of beam diagnostics
Regular crude checks of accelerator performance
  Beam Intensity
  Radiation levels
Standard regular measurements
  Emittance measurement
  Trajectories
  Tune
Sophisticated measurements e.g. during machine development sessions
  May require offline evaluation
  May be less comfortable

## Slide 5: Diagnostic devices and quantity measured
Instrument | Physical Effect | Measured Quantity | Effect on beam
Faraday Cup | Charge collection | Intensity | Destructive
Current Transformer | Magnetic field | Intensity | Non destructive
Wall current monitor | Image Current | Intensity Longitudinal beam shape | Non destructive
Pick-up | Electric/magnetic field | Position | Non destructive
Secondary emission monitor | Secondary electron emission | Transverse size/shape, emittance | Disturbing, can be destructive at low energies
Wire Scanner | Secondary particle creation | Transverse size/shape | Slightly disturbing
Scintillator screen | Atomic excitation with light emission | Transverse size/shape (position) | Destructive
Residual Gas monitor | Ionization | Transverse size/shape | Non destructive

## Slide 6: A beam parameter measurement

## Slide 7: Required Competence in a beam diagnostics group
Some beam physics in order to understand the beam parameters to be measured and to distinguish beam effects from sensor effects
Detector physics to understand the interaction of the beam with the sensor
Mechanics
Analogue signal treatment
  Low noise amplifiers
  High frequency analogue electronics
Digital signal processing
Digital electronics for data readout
Front-end and Application Software

## Slide 8: Bethe Bloch formula
with the following constants:
NA: Avogadro’s number
me and re: electron rest mass and classical electron radius
c: speed of light
the following target material properties:
ρ: material density
AT and ZT: the atomic mass and nuclear charge
and the particle properties:
Zp: particle charge
β: the particles velocity and
Dependance on

## Slide 9: High energy loss a low energies
Heavy ions at low energy are stopped within a few micro-meters
All energy is deposited in a very small volume

## Slide 10: Scintillating Screens
Method already applied in cosmic ray experiments
Very simple
Very convincing
Needed:
Scintillating Material
TV camera
In/out mechanism
Problems:
Radiation resistance of TV camera
Heating of screen (absorption of beam energy)
Evacuation of electric charges

## Slide 11: Test for resistance against heat-shock
Material | r g/cm3 | cp at 20ºC J/gK | k at 100ºC W/mK | Tmax ºC | R at 400 ºC Ω.cm
Al2O3 | 3.9 | 0.9 | 30 | 1600 | 1012
ZrO2 | 6 | 0.4 | 2 | 1200 | 103
BN | 2 | 1.6 | 35 | 2400 | 1014
Better for electrical conductivity (>400ºC)
Better for thermal properties
(higher conductivity, higher heat capacity)

## Slide 12: Degradation of screen
Degradation clearly visible
However sensitivity stays essentially
the same

## Slide 13: Screen mechanism
Screen with graticule

## Slide 14: Results from TV Frame grabber
For further evaluation the video signal is digitized, read-out and treated by program
First full turn
as seen by the BTV
10/9/2008
Uncaptured beam sweeps through he dump line

## Slide 15: Layout of a Faraday Cup
Electrode: 1 mm stainless steel
Only low energy particles can be measured
Very low intensities (down to 1 pA) can be measured
Creation of secondary electrons of low energy (below 20 eV)
Repelling electrode with some 100 V polarisation voltage pushes secondary electrons back onto the electrode

## Slide 16: Faraday Cup

## Slide 17: Electro-static Field in Faraday Cup
In order to keep secondary electrons within the cup a repelling voltage is applied to the polarization electrode
Since the electrons have energies of less than 20 eV some 100V repelling voltage is sufficient

## Slide 18: Energy of secondary emission electrons
With increasing repelling voltage the electrons do not escape the Faraday Cup any more and the current measured stays stable.
At 40V and above no decrease in the Cup current is observed any more

## Slide 19: Current Transformers
Beam current
Magnetic field
ri
ro
Fields are very low
Capture magnetic field lines with cores of high
relative permeability
(CoFe based amorphous alloy Vitrovac: μr= 105)

## Slide 20: The ideal transformer
Beam signal
Transformer output signal
Inductance L of the winding
RF
RL
CS
A
Ls
RL
CS
R
The AC transformer
The active AC transformer

## Slide 21: Principle of a fast current transformer
Diagram by H. Jakob
Image
Current
BEAM
Calibration winding
500MHz Bandwidth
Low droop (< 0.2%/ms)
Ceramic Gap
80nm Ti Coating
20W to improve
impedance

## Slide 22: Fast current transformers for the LHC

## Slide 23: Magnetic shielding
Shield should extend along the vacuum chamber length > diameter of opening
Shield should be symmetrical to the beam axis
Air gaps must be avoided especially along the beam axis
Shield should have highest μ possible but should not saturate
monitor
Soft iron (μ1)
Transformer steel (μ2)
Permalloy (μ3)

## Slide 24: Calibration of AC current transformers
The transformer is calibrated with a very precise current source
The calibration signal is injected into a separate calibration winding
A calibration procedure executed before the running period
A calibration pulse before the beam pulse measured with the beam signal

## Slide 25: Current transformer and electronics

## Slide 26: Display of transformer readings
First result from LHC FBCT
Measurement of bunch intensity
Diminishing intensity due to debunching
Beam losses will trigger machine protection system

## Slide 27: The DC current transformer
AC current transformer can be extended to very long droop times but not to DC
Measuring DC currents is needed in storage rings
Must provide a modulation frequency
Takes advantage of non/linear magnetisation curve
B
H

## Slide 28: Principle of DCCT
beam
Compensation
current Ifeedback=-Ibeam
modulator
V=RIbeam
Power supply
R
Synchronous
detector
Va-Vb
Vb
Va

## Slide 29: Modulation of a DCCT without beam
Modulation current has only odd harmonic frequencies since the signal is symmetric
1
2
B=f(t)
B
H
5
3
4

## Slide 30: Modulation of a DCCT with beam
1
B=f(t)
B
H
1
2
5
3
4
Sum signal becomes non-zero
Even harmonics appear

## Slide 31: Modulation current difference signal with beam
Difference signal has 2ωm
ωm typically 200 Hz – 10 kHz
Use low pass filter with ωc<< ωm
Provide a 3rd core, normal AC transformer to extend to higher frequencies

## Slide 32: Photo of DCCT internals

## Slide 33: Results from DCCT
Beam 2
DCCT sees first
circulating beam
Injections into LHC

## Slide 34: Measuring Beam Position – The Principle
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
Slide by R. Jones

## Slide 35: Wall Current Monitor – The Principle
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
V
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Ceramic Insert

## Slide 36: Wall Current Monitor – Beam Response
L
IB
V
R
C
Frequency
Response
0
0
IB

## Slide 37: Electrostatic Monitor – The Principle
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-

## Slide 38: Electrostatic Monitor – Beam Response
VB
V
R
C
Frequency (Hz)
Response (V)
0
0

## Slide 39: Position measurements
U
If the beam is much smaller than w, all field lines are captured and
U is a linear function with replacement
else: Linear cut (projection to measurement plane must be linear)
w
U
d
d

## Slide 40: Shoebox pick-up
Linear cut through a shoebox
UL
UR

## Slide 41: Doubly cut shoebox
Can measure horizontal and vertical position at once
Has 4 electrodes
a
b
c
d

## Slide 42: Simulatenous horizontal and vertical measurement
a
b
c
d
horizontal
vertical
a
b
c
d

## Slide 43: Interaction of particles with matter
Coulomb interaction
Average force in s-direction=0
Average force in transverse direction <> 0
Mostly large impact parameter => low energy of ejected electron
Electron mostly ejected transversely to the particle motion

## Slide 44: Profile measurements
Secondary emission grids (SEMgrids)
The ejected electrons are taken away by polarization voltage
When the beam passes secondary electrons are ejected from the wires
The current flowing back onto the wires is measured

## Slide 45: Profiles from SEMgrids
Charge density
projected to x or y axis is
measured
One amplifier/ADC per wire
Large dynamic range
Resolution is given by wire
distance
Used only in transfer lines

## Slide 46: Wire Scanners
A thin wire is quickly moved across the beam
Secondary particle shower is detected outside the vacuum chamber
on a scintillator/photo-multiplier assembly
Position and photo-multiplier signal are recorded simultaneously

## Slide 47: Wire scanner profile
High speed needed
because of heating.
Adiabatic damping
Current increase due to
particle speed increase
Wire speeds of up to
20m/s
=> 200g acceleration

## Slide 48: Stored Beam Energies
(Based on graph from R. Schmidt)
Quench Levels | Units | Tevatron | RHIC | HERA | LHC
Instant loss (0.01 - 10 ms) | [J/cm3] | 4.5 10-03 | 1.8 10-02 | 2.1 10-03 - 6.6 10-03 | 8.7 10-04
Steady loss (> 100 s) | [W/cm3] | 7.5 10-02 | 7.5 10-02 |  | 5.3 10-03
Notes: LHC extrapolation
Energy in beam 7 times Tevatron
Intensity of beam 30 times Tevatron
Number of LHC of moving elements in the LHC about 200
Energy stored in both beams: 0.7 GJoule
heat up from cryogenic temperature and melt
0.1 m3 of Copper

## Slide 49: Beam power in the LHC
The Linac beam (160 mA, 200μs, 50 MeV, 1Hz) is enough to burn a hole into
the vacuum chamber
What about the LHC beam: 2808 bunches of 15*1011 particles at 7 TeV?
1 bunch corresponds to a 5 kg bullet at 800 km/h

## Slide 50: Beam Dammage
Fermi Lab‘sTevatron has 200 times less beam power than LHC!

## Slide 51: Beam Loss Monitor Types
Design criteria: Signal speed and robustness
Dynamic range (> 109) limited by leakage current through insulator ceramics (lower) and saturation due to space charge (upper limit).
Ionization chamber:
  N2 gas filling at 100 mbar over-pressure
  Length 50 cm
  Sensitive volume 1.5 l
  Ion collection time 85 s
Both monitors:
  Parallel electrodes (Al, SEM: Ti) separated by 0.5 cm
  Low pass filter at the HV input
  Voltage 1.5 kV
Secondary Emission Monitor (SEM):
  Length 10 cm
  P < 10-7 bar
  ~ 30000 times smaller gain
Notes: Chamber design considerations: Signal speed and robustness against ageing (and loss of chamber gas)
BLMB: Planned to use standard ACEM detectors (or photo-multipliers), and acquisition systems of LHC BCT. Not part of baseline scenario, for refined beam observation.

## Slide 52: Quench levels

## Slide 53: Industrial production of chambers
Beam loss must be
measured all around
the ring
=> 4000 sensors!

## Slide 54: System layout

## Slide 55: Successive running sums