## Slide 1
PART IV:
The Superconducting Synchrocyclotron Project S2C2
1
MID 42331

## Slide 2: Outline
Intro (ProteusONE® and compact gantry)
Main features and parameters
Subsystems:
      Magnetic Circuit
      Superconducting coil
      RF System
      Ion Source+Central Region
      Extraction System
      Extracted beam line
Current status

## Slide 3: The New IBA Single Room Proton Therapy Solution: ProteusONE
Synchrocyclotron with superconducting coil: S2C2
New Compact Gantry for pencil beam scanning
Patient treatment room
High quality PBS cancer treatment: compact and affordable
30.4 m
12.8 m

## Slide 4: The new compact gantry for pencil beam scanning
Design aimed at reducing footprint and cost
Scanning magnets are placed upstream of the last bending magnet
ESS integrated in the 45 deg inclined part
Rotation angle 220°
Gantry has been fully manufactured and tested. Now installed at the customer site in Shreveport where the beam has already been transported to the gantry isocenter

## Slide 5: How a Synchrocyclotron differs from a Cyclotron
Isochronous cyclotron:
  Requires B to increase proportionally to m.
  Requires sector focusing for vertical stability.
  This leads to a smaller average magnetic field, thus a larger structure.
  All parameters being constant, operation is CW.
Synchro cyclotron:
  Requires B to decrease for weak focusing.
  Requires f to decrease during acceleration.
  Smaller structure due to high average magnetic field.
  Acceleration being frequency dependant, operation is pulsed.
5

## Slide 6: What is a Synchrocyclotron ?
A Synchrocyclotron is a circular accelerator in which the frequency of the accelerating electric field is modulated in order to compensate the decreasing period of revolution of the particles. This decrease is due to the relativistic mass increase and the decreasing field during acceleration.
6

## Slide 7: S2C2 overview
General system layout and parameters
An invited talk on this project was presented at the 2013 cyclotron conference in Vancouver
Several contributions can be found on the ECPM2012-website
Notes: The idea grew in more detailed drawings and specifications, thanks to a fully dedicated team of Physicists and engineers inside and outside IBA

## Slide 8: S2C2 overview
Main subsystems
RF system
Extracted beam line
Yoke lifting system
Magnet yoke
Superconducting coil
Ion source+central region
Extraction system
Vacuum system
Notes: As for every guided tour, a small map is always welcome…

## Slide 9: Magnetic circuit-modeling
OPERA3D full model with many details
Long and tedious optimization process
Yoke iron strongly saturated
Influence of external iron systems on the internal magnetic field
Stray-field => shielding of rotco and cryocoolers
pole gap < => extraction system optimization
Influence of yoke penetrations
Median plane errors
Magnetic forces
ITERATIVE PROCESS WITH STRONG INTERACTION TO BEAM SIMULATIONS

## Slide 10: Magnetic circuit modeling
Some examples of OPERA2D results
Pole-gap profile
Iron and coil fields
Bmax on the coil

## Slide 11: Magnetic circuit
Forces acting on the coil => implications for tierod design
MEASURED
SIMULATED
Unstable forces act on displaced cold mass
  2 tons/mm
Large forces for well-centered coil due to
  External elements
  Extraction system
  Vertical asymmetries

## Slide 12: Magnetic circuit: Centering of the cold mass (1)
An innovative method was developed
Measure profiles in between the coils where the field is coil-dominated
Three ports are available: RF-port, source port and vacuum port
Horizontal centering: measure Bz(r) where dBz/dr is maximum
Vertical centering: measure Br(r) where Bz=0

## Slide 13: Magnetic circuit: centering of the cold mass (2)
Minimize the vertical forces on the cold mass
Forces are measured with strain gauges placed in the tierod-assemblies
Opera calculations:
  Low vertical forces  small median plane errors
Vertical forces are reduced by coil-centering and by external shimming
Horizontal forces are large due to extraction system and yoke-asymmetries
During ramp-up coil moves 1.5 mm in to the centered position

## Slide 14: Magnetic circuit:
Dynamic behavior studied with Opera2D FE transient circuit model
Eddy current losses
AC losses in superconductor
Ramp-up characteristics
Quench behavior (qualitative)
Ramp up characteristics
Eddy current losses

## Slide 15: Superconducting coil
Designed and manufactured by ASG (Genua, Italy)
Sumitomo
SRDK-415D
cryocoolers
Beam exit port
RF system port
Vertical tie rods
NbTi wire in channel coil.
Suspended cold mass: 3tons.
Overall weight: 4tons.
Nominal current: 650A (56 A/mm²).
Nominal ampere-turns: 4.3x10+6At.
Conduction cooled by 4 SHI cryocoolers.
9 Inconel tension rods with strain gauges.
(radial ø14mm; upper ø8mm; lower ø6mm)
Cryostat is the cyclotron vacuum chamber.
Radial tie rod
Notes: Niobium Titane coils, suspended cold mass of 3 tons though 9 tension rods, energy stored of 10 Mjoules.
The coil is dry cooled thankx to 4 cryo coolers, with cryo compressors stored near by, for compactness, safety
The coil has been fabricated by ASG-geneva
2 on top of each other Cryocooler configuration is to improve conduction path from coil to cold head without affecting the flux return cross section.

## Slide 16: Superconducting coil
Cryostat penetrations
Four cryostat penetrations
  RF port
  Vacuum pumping/radial probe
  Ion source
  Beam extraction
Aluminium coil former

## Slide 17: Superconducting coil
Installation of the cryostat in the yoke
Notes: That was a real tight fit. 1mm clearance between yoke and outer wall of cryostat.

## Slide 18: RF-system
/2-structure operating in 1th harmonic mode: terminated by the 180°dee on one side and the rotco on the other side
Biased at 1 kVolt to supress mulipactor
Two side stubs provide fine-tuning of df/dt during capture
RF Frequency: 60~90MHz
Modulation frequency: 1kHz
Dee voltage: 3~12kV
Extensively modeled with CST
Placed outside yoke in shielded volume
Fully assembled in cyclotron and tested
A triode-based self-oscillating RF system
RF system on the test bench
Notes: Work of art of M.Abs and J.Mandrillon.
Self oscillating mecanism with adjustable frequency from 60 to 90 MHz, modulation done through a compact Rotating condenser rotating at 7500RPM.
1kHz pulse frequency. Probably the fasted in the world.
Very low consumption (low Dee and Peak power)

## Slide 19: RF-system
Adjustable stub
Oscillator
Pyrometer
Rotco
Servo motor
Turbo pump
layout
Liner
RF pick-up
Dee
Line through cryostat
Vacuum feedthrough
Notes: Stub allow to modify one end of the frequency span, bring in cooling water to the dee and bring a bias voltage to effectively kill all possibility of having multipactoring inside the RF structure.

## Slide 20: RF-system
Innovative/patented design:
excellent mechanical stability and good pulse reproducibility
Stator: 8 blades with a carefully designed profile to have the desired df/dt curve.
Rotor: wheel with 2x8 electrodes turning at 7500rpm (1 kHz pulse)
Rotco=> Coaxially mounted with 8-fold symmetry
Latest stator blade optimization
Notes: Difficult to see here what is fixed and was is supposed to rotate. The next slide will make things a bit clearer.

## Slide 21: Ion source
Use of cold cathode PIG internal ion source
For fast and precise PBS treatment we want:
  High dynamic range in output current (factor 100)
  Good pulse-to-pulse stability (<1%)
  Fast rise times (< 10 msec)
Best candidate is cold cathode PIG source:
  Slow thermal effects/decreased thermal drifts
  Long lifetime (less evaporation of cathodes)
  Pulse frequency 1 kHz (synchronized with rotco)
  Pulse duration about 50 msec
Test-stand measurements:
  Up to 6 mA H+ current extracted
  Good reproducibility of plasma impedance
  Very good arc-current stability at 5 Tesla

## Slide 22: Ion source and central region
The Ion Source and the central region, can be extracted as one assembly for easy maintenance and precise repositioning, without turning down the magnetic field.
Assembly
Insulator
Puller
Chimney
Part of dee
Polarized counter dee
Source shaft
Both dee and counter dee are biased at 1 kV
Notes: Central region can be adjusted outside and brought inside radially along with the ion source.
Advantage : easy adjustmen outside the machine, high yield (several mAmps)
Contacts to connect puller to dee.
Cones for vertical focusing.

## Slide 23: Ion source and central region
Due to the high magnetic field (5,74T) and the low dee voltage (11kV), the source has to be extremely compact:
Central region size
Puller
cm
Ø6mm
cm
Chimney
Chimney
Puller
Source diameter < 5 mm
Vertical gap in the center 6 mm
First 100 turns within a radius of 3 cm
Tapered vertical gaps in the center, to reduce transit time factors

## Slide 24: Layout of the extraction system
Fully passive system => only soft iron
Slow acceleration in a synchrocyclotron
=> Electrostatic deflector difficult to use.
Use resonant extraction based on 2Qh=2 resonance.
Strong local field bump produced by regenerator increases horizontal betatron frequency and locks it to unity.
Unstable orbit is pushed outward by first harmonic into the extraction channel.
Horizontal focusing by gradient corrector and permanent magnet quadrupole (PMQ) in strongly decreasing field.
Correction bars are needed to reduce strong first harmonic error during acceleration.

## Slide 25: Regenerative extraction based on 2Qh=2 resonance
Regenerator bump increases Qh and locks it to 1
Strong enough field bump => Qh=1
Avoid Walkinshaw resonance (Qh-2Qv)
A steady shift of the beam towards the extraction channel builts up

## Slide 26: Extraction system
Layout (1)
Harmonic coils
Correction bars
Regenerator
Graphite beam stops
Magnetic septum
Anti-septum
3 bar gradient corrector
Harmonic coils
Notes: Acceleration os done though 45 cm and Roughly 40000turns is accomplished.
Passive magnetic septum for simplicity, reliability and affordability. and allow the extraction Magnetic Field bump introducted to excite radial oscillation
Graphite Beam Stops to minimize activation of the machine (>50% extraction efficiency)

## Slide 27: Extraction system
Layout (2)
Aluminum support frame
Correction bars
Regenerator
Harmonic coil
Septum
Anti-septum
Graphite beam stop
3 bar corrector
Notes: Harmonic coils to steer the beam precisely into the extraction channel.
The extraction will not be without losses so we determined the distribution of these losses and introduced graphite beam stops in areas where they effectively catch most of the scattered beam.

## Slide 28: External beam line
layout
A variable energy degrader is placed at 2 meters from the yoke exit
A permanent magnet quad matches the beam phase space with respect to the beam line optics
A quadrupole doublet provides a 1 mm double waist (1-sigma) on the degrader.
A variable horizontal collimator between the two quads cuts the horizontal divergence providing constant optics independent of gantry angle
Full assembly can be shifted aside for access to the quads in the shielding wall

## Slide 29: Status after assembly and installation in the
Magnet:
  - Successfully cooled down and ramped up
  - Coil well centered in cryostat
  - Magnet fully mapped, no shimming needed
  - Extraction system fully mapped, no shimming needed
Vacuum system operational
Ion source operational; central region installed
Rotco tested and operational
RF-system installed in cyclotron and tested at full voltage
Cyclotron installed in the shielded bunker
First beam tests started middle of August
Cyclotron fully assembled and installed in shielded beam test vault (August 2013)
The assembled S2C2 in the IBA beam test vault, August 2013

## Slide 30: Extracted beam
First extracted beam observed end of August 2013
A few days after the beginning of commissioning, beam extraction was observed by colour change of a radiation sensitive foil at the beam exit window
After beam injection from the PIG source and 40000 turns of acceleration the beam was immediately ejected by the passive extraction system without any major design or other modifications
Measured beam emittances and beam stability are in full agreement with expectations.
Beam has been transported in the external beam line to the position of the degrader
Shipment of cyclotron to Nice March 2014

## Slide 31: Acknowledgements
A large team of IBA, AIMA and ASG experts have taken part in the design and testing of the S2C2...
Thank you!