## Slide 2: Vacuum Technology for Particle Accelerators JUAS 2013
Paolo Chiggiato
Technology Department
Vacuum Surfaces and Coatings Group
CERN, CH-1211 Geneva
A quick view of vacuum technology
The basis of vacuum technology: pressure, conductance, pumping speed
Calculation of simple pressure profiles
Gas sources: thermal outgassing and beam-induced desorption
Gas pumping: momentum-transfer and capture pumps
Instruments
Seminar: Example of vacuum system design for particle accelerators (by Roberto Kersevan)

## Slide 3
Karl Jousten (Ed.), Handbook of Vacuum Technology Book (Verlag-VCH, Weinheim, 2008)
James M. Lafferty (Ed.), Foundations of Vacuum Science and Technology (John Wiley and Sons, 1998)
Berman, Vacuum Engineering calculation, Formulas, and Solved Exercises (Academic Press, 1992)
P. A. Redhead, J. P. Hobson, E. V. Kornelsen, The Physical Basis of Ultrahigh Vacuum, (Chapman and Hall, 1968)
John F. O'Hanlon, A User's Guide to Vacuum Technology
Handbook of Materials and Techniques for Vacuum Devices (AVS Classics in Vacuum Science and Technology)
A. Berman, Total pressure measurement in vacuum technology, Academic Press, 1985
Selected books

## Slide 4
Part 1
A quick view of vacuum technology

## Slide 5
A quick view of vacuum technology
Beam pipe
Flange
Pump (sputter ion pump)
Vacuum systems are made of several components. The connection between them is obtained by flanges.

## Slide 6
Beam pipe & NEG coating
Flange
Pump (sputter ion pump)
Pressure measurement
Sector valves
cold
Room temp.
A vacuum sector is delimited by gate valves.
The pressure is monitored in each vacuum sector by gauges.
The beam pipes can be at different temperatures
A quick view of vacuum technology

## Slide 7
Beam pipe & NEG coating
Flange
Pump (sputter ion pump)
Gauge
Sector valves
A quick view of vacuum technology

## Slide 8
Sector valves
Example of vacuum components:
pumping ports
bellows
RF shield
flanges
A quick view of vacuum technology

## Slide 9
Sector valves
Example of vacuum components:
Copper beam pipes with stainless steel flanges
A quick view of vacuum technology

## Slide 10
Sector valves
Example of vacuum components:
Beryllium beam pipes in the collision area of the ATLAS experiment (CERN)
A quick view of vacuum technology

## Slide 11
Sector valves
In general the leak-tight connection is made by copper gaskets that are squeezed between two stainless steel knives.
A quick view of vacuum technology

## Slide 12
In general vacuum systems for particle accelerators are made of metals:
Austenitic stainless steel
Copper and its alloys
Aluminum and its alloys
Inconel (a Ni alloy)
Beryllium
In several components the characteristics of the surface is modified by chemical treatment and coating.

## Slide 13
Some example of surface treatments:
Chemical cleaning (mandatory!)
Electropolishing
Passivation
Elecroplating
Thin-film coating by sputtering

## Slide 14
Steps in the production of vacuum systems:
Identification of requirements
Choice of materials and treatments
Design, mechanical evaluation and pressure profile calculation
Machining & assembly (welding, brazing)
Surface treatments
Leak test and global assessment.
Installation and in situ final check.
A quick view of vacuum technology

## Slide 15
Example of vacuum system:
The H- source of the Linac4

## Slide 16
Mechanical engineering for vacuum technology
Buckling mode
Dynamic nodal forces
Principal stress
A quick view of vacuum technology

## Slide 17
Part 2
The basis of vacuum technology

## Slide 18
The basis of vacuum technology: pressure
Definition of pressure:
Unit of measurement: =
Still used in vacuum technology: ;
 | Pa | bar | atm | Torr
1 Pa | 1 | 10-5 | 9.87 10-6 | 7.5 10-3
1 bar | 105 | 1 | 0.987 | 750.06
1 atm | 1.013 105 | 1.013 | 1 | 760
1 Torr | 133.32 | 1.33 10-3 | 1.32 10-3 | 1
I
Conversion Table

## Slide 19
 | Pressure boundaries [mbar] | Pressure boundaries [Pa]
Low Vacuum LV | 1000-1 | 105-102
Medium Vacuum MV | 1-10-3 | 102-10-1
High Vacuum HV | 10-3-10-9 | 10-1-10-7
Ultra High vacuum UHV | 10-9-10-12 | 10-7-10-10
Extreme Vacuum XHV | <10-12 | <10-10
Degree of Vacuum
Pressures and gas quantities are correlated by the gas equation of state. In vacuum the ideal gas law is always fulfilled :
(thermodynamic)
T (statistical mechanics)
P pressure, V volume, T temperature, quantity of gas in moles (n) and number of molecules (N), R gas constant, kB Boltzmann constant
The basis of vacuum technology: pressure

## Slide 20
 | Pressure [Pa] | Gas Density 293 K | Gas Density 4.3K
Atmosphere | 1.013 105 | 2.5 1019 | 1.7 1021
Plasma chambers | 1 | 2.5 1014 | 1.7 1016
LINAC pressure upper limit | 10-5 | 2.5 109 | 1.7 1011
Lowest pressure ever measured at room T | 10-12 | 250 | 1.7 104
The basis of vacuum technology: gas density

## Slide 21
Gas quantities can be expressed in:
Number of molecules: N
Moles:
PV quantities if the temperature is known and constant :
Example: 1 Pa m3 at 293K contains
1 Torr l = 3.3 1019 molecules
1 mbar l = 2.47 1019 molecules
The basis of vacuum technology: gas quantity

## Slide 22
In the kinetic theory of gas the mean speed of a molecule is the mathematical average of the speed distribution:
m is the molecular mass [Kg]
M is the molar mass [Kg]
The basis of vacuum technology: gas kinetics

## Slide 23
Gas |  | 
H2 | 1761 | 213
He | 1244 | 151
CH4 | 622 | 75
N2 | 470 | 57
Ar | 394 | 48
Gas |  | 
H2 | 1761 | 213
He | 1244 | 151
CH4 | 622 | 75
N2 | 470 | 57
Ar | 394 | 48
The basis of vacuum technology: gas kinetics

## Slide 24
Gas | Pressure [mbar] | Impingement rate 293 K [cm-2s-1]
H2 | 10-3 | 1.1 1018
 | 10-8 | 1.1 1013
 | 10-14 | 1.1 108
N2 | 10-3 | 2.9 1017
 | 10-8 | 2.9 1012
Ar | 10-3 | 2.4 1017
 | 10-8 | 2.4 1012
The basis of vacuum technology: impingement rate

## Slide 25
The molecular collision rate w in a gas is:
where is the collision cross section.
For a single gas, in case of elastic collision of solid spheres:
and is the molecular diameter.
The mean free path , i.e. the average distance travelled by a molecule between collisions, is given by:
Gas | 
H2 | 0.27
He | 0.21
N2 | 0.43
O2 | 0.40
CO2 | 0.52
Gas | 
H2 | 0.27
He | 0.21
N2 | 0.43
O2 | 0.40
CO2 | 0.52
The basis of vacuum technology: mean free path

## Slide 26
The basis of vacuum technology: mean free path
Moon-Earth
distance
Earth
diameter
Typical dimension
of standard vacuum
system
Paris-New York
distance
Kn≈1

## Slide 27
The basis of vacuum technology: mean free path
When the mean free path is of the order of typical dimensions of the vacuum vessel, for example the diameter of cylindrical beam pipes, molecular collisions with the wall of the vacuum envelope become preponderant.
For even longer , the gas dynamics is dominated by molecule-wall collisions; intermolecular interactions lose any effect on the gas displacement.

## Slide 28
is the mean free path and D is a characteristic dimension of a vacuum system (p.ex. the diameter of a beam pipe).
Kn range | Regime | Description
Kn >0.5 | Free molecular flow | The gas dynamic is dominated by molecular collisions with the walls of the system
Kn <0.01 | Continuous (viscous) flow | The gas dynamic is dominated by intermolecular collisions
0.5<Kn <0.01 | Transitional flow | Transition between molecular and viscous flow
The basis of vacuum technology: Knudsen number
Except for ion source’s plasma chambers, vacuum systems of accelerators operate in free molecular regime; only this vacuum regime is considered in this lesson.

## Slide 30
The basis of vacuum technology: conductance
In molecular regime the net flux of molecules from one point (1) to another (2) is proportional to the pressure drop:
C is called the gas conductance from the two points. It does not depend on pressure.
It depends only on the mean molecular speed and vacuum system geometry.
If the gas flow units are expressed in terms of pressure-volume, the conductance is reported as volume per unit time, i.e. l s-1 or m3 s-1.

## Slide 31
The basis of vacuum technology: conductance
The simplest example is the gas flow through a small wall slot of area A and infinitesimal thickness
Gas flow 1 2 :
Gas flow 2 1 :
T, P1
T, P2
A
Net flow:
In PV units (

## Slide 32
The basis of vacuum technology: conductance
The conductance of the wall slots is inversely proportional to the square root of the molecular mass.
Therefore, for equal pressure drop the gas flow of H2 is the highest.
For gas molecules of different masses, the conductance scales as the square root of the inverse mass ratio:

## Slide 33
This table collects conductance values, for an orifice, per unit surface area (C’) at room temperature for common gas species.
The basis of vacuum technology: conductance
Conductance of a wall aperture in PV units, per unit area:
T= 293 K
Gas |  |  | 
H2 | 1761 | 440.25 | 44
He | 1244 | 311 | 31.1
CH4 | 622 | 155.5 | 15.5
H2O | 587 | 146.7 | 14.7
N2 | 470 | 11.75 | 11.75
Ar | 394 | 98.5 | 9.85
Gas |  |  | 
H2 | 1761 | 440.25 | 44
He | 1244 | 311 | 31.1
CH4 | 622 | 155.5 | 15.5
H2O | 587 | 146.7 | 14.7
N2 | 470 | 11.75 | 11.75
Ar | 394 | 98.5 | 9.85

## Slide 34
Example:
P1
P2
A
The basis of vacuum technology: conductance
If the gas was N2, the flow would have been times smaller

## Slide 35
For more complex gas flow restrictions, the transmission probability t is introduced.
P1
P2
A1
A2
Vessel 1
Vessel 2
Gas flow 1 2 :
Gas flow 2 1:
In PV units:
The basis of vacuum technology: conductance
In absence of net flow: and P1=P2

## Slide 36
The molecular transmission probabilities:
Depend only on the geometry of the vacuum system;
They are accurately obtained by Monte Carlo methods.
They can be calculated analytically only for simple geometry;
The basis of vacuum technology: conductance

## Slide 37
Example:
Tubes of uniform circular cross section ( length, R radius); Santeler formula (max error 0.7%):
For long tubes ():
For N2 and
The basis of vacuum technology: conductance

## Slide 38
The basis of vacuum technology: conductance
Monte Carlo

## Slide 39
roberto.kersevan@cern.ch
Other analytical expressions for the transmission probability can be found for ducts of circular, rectangular and elliptical cross section (see for example J. M. Lafferty, Foundation of Vacuum Science and Technology, Wiley Interscience).
For more complicated geometry, Test-Particle Monte Carlo methods (TPMC) are used.
The basis of vacuum technology: conductance

## Slide 40
The TPMC:
generates “random” molecules according to the cosine distribution at the entrance of the tube;
then follows their traces until they reach the exit of the tube
each collision with the walls is followed by a random emission
many simulated molecules are needed in order to reduce the statistical scattering
The software used at CERN is Molflow+, written and maintained by Roberto Kersevan and Marton Szakacs (CERN, TE-VSC)
The basis of vacuum technology: conductance

## Slide 41
P1
P2
P3
C1
C2
It can be easily verified that: and :
In general for N vacuum components traversed by the same gas flux, i.e. placed in series :
In stable conditions, there is no gas accumulation in the whole system:==
The basis of vacuum technology: conductance
Conductance of components connected in series:

## Slide 42
P1
P2
C1
C2
=+
For components connected in parallel (same pressures at the extremities):
The basis of vacuum technology: conductance

## Slide 43
In vacuum technology a pump is any ‘object’ that remove gas molecules from the gas phase.
The pumping speed S of a pump is defined as the ratio between the pump throughput QP (flow of gas definitively removed) and the pressure P at the entrance of the pump:
vacuum vessel
Ap
pump
aperture
The basis of vacuum technology: pumping speed

## Slide 44
The gas removal rate can be written as:
is the area of the pump aperture
: is the conductance of the unit surface area
n: the gas density
: the capture probability, i.e. the probability that a molecule entering the pump is definitively captured
As usual, in term of pressure and PV units:
From the definition of pumping speed:
The basis of vacuum technology: pumping speed

## Slide 45
vacuum vessel
Ap
pump
aperture
ID [mm] | H2 | N2 | Ar
36 | 448 | 120 | 100
63 | 1371 | 367 | 307
100 | 3456 | 924 | 773
150 | 7775 | 2079 | 1739
Maximum pumping speed [l s-1]for different circular pump apertures
The maximum pumping speed is obtained for and is equal to the conductance of the pump aperture.
S depends on the conductance of the pump aperture and the capture probability s
is in general not a constant; it may depend on many parameters including pressure, kind of gas and quantity of gas already pumped.
The basis of vacuum technology: pumping speed

## Slide 46
A gas flow restriction interposed between a pump and a vacuum vessel reduces the ‘useful’ pumping speed.
The effective pumping speed Seff seen by the vacuum vessel is easily calculated:
vacuum vessel
P1
pump
aperture
S
P2
C
vessel
aperture
For C<<S:
The basis of vacuum technology: pumping speed

## Slide 47
Example:
Vessel and pump connected by a 10 cm diameter tube; N2, S=250 l/s and 1000 l/s.
The basis of vacuum technology: pumping speed
Chart: 
| 1 | 2 | 3 | 4 | 4.5 | 5 | 10 | 15 | 20 | 25 | 30
Conductance | 471.24 | 332.64 | 254.4696 | 207.9 | 190.9363 | 176.4295 | 100.716 | 71.148 | 54.8856 | 44.766 | 37.8036
S=250 l/s | 163.3437 | 142.7296 | 126.1075 | 113.5073 | 108.2562 | 103.4341 | 71.7931 | 55.3857 | 45.0051 | 37.9674 | 32.838
S=1000 l/s | 320.3012 | 249.6098 | 202.8504 | 172.1169 | 160.3245 | 149.9703 | 91.5004 | 66.4222 | 52.0299 | 42.8478 | 36.4266
Chart: 
| 1 | 2 | 3 | 4 | 4.5 | 5 | 10 | 15 | 20 | 25 | 30
Conductance | 471.24 | 332.64 | 254.4696 | 207.9 | 190.9363 | 176.4295 | 100.716 | 71.148 | 54.8856 | 44.766 | 37.8036
S=250 l/s | 163.3437 | 142.7296 | 126.1075 | 113.5073 | 108.2562 | 103.4341 | 71.7931 | 55.3857 | 45.0051 | 37.9674 | 32.838
S=1000 l/s | 320.3012 | 249.6098 | 202.8504 | 172.1169 | 160.3245 | 149.9703 | 91.5004 | 66.4222 | 52.0299 | 42.8478 | 36.4266
Seff and C [l/s]
vacuum vessel
P1
pump
aperture
S
P2
C
vessel
aperture

## Slide 48
Efficient pumping?

## Slide 49
Part 3
Calculation of simple pressure profiles

## Slide 50
P
S
Q
P
S
C
Q
Simple pressure profiles: stationary values
P0 is the pressure attained in the vacuum system without any gas load. It is an
intrinsic limitation of the pump.

## Slide 51
P1
S1
P2
S2
P3
S3
P4
S4
Q
C1
C2
C3
Flux balance at the connexions (node analysis):
+
S
Cx
Q
x
L
Simple pressure profiles: stationary values

## Slide 52
Pressure profiles with distributed gas load can be calculated analytically (for simple geometry), by electrical analogy or by Monte Carlo simulation.
P(0)
S
P(x)
x
x
L
Distributed gas load, one pump
at one extremity only
x
x+Dx
Distributed outgassing
Simple pressure profiles: stationary values

## Slide 53
P(0)
S
P(x)
x
x
L
Distributed outgassing
P(0)
S
PMAX
Distributed gas load, one pumps
at each extremity
Simple pressure profiles: stationary values

## Slide 54
t
From the ideal gas equation:
A gas balance equation can be written as:
;
in PV units:
The pumped gas rate is:
P
S
Qin
Qout
Simple pressure profiles: time dependent values
In time dependent condition:
Therefore, the variation of the gas density in the system has to be taken into account.

## Slide 55
t
P
tp
P
S
Qin
Qout
characteristic time of pumping
A = integration constant
Simple pressure profiles: time dependent values
P0
tp characterizes all pressure transients in a vacuum system

## Slide 56
When Qin is a function of time:
t
Q
t
P
A = integration constant
For a network of vacuum chambers, systems of coupled differential equations for each chamber have to be solved.
However, a simpler method exists. It is based on the analogy between vacuum systems and electrical networks. Very powerful software is available for the time dependent analysis of electrical networks.
Simple pressure profiles: time dependent values
P
S
Qin
Qout

## Slide 57
Vacuum element | Electrical elements
Conductance C | Conductance 1/R
Gas Flow Q | Current I
Pressure P | Voltage V
Volume V | Capacitance C
Pump | Conductance to ground
Gas source | Current generator
Constant pressure source | Voltage supply
Vacuum chamber with conductance and volume | 
The ground potential is equivalent to zero pressure.
Long tubes are subdivided in smaller units and considered as single vacuum chambers (conductance + volume) in series.
The conductance of a single small unit is equal to the conductance of the entire vacuum chamber times the number of units
Non-linear electric characteristics can be used to simulate pressure and time dependent conductances and pumping speeds.
Simple pressure profiles: time dependent values

## Slide 58
Electrical analogy of a vacuum chamber with gas injection and pumping
Electrical analogy of a vacuum system with differential pumping
Simple pressure profiles: time dependent values
S
Q
Q
P1
P2
C
V1
V2
S1
S2
The electrical network is solved by dedicated software, for example LTSpice

## Slide 59
Simple pressure profiles: time dependent values
Example 1: the Linac4 ion source

## Slide 60
Example 1: the Linac4 ion source
Simple pressure profiles: time dependent values

## Slide 61
Simple pressure profiles: time dependent values
Example 2: the MedAustron facility

## Slide 62
Simple pressure profiles: time dependent values
Example 2: the MedAustron facility

## Slide 63
Equivalent electrical network for sector 560.
Lay-out of the SPS vacuum sector 560.
Simple pressure profiles: time dependent values
Example 3: pumpdown of one vacuum sector of the SPS

## Slide 64
Comparison between calculated pumpdown curves for original (green) and halved (orange) longest SPS vacuum sector.
Current dependent resistances simulate conductance variations in the whole pressure range, even in viscous regime.
Simple pressure profiles: time dependent values
Example 3: pumpdown of one vacuum sector of the SPS

## Slide 65
Part 4
Gas sources in particle accelerators

## Slide 66
Outgassing is the spontaneous evolution of gas from solid or liquid.
Degassing is the deliberate removal of gas from a solid or a liquid.
Desorption is the release of adsorbed chemical species from the surface of a solid or liquid.
Gas sources

## Slide 67
Solvents: their molecules interact and transport contaminants away by diffusion (dilution) -> quite selective! (C2Cl4, wide spectrum; HFC, more restricted action)
Detergents in water: allows organics and water to combine by forming micelle (surfactant: surface acting agent ). Based on molecule with hydrophilic heads and lipophilic tail: less selective than solvents
Gas sources
After production, the surface of vacuum components is always contaminated.
They must be thoroughly cleaned before installation.

## Slide 68
The bulk of materials is also a source of gas.
Gas molecules dissolved in the materials diffuse towards the surfaces and are then released
Polymers dissolve huge quantity of molecules in the voids between the molecular chains.
Metals can dissolve only limited quantities of small atoms that are immobile in the lattice at room temperature, except for hydrogen.
In one day, H atoms travel in average 4 mm in austenitic stainless steels, while O atoms travel the same distance in 1000 years.
Therefore, amongst the dissolved elements in the bulk of metals, only H2 is released at room temperature.
For comparison, in one day, H2O molecules move along about 20 mm in PEEK, a high performance polymer.
Gas sources: outgassing

## Slide 69
Materials | Characteristics that determine outgassing rates
Polymers | Dissolve entire molecules, for example H2O, O2, CO2, solvents, etc.
 | The solubility in the bulk can be very high in particular for H2O (up to a few weight percentage)
 | The dissolved molecules have a relatively high mobility
Metals | Only single atoms are dissolved (H, O, C, S, etc.)
 | The atom solubility is in general very low at room temperature (typically up to 103 at. ppm)
 | Only H has a significant mobility at room temperature
Gas sources: outgassing

## Slide 70
Metals
After state-of-art surface cleaning
If not heated in situ: mainly H2O for the first months in vacuum, then also H2.
The source of H2O is recharged after each venting to air.
If heated in situ (baked-out): mainly H2. The outgassing rate can be assumed as constant; it depends on the accumulated effect of the previous thermal treatments
Organics (Polymers)
High solubility of gas in the bulk, in particular H2O.
In general, the outgassing process is dominated by H2O release.
In the initial phase of pumping:
Heavier gas molecules can be outgassed (remnant of polymerisation, fraction of polymeric chains)
The permeation of light molecules is not negligible, in particular He.
Gas sources: outgassing

## Slide 71
Neoprene (10 h pumping):
qH2O ≈ 10-5 Torr l s-1 cm-2
qH2O=3.3x1014 molecules cm-2
Unbaked stainless steel (10 h pumping):
qH2O=2x10-10 Torr l s-1 cm-2
qH2O=6.6x109 molecules cm-2
Baked stainless steel (150º C x 24 h):
qH2=2x10-12 Torr l s-1 cm-2
qH2=6.6x107 molecules s-1 cm-2
Baked OFS Copper (200º C x 24 h):
qH2=2x10-14 Torr l s-1 cm-2
qH2=6.6x105 molecules s-1 cm-2
Gas sources: outgassing
Some examples of outgassing rates

## Slide 72
H.F. Dylla, D. M. Manos, P.H. LaMarche
Jr. J. Vac. Sci. and Tech. A, 11(1993)2623
Outgassing rate of water vapour for unbaked metals
Gas sources: outgassing of metals
Water vapour outgassing rate of stainless steel that underwent four different surface treatments
The most effective way to accelerate the water vapour outgassing rate is by in situ bakeout.
At least 120°C for 12 hours

## Slide 73
Gas sources: outgassing of metals
When water vapour outgassing is strongly reduced, by either long pumping or bakeout, the outgassing process is led by H2.
This gas is dissolved in metals as single H atoms. Its diffusion is relatively fast and, after recombination on the surface, it can be released as molecular hydrogen
Most of the H atoms are dissolved in metals at the liquid state, during the production process.
H atoms mobility and solubility in the liquid are higher than in the solid.
Typical sources of H are:
metals ores;
tools needed for fusion;
refractory materials of furnaces;
combustion and treatment gas;
water vapour and fluids used for quenching (for example the hyperquench of austenitic stainless steels is carried out from 1100°C in water, air, or oil).

## Slide 74
Gas sources: outgassing of metals
Example of a process of H2 dissolution
in liquid Al
If the solidification is fast, the gas is
trapped in the solid far form the
equilibrium
Typical H2 contents are about 1 wt. ppm for
copper, aluminum, and austenitic stainless steel

## Slide 75
Gas sources: outgassing of metals
As for water vapour, hydrogen-outgassing rate is reduced by heating.
The high temperatures increase the H atoms mobility and, as a result, accelerate the depletion of the residual hydrogen content.
However, there is a crucial difference between water vapour and hydrogen.
Each time the vacuum system is exposed to air, water molecules re-adsorb on the surface, while hydrogen is not recharged in the bulk of the metal.
For most of the materials used for the manufacturing of vacuum chambers, the H solubility is very low in the solid state.
For example, to recharge 1 wt. ppm of hydrogen at room temperature in stainless steel, the material has to be in equilibrium with the gas at 7 bar. The hydrogen pressure in air is roughly 10-4 mbar, which gives a maximum recharging of about 2.10-4 wt. ppm.

## Slide 76
For copper and aluminium alloys, a few bakeout at 150-200°C for 24 hours are sufficient to reduce the hydrogen-outgassing rate to less than 10-13 mbar l s-1 cm-2.
For austenitic stainless steel, higher temperatures are needed to have a similar effect for a few mm thick vacuum chambers. Repeated bakeout at temperature higher than 200°C may have a significant influence.
Gas sources: outgassing of metals
Materials | Bakeout T[°C] x 24 h | q [mbar l s-1 cm-2]
Austenitic st. steel | 150 | 3 10-12
Austenitic st. steel | 200 | 2 10-12
Austenitic st. steel | 300 | 5 10-13
Copper Silver added (OFS) | 150 | 3 10-12
Copper Silver added (OFS) | 200 | ≈ 10-14
Beryllium | 150 | < 10-14
Al alloys | 150 | < 10-13

## Slide 77
For austenitic stainless steels, a radical effect is obtained by heating in a vacuum furnace to temperature up to about 1000°C.
Such a treatment is called ‘vacuum firing’. At CERN it is carried out at 950°C for 2 h in a vacuum better than 10-5 mbar at the highest temperature.
Gas sources: outgassing of metals
The CERN’s large furnace: useful height and diameter: 6 m and 1 m, respectively. Maximum charge weight: 1000 Kg. Ultimate pressure: about 10-7 mbar; pressure at the end of the 950°Cx2h treatments: 10-5mbar.

## Slide 78
Diffusion model of H2 outgassing
Eb
Diffusion, in most of the cases of interest, is described by the Fick’s equations:
where c(x,t) is the concentration in the solid and G is the flow of molecules per cm2
In the limit of this model, the outgassing rate is equal to a half of the flux of atoms arriving at the surface by diffusion (2 H atoms = 1 H2 molecule):
Gas sources: outgassing of metals

## Slide 79
co
c(x,t)
x
diffusion length
The t-0.5 evolution holds also for solid of finite dimension L (slab) when L>>(Dt)0.5
The total amount of substance M(t) which has left the solid is:
Diffusion model of H2 outgassing: semi-infinite solid approximation
Gas sources: outgassing of metals

## Slide 80
Diffusion model of H2 outgassing: slab approximation
Gas sources: outgassing of metals
L
c(x,t)
x
For Dt>0.05 L2 only the first term of the series is relevant :

## Slide 81
Diffusion model of H2 outgassing: Sievert’s law
Gas sources: outgassing of metals
If hydrogen in the gas phase and on the surface are in equilibrium, the surface concentration cw may be calculated by the Sievert’s law:
For stainless steel, Es is positive; therefore, the solubility increases with temperature. At room temperature and for typical hydrogen pressure, cw is negligible with respect to the residual hydrogen content. This is not true at higher temperature; for example during a vacuum firing treatment.
PH2 is the pressure of hydrogen in the gas phase
Ks(T) is the hydrogen solubility in the metal
Es is the H2 solubility energy
K0 a constant.
Sievert’s law for austenitic stainless steel

## Slide 82
Diffusion model of H2 outgassing: Fourier number
Gas sources: outgassing of metals
If the slab is heated to a temperature TH for a time tH, the outgassing rate, when back at room temperature is:
T
t
TH - tH
TRT
For an arbitrary thermal cycle T=T(t), it can be show that:
The dimensionless number:
is called the Fourier number

## Slide 83
Fo ‘records’ the thermal history of the material and determines how much of the initial concentration is depleted.
When thermal treatment of Fo≈ 3 are applied, the solid is actually emptied or in equilibrium with the surrounding gas phase.
Diffusion model of H2 outgassing: consecutive bakeouts
Gas sources: outgassing of metals
In case of multiple identical thermal treatments (for example bakeouts) with a given constant F0, the calculated ratio of the outgassing rate at room temperature after two successive thermal cycles is:
That is to say that each bakeout reduces the outgassing rete of hydrogen by the same constant factor. This is in line with the experimental results.

## Slide 84
1.00E-16
1.00E-15
1.00E-14
1.00E-13
1.00E-12
1
2
3
4
5
Bakeout cycles
bakeout at 300°C, 24 h
H2 outgassing rate [Torr l s-1 cm-2]
Diffusion model of H2 outgassing: consecutive bakeouts
Gas sources: outgassing of metals

## Slide 85
Parameters | Symbol | Values
Temperature of the firing treatment | Tf | 950°C
Duration of the firing treatment | tf | 2 hours
In situ bakeout temperature | Tbo | 150°C
Duration of the in situ bakeout | tbo | 24 hours
Initial content of residual hydrogen | co | 1 ppm wt (≈50 ppm at.)
Hydrogen equilibrium concentration on slab surfaces during firing | cw | 0.06 ppm wt. Equivalent to PH2=1.3x10-5 mbar
Diffusion model of H2 outgassing: vacuum firing
Gas sources: outgassing of metals
semi-infinite model
For zero pressure in the furnace
10-5 Torr pressure in the furnace
q [Torr l s-1 cm-2]
L [cm]

## Slide 86
Gas sources: outgassing of polymers
Polymers, in particular their amorphous structures, can dissolve huge quantities of gas.
The water vapour solubility is very high; for example for common materials like Viton, PEEK, and Vespel the content of water in equilibrium with 50% humidity air at 20°C is 0.21, 0.2, and 1 wt. %, respectively.
The huge quantity of dissolved gas and the relatively high mobility through the polymeric chains result in much higher outgassing rates than the ones for metals. This is particularly marked for water vapour.
The maximum bakeout temperature depend on the type of polymer; it is limited to about 200 °C for Viton®.

## Slide 87
Gas sources: outgassing of polymers
The outgassing rates of thick ( L>Dt) slab of polymers, as expected by the diffusion model, decrease with the inverse of the square root of the pumping time t:
Compared to unbaked metals, thick polymers experience a much slower reduction of water outgassing rate.

## Slide 88
R. N. Peacock, J. Vac. Sci. Technol., 17(1), p.330, 1980
baked Viton
unbaked Viton
10-10
10-6
Outgassing rate of water vapour for polymers
Gas sources: outgassing of polymers

## Slide 89
Another important limitation of polymers used as seals is the high gas permeability.
Gas penetrates into the material and diffuses towards the vacuum system. The permeation flow may limit the ultimate water vapour pressure in vacuum systems and affect the sensitivity of helium leak-detection.
As an example, the permeation flow of atmospheric water through a Viton O-ring, 5 mm cross section diameter, 6 cm torus diameter, is about 10-7 mbar l s-1 This value is equivalent to the outgassing of 1 m2 of stainless steel after 100 h of pumping.
Gas sources: outgassing of polymers

## Slide 90
Part 3: Outgassing of polymers
Since the permeability of metals is negligible for all gases, metallic coatings should entirely block the polymer outgassing and permeation.
However, experimental results show that only a partial reduction of the flux is attained.
This is attributed to defects on the coating (pinholes or scratches) that cause discontinuity on the surface coverage.
Pinholes are produced during the deposition process and they are presumably due to atmospheric dust particles.
Application of polymers in UHV technology is hindered by the huge outgassing and permeability.
This drawback has been partially overcome in other technological domains (i.e. packaging) by coating polymers with metals or metal oxides (more recently also with a:C-H).
How effective are coatings in reducing outgassing and permeation rate?
Gas sources: outgassing of polymers

## Slide 91
Part 3: Outgassing of polymers
2.00 mm
20.0 mm
Examples of pinholes on Al coating deposited on PEEK
Gas sources: outgassing of polymers

## Slide 92
Normalised uncoated surface areas of the order of 10-4 are measured
Transmitted light optical microscopy has been used to detect uncoated surfaces on metal films, showing that they mainly consist of pinholes of different diameters.
Part 3: Outgassing of polymers
Al/PEEK
Gas sources: outgassing of polymers

## Slide 93
Part 3: Outgassing of polymers
However, experiments have shown that the permeability is reduced much less than expected from the uncoated fraction
Example:
Al/PEEK(0.125 mm thick)
The uncoated fraction is 10-4, but the permeability of He is 10-2 of that of the uncoated polymer.
Gas sources: outgassing of polymers

## Slide 94
r =Normalised permeability  order of 10-2
Q =Normalised uncoated area  order of 10-4
This apparent inconsistency can be justified considering that the pinhole gas throughput is enhanced by lateral diffusion.
L
R0
1.18
Amplification factor
* After W.Prins and J.J.Hermans, J.Phys.Chem., 63 (1959) 716.
For L/R0 >0.3 :
*
Part 3: Outgassing of polymers
Gas sources: outgassing of polymers

## Slide 95
Barrier improvement factor:
Part 3: Outgassing of polymers
In the literature the barrier efficiency is called “barrier improvement factor” or BIF. It is the inverse of normalized permeability.
The fact that r depends on L means that for L>10R the flux value does not depend on the polymer thickness. In other terms, all the concentration gradient is localized near the pin-hole.
It follows also that, for similar coating, the improvement due to the coating is less significant for thicker polymer substrates.
Normalized permeability:
Gas sources: outgassing of polymers

## Slide 96
Part 3: Outgassing of polymers
The BIF for Al coated PEEK.
For thickness larger than 5 mm, the coating is useless.
Gas sources: outgassing of polymers
In conclusion, metallic coatings are very efficient in the reduction of the permeation and outgassing rate in thin polymers. For polymers thicker than 5 mm, the benefit of the coating is negligible.

## Slide 97
The impingement of electrons and photons, of energy higher than some eV, on solid surfaces can result in ions and neutral molecules desorption.
The first observations of induced desorption are very old and they go back to the origin of modern vacuum technology.
In 1918, Dempster observed ion desorption from electron bombarded salts (Phys. Rev. 11, 323)
Gas sources: beam induced desorption
ESD: electron stimulated desorption

## Slide 98
Millikan reported the first evidence of photon induced desorption in 1909 during the measurement of the photoelectric current of metals exposed to ultraviolet radiation. The first interpretation is given by Winch in 1930 (Phys. Rev. 36, 601).
Winch was the first to see the implication of photoelectrons on photon induced desorption.
Gas sources: beam induced desorption
PSD: photon stimulated desorption

## Slide 99
Gas sources: beam induced desorption
The desorption yield h, i.e. the number of molecules desorbed per impinging particle, is needed to design the vacuum system of particle accelerators.
depends on many parameters, in particular:
on the nature and energy of the impinging particle;
the material of the vacuum chamber;
the nature of the desorbed gas;
quantity of particles that have already impinged on the surface, namely the dose D .
The cleanliness of the surfaces has also a crucial influence.

## Slide 100
Typical values of h for baked structural vacuum materials vary between 10-1 and 10-3 molecules per electron at 500 eV. Neutral desorption is much more likely than ion desorption, about 100 times higher.
0
200
400
600
800
1000
1200
h
Electron Energy [eV]
500 eV
Gas sources: electron induced desorption
For typical metals used in vacuum technology, h of hydrogen is the highest followed by those for CO, CO2 and CH4, in decreasing order.

## Slide 101
where a is between 0.6 and 1.
Ee=300 eV
J. Gomez-Goñi and A. G. Mathewson J. Vac. Sci. Technol. A 15, 3093 (1997)
a=0.77
The samples were baked at 150 C for 24 h and at 300 C for 2 h
Gas sources: electron induced desorption
Electron bombardment cleans metallic surfaces by gas desorption.
As a consequence, the desorption yields decrease when the electron dose increases.
he values have a power law dependence on the dose of electrons De (for doses higher than 1018 e- cm-2):

## Slide 102
J. Gomez-Goñi and A. G. Mathewson J. Vac. Sci. Technol. A 15, 3093 (1997)
Ee=300 eV
a=0.62
a=0.68
Gas sources: electron induced desorption

## Slide 103
J. Gomez-Goñi and A. G. Mathewson J. Vac. Sci. Technol. A 15, 3093 (1997)
1ML=2x1015 molecules/ cm2
The total quantity of gas desorbed exceed by far one monolayer.
In order to explain the huge quantity of gas desorbed, some authors considered stimulated gas diffusion from the oxide layer
Gas sources: electron induced desorption
The accumulated dose effect is lost when the surface is vented to air for long time. For short exposure to air (for example a few hours) a residual effect is measured; it is called ‘memory effect’: h values similar to those attained before venting are achieved after a smaller dose.

## Slide 104
For H2 the fraction of energy transferred to the molecules is about . For typical electron energy of about 300 eV, DE is of the order of 0.1 eV.
These values are much lower than the measured energies of desorbed molecules, typically in the 2-10 eV range. Therefore, the direct momentum transfer is not the dominant process in ESD. Electronic energy transfers are then considered.
Gas sources: electron induced desorption
The ESD process can be considered as an isolated interaction between an electron and an adsorbate.
The maximum energy DE transferred during the collision between a low energy electron of mass me and a molecule of mass M is calculated by classical kinematics. For hard-sphere scattering and me<<M, one can estimate that:

## Slide 105
One of the earliest and general models to explain ESD from surfaces was introduced independently by Menzel, Gomer, and Redhead (MGR) in 1964.
The interaction with the incident electron provokes an adiabatic transition (Franck-Condon principle) to exited states of different nature:
Gas sources: electron induced desorption
The adiabatic transition implies a change of electronic energy in about 10-16 s with a frozen distance z between nuclei. The transition is vertical in potential energy vs. distance diagram.
antibonding
metastable
ionic state

## Slide 106
After the excitation, nuclear motion may occur over a time scale of about 10-13-10-14 s: part of the potential energy is converted into kinetic energy.
The crossing of potential curves is possible resulting in different de-excitation pathways.
Gas sources: electron induced desorption

## Slide 107
Whenever the original state is reached during the downward transition (rebonding), the molecule can escape if it has acquired enough kinetic energy to overpass the potential binding energy at the position of transition.
The desorption as a stable neutral is possible when the downward transition occurs at a distance from the surface which is larger than a critical distance rc .
Gas sources: electron induced desorption

## Slide 108
Gas sources: photon induced desorption
Particle beams emit photons whenever they are accelerated, namely whenever their velocity vectors are changed.
Synchrotron radiation (SR) is emitted when charged particles moving with relativistic speeds are forced to follow curved trajectories in magnetic fields.
The emitted photons impinge on the wall of the vacuum system and may desorb gas molecules.
There is experimental evidence that the photon stimulated desorption is generated by a two-step process:
The impinging photons extract photoelectrons.
The extraction and recapture of the photoelectrons stimulate
the desorption of gas molecules by the ESD mechanism.

## Slide 109
Gas sources: photon induced desorption
A particle of charge ‘e’, energy E and rest mass m0, moving on a circular orbit (radius r) radiates electromagnetic radiation with the following power Prad:
where e0 and c are the vacuum permittivity and the speed of light, respectively.
The emitted power depends strongly on the beam energy and the radius of the bent trajectory.
Consequently, electrons emit much more synchrotron radiation power than protons for the same bending radius and energy:

## Slide 110
Gas sources: photon induced desorption
Synchrotron emission is strongly beamed along the direction of motion, which is perpendicular to the acceleration.
The emission is concentrated into an angle of the order of 2/g rad along the direction of motion:

## Slide 111
Gas sources: photon induced desorption
Synchrotron radiation from bending magnets has a very broad energy spectrum, which is characterised by the critical energy eC:
The critical energy subdivide the photon spectrum in two parts of equal emitted power.
For electrons
For protons

## Slide 112
Gas sources: photon induced desorption
Courtesy of Roberto Kersevan

## Slide 113
Gas sources: photon induced desorption
Courtesy of Roberto Kersevan

## Slide 114
Gas sources: photon induced desorption
It can be shown that the number of emitted photon per unit time:
The linear flux, i.e. the emitted photons per second and meter, is given by
Mean photon energy: max of the distribution
For electrons
For electrons

## Slide 115
J. Gomez-Goñi, O. Groebner and A. G. Mathewson J. Vac. Sci. Technol. A 12, 1714 (1994)
Photons with energy lower than the photoelectron extraction threshold (workfunction, about 5 eV for metals used in UHV) do not stimulate gas desorption.
hph depends linearly on the critical energy of the photon spectrum when ec < 280 eV:
Gas sources: photon induced desorption
0.74< b <1.12

## Slide 116
Gas sources: photon induced desorption
hph does not depend significantly on the photon dose D up to about 1020 photons/m. For higher doses, hph varies as a power law function of the dose:
Photon desorption yields for baked stainless steel as a function of the accumulated dose of photons.

## Slide 117
unbaked chamber
Gas sources: photon induced desorption

## Slide 118
Methods for the reduction of electron and photon induced desorption
Gas sources: e- & ph induced desorption
The gas desorbed by electrons or photons is extracted from the surface of vacuum chambers or from the first tenths of atomic layers underneath.
Therefore, a state of art surface cleaning is essential to avoid excessive desorption.
A beneficial effect is obtained by in situ bakeout at temperature higher than 120°C; heating the material results in a cleaner surface; however, the effect on the h values is limited and, in general, less than a factor of 10.

## Slide 119
Electropolishing and vacuum firing have a very limited effect on the h values.
Air exposure after these treatments oxidizes again the surface and regenerates the gas source.
Methods for the reduction of electron and photon induced desorption
Gas sources: e- & ph induced desorption
Electron energy: 500 eV

## Slide 120
Electropolishing does not influence the desorption yield of H2. Very small variations are recorded for the other gasses.
Electron energy: 500 eV
Part 8: Induced desorption
Gas sources: e- & ph induced desorption

## Slide 121
The accumulated dose effect lead to a significant reduction of the desorption yields. Power law reductions are measured for both accumulated electron and photon impingements.
Gas sources: e- & ph induced desorption
Particle accelerators, whose performance is limited by beam induced desorption, are run at progressively increasing current to accumulate photon or electron dose.
This operation is called ‘beam conditioning’ or ‘beam scrubbing’ of the vacuum chambers. It can last several hours to several days.
LEP

## Slide 122
Gas sources: e- & ph induced desorption
Another radical method to reduce stimulated desorption is coating the inner wall of the beam pipes with a thin non-evaporable getter film.
After in situ bakeout at temperature higher than 180°C, the oxide layer covering the film is dissolved leaving an extremely clean surface.
This treatment is considered in the chapter dedicated to pumps.

## Slide 123
Gas sources: ion induced desorption
There are two sources of energetic ions:
Beam particles may collide with the residual gas and create ions (H2+, CH4+, CO+, CO2+, etc.). If the beam is positively charged, the ions are accelerated by the beam potential and collide on the nearby walls with energy between 1 eV to several KeV. The collision results in gas desorption.
When heavy ions are accelerated, beam losses may lead to collision of high-energy heavy ions with the wall of the vacuum system. Typical ions are Pb53+, U73+, and Ar10+. Experimental studies have been carried out with beam energy in the range from 1 MeV/nucleon to 100 GeV/nucleon. In this case, the desorption yields may be orders of magnitude higher than those for residual-gas ions.

## Slide 124
Ions are more effective in desorbing gas than electrons and photons. Typical hi values for baked copper and ion energy of about 1 KeV are about 1 for H2 and CO; 5 and 10 times lower for CO2 and CH4, respectively.
Gas sources: ion induced desorption
Residual-gas ion

## Slide 125
In the last 15 years, several experiments have shown that high energy (E > 1 MeV/nucleon) ions can induce the desorption of a huge quantity of gas molecules.
Desorption yield up to 105 molecules/ion were measured for In49+ at 158 GeV/nucleon at the CERN SPS.
Gas sources: ion induced desorption
High-energy heavy-ions

## Slide 126
Gas sources: ion induced desorption
High-energy heavy-ions
The accumulated dose has a beneficial effect also on the ion induced desorption
Pb 53+ ions
Grazing incidence:89.2°
4.2 MeV/u

## Slide 127
Ion induces desorption can trigger a rapid pressure rise in accelerators where positive beams circulate.
This phenomenon was shown first at the CERN ISR when increasing the proton current to about 1 A
Gas sources: ion induced desorption

## Slide 128
Gas sources: ion induced desorption
The residual gas is ionised by the positive-particle beam
The ionised molecules are accelerated by the beam potential
The ions collide on the surface and extract gas
The pressure rise is generated by a positive feedback process that can be depicted with the following block diagram.

## Slide 129
The critical current at which the runaway process starts can be obtained in the following way:
The ionisation probability for one proton is and the flow of circulating protons is . Therefore, the total gas flow of desorbed molecules per unit length of vacuum chamber is:
The pressure in the beam pipe is:
The pressure diverges when:
Gas sources: ion induced desorption

## Slide 130
Gas sources: ion induced desorption
Critical current:
Gas | si [10-18 cm2] 26 GeV | si [10-18 cm2] 7 TeV
H2 | 0.22 | 0.37
He | 0.23 | 0.38
CH4 | 1.2 | 2.1
CO | 1.0 | 1.8
CO2 | 1.1 | 2.0
Ar | 1.6 | 2.8
Example:
For an effective pumping speed of 10 l s-1 and an ionisation cross section of 10-18 cm2 (CO ionised by protons at 26 GeV), the critical current is about 16 mA if the desorption yield is 1.

## Slide 131
Part 5
Gas pumping in particle accelerators

## Slide 132
In molecular regime:
gas molecules cannot be removed by suction: the molecules do not transfer energy and momentum amongst them; pumps act on each molecule singularly;
pumps are classified in two families:
        momentum transfer pumps;
        capture pumps.
In the first, molecules receive a momentum component pointing towards the pump outlet (foreline) where the gas is compressed and evacuated by pumps working in the viscous regime (for example: rotary vane, diaphragm, and scroll pumps).
Capture pumps remove molecules from the gas phase by fixing them onto an internal wall.
Gas pumping

## Slide 133
In molecular pumps, molecules are headed for the outlet by moving surfaces.
Molecules impinge and adsorb on the moving surface; on desorption, the velocity distribution is superimposed to the drift velocity of the wall
As a result, molecules are preferentially redirected towards the direction of the wall movement, i.e. the density of molecules increases in the same direction.
A moving wall produces a gas flow.
Gas pumping: momentum transfer
Molecular pumps
u≠0

## Slide 134
The most important characteristics of molecular pumps are:
    Pumping speed S
    Maximum compression ratio MAX
The parameters affecting S and K0 can be identified by a simple model.
Gas pumping: momentum transfer
Molecular pumps
The first molecular pump was invented by W. Gaede. In the original design, the moving surface is a rotor revolving at high frequency
To prevent backstreaming, the inlet and the outlet were separated by a very thin slot (locking slot). The slot height was of the order of 10-2 mm.

## Slide 135
Gas pumping: momentum transfer
Molecular pumps
At any time, half of the molecules have just collided with the moving surface (rotor) and drift in the direction of the velocity . The other half hits the stator where the drift component is lost.
As a result, the drifted molecular flow Qp towards an imaginary section A is:
Where ‘bh’ is the area of the cross section A and ‘ubh’ is the gas volume drifting through section A in one second.

## Slide 136
The pumping speed of molecular pumps:
Depends linearly on the speed of the moving wall
Does not depend on the nature of the gas (in the frame of this model)molecular pumps are not selective!
Gas pumping: momentum transfer
Molecular pumps
Converting in pressure-volume units and dividing by P, the pumping speed is obtained:

## Slide 137
Gas pumping: momentum transfer
Molecular pumps
The moving wall produces a higher gas density near the pump outlet. In turn, the pressure gradient generates a gas backflow Qbf that can be written as:
where is the conductance of the unit-length duct of cross section area bh. For h<<b:
The net flow is therefore given by:
When the net flux is zero, the maximum compression ratio is reached; in this case:

## Slide 138
Integrating between inlet and outlet (distance L) the maximum compression ratio is obtained:
High compression ratio for molecular pumps can be obtained for:
Fast moving surfaces: at least of the same order of the average molecular velocity (>100 ms-1).
Narrow and long pumping ducts (lower backstreaming).
Heavy masses the maximum compression ratio depends strongly on the molecular mass, the lowest being for H2 The ultimate pressure of molecular pumps is dominated by H2.
L and h length and width of the pump duct
Gas pumping: momentum transfer
Molecular pumps

## Slide 139
Gas pumping: momentum transfer
Turbomolecular pumps
For many years, the industrialisation of molecular pumps was hindered by the machining of the locking slot. The narrow gap between rotor and stator created mechanical issues due to tolerances and thermal expansion.
The problem was overcome in 1957 when W. Becker invented the turbomolecular pump (TMP). In the TMP, the momentum transfer is produced by rapidly rotating blades rather than parallel surfaces.
The molecules seen from the blades have a velocity oriented toward the blades’ channels when they come from space 1. From space 2, most of the molecules hit the blades and are backscattered
A significant gas flow is set only if at least .

## Slide 140
In the real turbomolecular pump, the gas is compressed by several series of rotating blades.
Every series of rotating blades is followed by a series of static blades. Molecules transmitted through the rotating blades’ channels hit the static blades; as a result, the angular distribution of velocity is randomized and the molecules are ready for the next compression stage.
The momentum transfer is effective only if the molecules do not experience inter-molecular collisions after hitting the blades; this is equivalent to say that the mean free path has to be larger than the blade distances. As a result, this type of pumps works at full pumping speed only in molecular regimes (P < 10-3 mbar).
Gas pumping: momentum transfer
Turbomolecular pumps

## Slide 141: Maglev Turbopump Design
High vacuum flange
Turbo blades
Holweck stage
Drive- and levitation unit
Maglev electronic
Purge Gas
Cooling Water
Fore Vacuum Flange
Courtesy of Florian Hess

## Slide 142: Different Flow Areas
Molecular flow
Kn > 10
Knudsen flow
Kn = 0,1…10
Viscous flow
Kn < 0,1
Courtesy of Florian Hess

## Slide 143: Theory of Turbo Pumping Effect
α
1
2
t
h
b
Compression ratio
u
g = geometric factor
Source: Karl Jousten, Wutz Handbuch Vakuumtechnik, edition 9, Vieweg, p. 329
Courtesy of Florian Hess

## Slide 144: Gas Type dependant Compression Ratio
Courtesy of Florian Hess

## Slide 145: Theory of Turbo Pumping Effect
Pumping Speed
α
1
2
t
h
b
u
Source: Karl Jousten, Wutz Handbuch Vakuumtechnik, edition 9, Vieweg, p. 328 f

## Slide 146: Gas Type dependant Pumping Speed

## Slide 147: History of the Turbopump
In 1958 Dr. W. Becker invents the worlds first Turbopump
Invention was done by chance
TVF 400
Pumping speed 140 l/s for N2
Weight: approx. 60 Kg
Courtesy of Florian Hess

## Slide 148: Turbopumps today
HiPace 300
Pumping speed 260 l/s for N2
Weight: 6.2 Kg
Courtesy of Florian Hess

## Slide 149: General Turbopump Design
Classical Design
Modern Design
Single flow design
Double flow design
Courtesy of Florian Hess

## Slide 150: Different Bearing Concepts
Hybrid Bearing
Robust design
Less vibrations
Easy maintenance
Oil lubrication
Cantilever Bearing
Maglev Bearing
No field service
Grease lubrication
Compact size
Optimized crash design
Courtesy of Florian Hess

## Slide 151: Different Rotor Concepts
Pure Turbo
Gas flow optimized
Suitable for process applications
With Holweck Stage
Compression optimized
Ultra High Vacuum
Courtesy of Florian Hess

## Slide 152
S: no strong influence of gas molecular mass + linear dependence on revolution speed
K0: exponentially dependent on: revolution speed and square root of gas molecular mass
High pumping speed
High compression ratio
Courtesy of Pfeiffer Vacuum
The indications given by the simplified model for molecular pumps are valid also for the TMP:
Gas pumping: momentum transfer
Turbomolecular pumps

## Slide 153
TMP pumping speeds are in the range from 10 l/s to 25,000 l/s.
Their ultimate pressure (H2) is of the order of 10-10, 10-11 mbar
Courtesy of Pfeiffer Vacuum
http://www.pfeiffer-vacuum.com
Vacuum Technology KnowHow
Gas pumping: momentum transfer

## Slide 154
Gas pumping: momentum transfer
Turbomolecular pumps
At the pump outlet, the compressed gas is evacuated by a mechanical pump (often called backing pump) operating in viscous regime. The turbomolecular and its backing pump are in general assembled in a single unit that includes power supplies, controls and instrumentation.

## Slide 155
The main drawback of TMP is related to possible mechanical failures leading to definitive damage of the high-speed rotor.
Gas pumping: momentum transfer
Turbomolecular pumps

## Slide 156
Advantages of TMP:
    constant pumping speed in a large range of pressure
    no memory effect (the gas is definitively evacuated) nor gas selectivity
    start working at relatively high pressure (as soon as molecular regime is attained)
Disadvantages of TMP:
    mechanical fragility
    risk of contamination from the backing pump
    need of venting anytime the pump is stopped to block backstreaming of contaminationsneed of valve between TMP and vacuum vessel
    intrinsic limitation in ultimate pressure of H2
Present trend:
    Use of dry pumps as backing pumps (but lower compression ratio than oil pumps)
    Increase compression ratio by adding molecular drag stages below the set of TMP blades (very compact design)
    Remove all lubricated mechanical bearing by magnetic rotor suspension (higher cost).
Gas pumping: momentum transfer
Turbomolecular pumps

## Slide 157
http://www.youtube.com/watch?v=1xZe1H2XHhM&feature=youtube_gdata_player
Gas pumping: momentum transfer
Turbomolecular pumps
http://www.youtube.com/watch?v=XOca57gPsq4

## Slide 158
To do so the sojourn time on the wall has to be much longer than the typical time of the accelerator run. An estimation of sojourn time is given by the Frenkel law J. Frenkel, Z. Physik, 26, 117 (1924):
where Ea is the adsorption energy and s.
Ea >> kBT Chemical pumps (getter pumps)
T << cryopumps
Gas pumping: capture pumps
Capture pumps remove gas molecules by fixing them on an internal wall.
To be efficient, capture pumps must block the gas molecules for an average time (sojourn time ts) much longer than the typical running time of the accelerator.
Ea

## Slide 159
Gas pumping: capture pumps
In chemical pumps, the binding force involves electron exchange between gas molecules and surfaces (chemisorption). T
Typical binding energies for getter pumps are higher than 1 eV/molecule.
In cryogenic pumps, Van der Waals interactions might be sufficiently strong to fix molecules to cold surfaces (physisorption).
In this case, the binding energy can be much lower than 0.5 eV/molecule.

## Slide 160
In SIP the residual gas is ionized in a Penning cell.
The ions are accelerated towards a cathode made of reactive metal
The collisions provoke sputtering of reactive metal atoms that are deposited on the nearby surfaces.
Gas pumping: capture pumps / sputter ion pumps

## Slide 161
In the diode SIP, the anode is composed of several open cylinders made of stainless steel at a positive electrical potential (from 3 to 7 kV). The cathodes are plates of Ti -at ground potential- placed a few millimetres from both anode extremities.
A magnetic field parallel to the cells’ axis is generated by external permanent magnets (about 0.1 T). In this configuration, the crossed electrical and magnetic fields trap electrons in long helical trajectories resulting in an increased probability of gas ionisation.
Gas pumping: capture pumps / sputter ion pumps
http://www.youtube.com/watch?v=_snzYepQTjI

## Slide 162
Gas pumping: capture pumps / sputter ion pumps
The pumping action is given by:
    chemical adsorption onto the reactive metal layer and subsequent burial by additional metallic atoms of gas molecules: all gases except rare gases
    implantation of gas ions in the cathode and of energetic neutrals bounced back from the cathode in the deposited film: only mechanism of pumping for rare gases
    diffusion into the cathode and the deposited film: only H2
http://www.youtube.com/watch?v=_snzYepQTjI

## Slide 163
Gas pumping: capture pumps / sputter ion pumps
Molecular implantation in the cathodes is not a permanent pumping; in fact, the progressive erosion due to gas ion sputtering soon or later frees the implanted gas.
Once released from the cathodes, reactive gas species are chemisorbed while rare gases remain in the gas phase until the next implantation.

## Slide 164
Pressure instabilities have been reported after excessive pumping of rare gas:
the continuous erosion liberates gas that increases the pressure, which in turn increases the erosion.
Consequently, a pressure rise is generated; this is stopped when most of the rare gas is implanted in a deeper zone of the cathode.
This type of instability is typically observed when large quantities of air are pumped (Ar 1% in air) and are known as ‘argon disease’.
Gas pumping: capture pumps / sputter ion pumps

## Slide 165
To increase the pumping efficiency of noble gas, the rate of ions implantation in the cathode has to be reduced while increasing the rate of energetic neutrals impingement on the anode and their burial probability.
Two different approaches:
  Heavier atoms for the cathode
  Ta (181 amu) is used instead of Ti (48 amu). The ions, once neutralized, bounce back at higher energy and rate these pumps are called ‘noble diode’
Gas pumping: capture pumps / sputter ion pumps

## Slide 166
Gas pumping: capture pumps / sputter ion pumps
  Different geometry of the Penning cell.
    Three electrodes are used: triode pumps. The cathodes consists of a series of small platelets aligned along the cell axis.
    The collisions ion-cathode are at glancing angle higher sputtering rate of Ti atoms + higher probability of neutralization + higher energy of bouncing + lower probability of implantation in the cathode.

## Slide 167
Gas pumping
DIODE
TRIODE

## Slide 168
Gas pumping: capture pumps / sputter ion pumps
An improved triode pump is the StarCell® produced by Agilent. Two sheets of Ti that are cut and bent to form radial blades (see Fig 55) replace the series of small plates. This increases the rigidity of the cathode and reduces the risks of short circuits

## Slide 169
Nominal pumping speed for N2:
Agilent StarCell ®
DN | S [ l s-1]
63 | 50
100 | 70/125
150 | 240/500
Pumping speed for SIP depends on the pressure at the pump inlet and the nature of the gas.
GAS | DIODE PUMPS | TRIODE PUMPS
AIR | 1 | 1
N2 | 1 | 1
O2 | 1 | 1
H2 | 1.5-2 | 1.5-2
CO | 0.9 | 0.9
CO2 | 0.9 | 0.9
H2O | 0.8 | 0.8
CH4 | 0.6-1 | 0.6-1
Ar | 0.03 | 0.25
He | 0.1 | 0.3
Pumping speed normalized to air
Gas pumping: capture pumps / sputter ion pumps
Nominal pumping speed

## Slide 170
Courtesy of Agilent Vacuum
http://www.chem.agilent.com/en-US/Products/Instruments/vacuum
Gas pumping: capture pumps / sputter ion pumps

## Slide 171
Hydrogen pumping by SIP
H2 is mainly pumped by diffusion into the cathode.
To be adsorbed, H2 must be dissociated. Only 2.5% of the ions created in a low-pressure H2 Penning discharge are H+ ions.
The dissociation is possible only on atomically clean Ti.
H2+ ions have poor sputtering yield: 0.01 at 7 KeV on Ti.
When H2 is the main gas, it takes a long time to clean the cathode surface by sputtering.
As a consequence, at the beginning of the operation the pumping speed for H2 is lower than the nominal and increases gradually with time.
Gas pumping: capture pumps / sputter ion pumps

## Slide 172
Hydrogen pumping by SIP
The simultaneous pumping of another gas has strong effects on H2 pumping speed.
        Higher sputtering yieldfaster cleaning higher pumping speed
        Contaminating of the Ti surface lower pumping speed
        Desorption of implanted H ions lower pumping speed
When the concentration of H2 is higher than the solubility limit in Ti, hydride precipitates are formed Ti expansion and hydrogen embrittlement short circuits and cathode brittleness (for 500 l/s pumps: typical value are 13000 mbar.l of H2)
Gas pumping: capture pumps / sputter ion pumps

## Slide 173
High Pressure Operation
High pressure (>10-5 mbar) operation can generate thermal run-away. It is frequently noticeable during the pumping of H2 or after the absorption of high quantity of H2 (for example due to pumping of H2O).
The Penning discharge heats the cathode and provokes gas desorption, which enhance the discharge. This positive feedback mechanism can melt locally the cathode.
The total electrical power given to the pump has to be limited at high pressure.
Gas pumping: capture pumps / sputter ion pumps

## Slide 174
Pressure measurement by ion pumps
The discharge current of the penning cells can be used for pressure measurement.
In the low pressure range, the current measurement is limited by field emission (leakage current): pressure reading limitation in the 10-9 mbar range.
By reducing the applied voltage in the lower pressure range, the pressure measurement is possible down to 10-10 mbar.
Courtesy of Agilent Vacuum
Gas pumping: capture pumps / sputter ion pumps

## Slide 175
The surface of getter materials reacts with gas molecules by forming stable chemical compounds.
This is possible only if the surface is clean, free of contamination and native oxide.
The clean metallic surface is obtained by:
Sublimating the reactive metal in situ Evaporable Getters, Sublimation Pumps
Dissolving the surface contamination into the bulk of the getter material by heating in situ (activation): Non-Evaporable Getters NEG.
Gas pumping: capture pumps / getter

## Slide 176
Getter surfaces are characterized by the sticking probability a :
For a=1, the pumping speed of the surface is equal to its maximum theoretical pumping speed.
Getter materials do not pump rare gases and methane at room temperature.
In that respect, they always need auxiliary pumping to keep a stable pressure
Gas pumping: capture pumps / getters
Sticking probability

## Slide 177
Gas pumping: capture pumps / getters
The sticking probabilities of homonuclear diatomic molecules are in general lower than that of molecules composed of dissimilar atoms.
The values of a for common gas species in vacuum systems may be ordered in a growing series as follows:
N2 < H2 < O2 < H2O < CO < CO2
For N2 typical a are in the high 10-3 to low 10-2 range; for CO and CO2, a close to 1 may be measured (for negligible coverage).
Sticking probability

## Slide 178
Gas pumping: capture pumps / getters
The sticking probability is strongly dependent on the gas surface coverage, i.e. the quantity of gas previously pumped on the getter material.
As the gas surface coverage increases, the sticking probability decreases. The pumping becomes negligible when the surface is fully covered by gas molecules. The pumping is restored by regenerating the surface.
As an example, complete saturation of a smooth NEG surface is obtained when about 1015 CO molecules cm-2 are adsorbed. Quantities about 5-10 times higher are reported for H2O and O2.
H2 dissociates on the getter surface and diffuses into the bulk of the NEG material. As a result, its pumping does not block the adsorption of other molecules.
The pumping capacity of hydrogen is orders of magnitude higher than that for other gas species.
Sticking probability

## Slide 179
Due to the high mobility of H atoms and the relatively low binding energy, the pumping of this gas is reversible: H2 may be released by heating.
The opposite effects of pumping and desorption set up an equilibrium pressure PH2 that depends on heating temperature T and hydrogen concentration cH.
The three variables are correlated by the Sievert’s law. In the logarithmic form:
A and B are typical values for a given getter material.
Gas pumping: capture pumps / getters
Sievert’s law

## Slide 180
For particle accelerators Ti is the sublimated metal.
Ti alloy rods are heated up to 1500°C attaining a Ti vapour pressure of about 10-3 mbar.
The Ti thin film is deposited onto the pump vessel.
In some cases, the wall of the sublimation pumps may be cooled to liquid nitrogen temperature. In this case, the sublimated film is rougher and consequently the sticking probabilities are higher.
Courtesy of Kurt J. Lesker Company
http://www.lesker.com/newweb/Vacuum_Pumps
Gas pumping: capture pumps / sublimation pumps

## Slide 181
Gas pumping: capture pumps / sublimation pumps
The sticking probabilities depend on the nature
of the gas and the quantity of gas already pumped.
The sticking probability is negligible:
For CO, one monolayer adsorbed
For of O2 several monolayer
For N2 fraction of monolayer
A new sublimation is needed.
Hydrogen diffuses in the Ti filmmuch higher
capacity
max
A K Gupta and J H Leek, Vacuum, 25(1975)362

## Slide 182
Heating in vacuum
Oxide layer dissolution-> activation
T = Ta
T = RT
Surface oxide
Active surface
T = RT
No pumping
Pumping
The dissolution of the native oxide layer (ACTIVATION) is possible only in metals having very high oxygen solubility limit, namely the elements of the 4th group: Ti, Zr and Hf.
Gas pumping: capture pumps / NEG

## Slide 183
Gas pumping: capture pumps / sublimation pumps
La
High O
solubility limit

## Slide 184
The activation temperature of the 4th group elements can be decreased by adding selected elements that increase oxygen diffusivity.
NEG materials are produced industrially by powder technology. Small grains are sintered to form pellets, discs or plates. The grains can also be pressed at room temperature on metallic ribbon.
A typical alloy produced by SAES Getter is St707:
Element | Concentration [wt. %] | Main role in the alloy
Zr | 70 | High O solubility limit. Chemical reactivity
V | 24.6 | Increases O diffusivity, Chemical reactivity
Fe | 5.4 | Reduces pyrophoricity
Full pumping speed is obtained after heating at 400°C for 45’ or 300°C for 24h
Gas pumping: capture pumps / NEG

## Slide 185
Courtesy of SAES Getters, www.saesgetters.com
Different shapes of NEG pumps
Gas pumping: capture pumps / NEG

## Slide 186
The high porosity of NEG materials allows pumping of relatively high quantities of gas without reactivation: for CO about 100 times higher than those for sublimation pumps per unit of geometrical surface of active metal.
Courtesy of SAES Getters, www.saesgetters.com
Gas pumping: capture pumps / NEG

## Slide 187
The maximum H2 sorption capacity is limited by H2 embrittlement of the NEG elements.
In general a safe limit is 20 Torr l/g is given by the supplier.
The stored H2 can be desorbed by heating and pumping out with an auxiliary pump (for example a turbomolecular pump).
Gas pumping: capture pumps / NEG

## Slide 188
Courtesy of SAES Getters, www.saesgetters.com
Gas pumping: capture pumps / NEG

## Slide 189
The time tR needed to regenerate a NEG pump is given by:
Where:
qf and qi are the final and initial concentrations in Torr l/g
Saux is the pumping speed of the TMP
MNEG the mass of the NEG in g
A and B typical values of the NEG material
For St707: A=4.8, B= -6116
The equilibrium pressure during the regeneration heating is given by:
q is the actual hydrogen concentration
Gas pumping: capture pumps / NEG

## Slide 190
In general in particle accelerators the gas source is uniformly distributed along the beam pipe
Pumps located at regular distances from each other give rise to a parabolic pressure profile with minima at the pumps location.
Even for large aperture, the average pressure is quickly limited by the chamber conductance when increasing the spacing between the pumps: the benefit of increasing the lumped pumping speed is minimal.
It is more efficient to increase the number of pumps. But this can be very expensive or not possible due to space limitation. Linear pumping is the solution.
Gas pumping: capture pumps / NEG

## Slide 191
LEP dipole vacuum chamber
Linear pumping may be obtained by NEG ribbons.
Gas pumping: capture pumps / NEG
The first application was in the LEP.

## Slide 192
The trend in vacuum technology consists in moving the pump progressively closer to the vacuum chamber wall.
The ultimate step of this process consists of transforming the vacuum chamber from a gas source into a pump.
One way to do this is by “ex-situ” coating the vacuum chamber with a NEG thin film that will be activated during the “in situ” bakeout of the vacuum system.
NEG film
NEG strip
NEG films do not need space, electric power, insulation and feedthroughs
⇒ simplified design and larger pumping speed
Gas pumping: capture pumps / NEG
The coating techniques is sputtering.

## Slide 193
manifold
chambers
extensions
Solenoid
L=8m
f=60cm
NEG coating unit at CERN
Ti-Zr-V coating production for the LHC: coating facility
Gas pumping: capture pumps / NEG

## Slide 194
manifold
chambers
extensions
Solenoid
L=8m
f=60cm
3mm wires of Ti, Zr and V
NEG coating unit
Gas pumping: capture pumps / NEG

## Slide 195
Gas pumping: capture pumps / NEG

## Slide 196
Intertwisted Ti, Zr, and V elemental wires give rise to Ti30 -Zr30 -V40 film composition
Which is the necessary condition for the exceptional fast activation in the TiZrV system?
The lowest activation temperature has been found in a wide range of composition in the Ti-Zr-V system: 180 °C (24 h heating).
Gas pumping: capture pumps / NEG

## Slide 197
Nanocrystalline (3÷5 nm grain size) structure as confirmed by TEM
Well crystallised structure (grain size ≥ 100 nm)
30
40
50
60
TiZrV
Substrate
 (degrees)
30
40
50
60
TiZrV
Substrate
 (degrees)
Gas pumping: capture pumps / NEG

## Slide 198
About 1400 chambers for the warm zones of the LHC were coated with a Ti-Zr-V film.
Ti-Zr-V coating production for the LHC
Gas pumping: capture pumps / NEG

## Slide 199
TiZrV/St.St.
TiZrV/Smooth Cu
1 µm
TiZrV/Glidcop.
TiZrV/Al
Gas pumping: capture pumps / NEG

## Slide 200
TiZrV/Al
TiZrV on rough copper: Cu6
Gas pumping: capture pumps / NEG

## Slide 201
300°C
250°C
200°C
1 m
100°C
150°C
350°C
Gas pumping: capture pumps / NEG

## Slide 202
After activation a NEG film surface is very clean resulting in both a large pumping speed and reduced degassing (both thermal and ion/radiation/electron induced).
Very important bonus: low secondary electron yield.
NEG films trap the gas coming from the substrate material
Gas pumping: capture pumps / NEG

## Slide 203
PSD at ESRF
Gas pumping: capture pumps / NEG

## Slide 204
Cryopumps rely on three different pumping mechanisms:
Cryocondensation: is based on the mutual attraction of similar molecules at low temperature:
  the key property is the saturated vapour pressure, i.e. the pressure of the gas phase in equilibrium with the condensate at a given temperature. The attainable pressure is limited by the saturated vapour pressure.
  Only Ne, H2 and He have saturated vapour pressures higher than 10-11 Torr at 20 K.
              The vapour pressure of H2 at 4.3 K is in the 10-7 Torr range, at 1.9 lower than 10-12 Torr.
              Large quantity of gas can be cryocondensated (limited only by the thermal properties of the condensate phase)
Gas pumping: capture pumps / cryo

## Slide 205
Gas pumping: capture pumps / cryo

## Slide 206
Cryosorption: is based on the attraction between molecules and substrate. The van der Waals forces are much stronger than those between similar molecules:
  Gas molecules are pumped at pressure much lower than the saturated vapour pressure providing the adsorbed quantity is lower than one monolayer.
  Porous materials are used to increase the specific surface area; for charcoal about 1000 m2 per gram are normally achieved.
  The important consequence is that significant quantities of H2 can be pumped at 20 K and He at 4.3 K.
  Submonolayer quantities of all gases may be effectively cryosorbed at their own boiling temperature; for example at 77 K all gases except He, H2 and Ne.
Gas pumping: capture pumps / cryo

## Slide 207
Cryotrapping : low boiling point gas molecules are trapped in the layer of an easily condensable gas. The trapped gas has a saturation vapor pressure by several orders of magnitude lower than in the pure condensate. Examples: Ar trapped in CO2 at 77 K; H2 trapped in N2 at 20 K.
Gas pumping: capture pumps / cryo

## Slide 208
Modern cryopumps take advantages of the first two mechanisms.
The cryocondensation takes place on a cold surfaces, in general at 80 K for H2O and 10 or 20 K for the other gases.
The cryosorption of H2, Ne and He is localised on a hidden surface where a porous material is fixed. This surface is kept away from the reach of the other molecules.
Helium is the working fluid of modern cryopumps.
Gas pumping: capture pumps / cryo

## Slide 209
Courtesy of Oerlikon Leybold Vacuum www.oerlikon.com/leyboldvacuum
Gas pumping: capture pumps / cryo

## Slide 210
Characteristics of Cryopumps
Starting Pressure
  Cryopumps should be started when the mean free path of molecules is higher than the pump vessel diameter: P<10-3 mbar. Otherwise the thermal load is too high.
  In addition a thick condensate layer must be avoided.
  They always need auxiliary pumps.
Pumping speed
  High effective pumping speed for all gases. Pumping speed from 800 l/s up to 60000 l/s are commercially available .
  Pumping speed for water vapour close to the theoretical maximum.
Gas pumping: capture pumps / cryo

## Slide 211
Maximum Gas Intake (Capacity)
  At the maximum gas intake, the initial pumping speed of the gas is reduced by a factor of 2.
  Condensed gases: the limitation is given by the thermal conductivity of the gas layer.
  Adsorbed gases: the capacity depends on the quantity and properties of the sorption agent; it is pressure dependent and generally several orders of magnitude lower compared to that of condensable gases.
Gas pumping: capture pumps / cryo

## Slide 212
Cryopumps require periodic regeneration to evacuate the gas adsorbed or condensed.
To remove all captured gas, the pump is warmed at room temperature. The desorbed gas is removed by mechanical pumps (in general, for accelerators, mobile TMP).
During regeneration, the rest of the system must be separated by a valve.
In the majority of application, the performance deterioration is given by the gas adsorbed on the second stage (10-20 K). A partial regeneration may be carried out for a shorter time while water vapour is kept on the first stage at temperatures lower than 140 K.
Gas pumping: capture pumps / cryo

## Slide 213
Courtesy of Oerlikon Leybold Vacuum www.oerlikon.com/leyboldvacuum
Gas pumping: capture pumps / cryo

## Slide 214
 | Advantages | Disadvantages
TMP | No memory effects Constant pumping speed for pressures lower than 10-3 mbar Pumping speed independent of total gas load Starts working at high pressures (molecular regime) | Mechanical fragility Risk of contamination from the backing pump Need of venting anytime the pump is stopped Need of valve on the main flange Intrinsic limitation in ultimate pressure of H2 Possible vibrations Maintenance
SIP | Clean pumping No maintenance No vibrations Installation in any orientation Relatively long lifetime Relatively low cost Limited but high H2 capacity The pump current gives a pressure reading | Low capture probability Gas Selectivity and limited capacity Memory effects (in particular for rare gases) Ignition in 10-5 mbar range Bulky Difficult starting for old pumps Production of charged particles in particular at start-up Field emission problems for old pumps Fringing magnetic field Safety issue: high voltage
Gas pumping

## Slide 215
 | Advantages | Disadvantages
NEG pumps | Clean vacuum High pumping speed for reactive gases With SIP, extremely low vacuum can be achieve High gas capacity for porous NEG Low cost Electrical power needed only for activation; it works in case of power cut No maintenance No vibration | Selective pumping (no pumping of rare gases and methane) H2 enbrittlement if regeneration is not applied Formation of dust particles is not excluded Safety issue: pyrophoric, it burns when heated in air at high temperature
Gas pumping

## Slide 216
 | Advantages | Disadvantages
Sublimation Pumps | Clean vacuum High pumping speed for reactive gases With SIP, extremely low vacuum can be achieve Low cost Electrical power only for sublimation; it works in case of power cut Limited maintenance (filament change) No vibration | Very limited capacity Need frequent sublimations at high pressure Ti film peel-off for high sublimation rates Selective pumping (no pumping of rare gases and methane) Risk of leakage current in high voltage insulators Ideal for low pressure applications
Gas pumping

## Slide 217
 | Advantages | Disadvantages
Cryopumps | Very large pumping speed for all gases Clean vacuum High pumping capacity Limited selectivity | Cost and maintenance Relatively large volume needed (including refrigerator) Gas release in case of power cut Reduced pumping efficiency for H2 for high quantity of gas adsorbed: regeneration needed Need of valve on the main flange
Gas pumping

## Slide 218
Part 6
Instrumentation

## Slide 219
Instrumentation

## Slide 220: Thermal Conductivity Gauges
  Four possible mechanisms to take heat away from sensing element:
    Radiative heat transfer to surroundings
    Conduction along the sensing element
    Conduction through the residual gas
    Convective heat flow to the residual gas
Hot wire
1/2 WC
1/2 WC
Thermocouple
joint
WR
WG
WCNV
Heating
current
Instrumentation
Principle of operation:
  Pirani gauge: temperature dependent resistance of a hot wire
  Thermocouple gauge: temperature dependent voltage from a thermocouple joint

## Slide 221
Radiative heat transfer: does not depend on pressure
Conduction along the wire: does not depend on pressure
Convection for pressure higher than 100 mbar
Conduction through the residual gas: linear
with pressure. Gas properties determine
variation in sensor response to
pressure
Thermal Conductivity Gauges
Instrumentation

## Slide 222
Constant current mode:
  As pressure decreases, less heat is removed by conduction through the gas from the sensing element, and temperature increases.
  Measure temperature dependent change of resistance (or TC voltage).
Constant temperature mode:
  Adjust current to heater to maintain constant sensor property (resistance or voltage)
  Use Wheatstone bridge arrangement to optimize sensitivity
  Correlate current flow against pressure
Thermal Conductivity Gauges
Instrumentation

## Slide 223
Thermal Conductivity Gauges
Instrumentation

## Slide 224
Hot cathode gauge
  Thermionic source (electrons) cause inelastic collisions with gas, producing ions
  Ions collected at (-) biased surface
  Electrons travel to a grid
Pre-1950 lower limit of 10-8 mbar
  Soft x-rays produced at grid result in photo-electron flow away from collector (same as ion flow to collector)
  Bayard-Alpert modification: fine wire collector surrounded by grid  range extended to 10-12 mbar
+
-
-
collector
grid
filament
Ionization Gauges
Instrumentation

## Slide 225
Bayard-Alpert gauges are used for vacuum measurement purposes in the range 10-5 -10-12 mbar.
It is a hot filament ionization gauge. Electrons emitted by the filament perform oscillations inside the grid and ionize the molecules of the residual gas. Ions are then collected by an electrode.
  Where :
    I+ is the ion current
    I- is the filament current
    σ is the ionisation cross section
    n the gas density
    L the electron path lenght
The gauge needs to be calibrated
X-ray limit of a few 10-12 mbar

## Slide 226
SVT Bayard Alpert modulated gauge 10-12 Torr
Ionization Gauges
Instrumentation

## Slide 227
Cold Cathode Gauge: Invented by Penning in 1937 - electrical configuration like diode SIP
1952 design of wire-in-cylinder with reverse polarity (+ wire) provides key performance improvement
“Inverted Magnetron” - Magnetic field produces long helical path for electrons & ions  many collisions
Ionizing electrons part of self-sustaining gas discharge
No background current to mask the ion current
Power into CCG ~ 0.1 W
Ionization Gauges
Instrumentation

## Slide 228
“Striking time” is delay before discharge starts.
At 10-4 Pa, ts ~ seconds; at 10-8 Pa, ts ~ hours or days
Nearly linear dependence of ion current with pressure:
Typical anode voltage +3 kV
Ionization Gauges
Instrumentation

## Slide 229
Residual gas analyzers
Instrumentation

## Slide 230
RGA