## Slide 1: 3 MeV instrumentation commissioning
U. Raich, Federico Roncarolo, F. Zocca
Linac-4 Beam Coordination Committee 10.1.2013

## Slide 2: Current Setup

## Slide 3: LEBT
Faraday Cup
BCT
Wire Grids

## Slide 4: MEBT (chopper line)
BCT
BCT
BWS
BWS

## Slide 5: Measurement Bench
Emittance
Meter Slit
Emittance
Meter Grid
BCT
BSM
Spectrometer
Magnet
Spectrometer
Grid
Halo
Monitor
BCT
BPM

## Slide 6: Instruments installed
LEBT:
  Faraday Cup
  Wire Grid
MEBT (chopper line)
  2 wire scanners
  2 Beam Current Transformers (BCTs)
Measurement Line
  Slit/Grid Emittance meter
  2 BCTs
  2 (3) Beam Position Monitors (BPMs)
  Bunch Shape Monitor
  Spectrometer
  Halo Monitor

## Slide 7: Measurement campaigns
~ 3 months at the 3 MeV test stand
May – October move and re-installation into the L4 tunnel
Since November ~ 3 weeks of measurements

## Slide 8: Faraday Cup
Connected to remotely accessible Oscilloscope
Guard ring polarization set to -1000V

## Slide 9: Wire Grids in LEBT
24 wires per plane with spacing 1mm, 3mm, 5mm
Wire signals sampled with 250 kHz

## Slide 10: Wire Grid in LEBT
Can polarize frame and wires
Measurements for different solenoid settings
Add wire signals and compare to FC
Frame needs to be polarized, wire polarization is less important

## Slide 11: Beam Current Transformers
Instrument mostly used to see if beam arrives
Signal sampled at max. 200 MHz (2 samples added -> 100 MHz)
Multiple level shielding
Calibration after each beam pulse

## Slide 12: BCT signals
White areas show integration windows

## Slide 13: BCTs give an idea on chopper rise time

## Slide 14: Wire scanners
40 um carbon wires
First X-shaped, now L-shaped
Wires can be polarized
Readout at 205 kHz
Now configured in L-shape

## Slide 15: Wire Scanners and Chopper

## Slide 16: Crosstalk between wires
On L-shaped wire scanners crosstalk from one wire to the next can be separated
Measured crosstalk when increasing positive wire polarization
Profile size shows ~ 2% dependence on wire polarization up to ~ 40V.

## Slide 17: Crosstalk versus polarization
WS2 – 0V
WS2 – 5V
WS2 – 10V
WS2 – 15V
WS2 – 20V
WS2 – 40V
WS2 – 60V
WS2 – 80V
WS2 – 100V
WS2 – 120V

## Slide 18: Emittance Meter
Labview program sets
Motor position of slit and grid
(controlled by PLC and FESA)
Accesses FESA class to control
Amplifier gain, plane ...
Reads ADCs and saves data in
.csv format for further evaluation
1 emittance plot every 4 µs

## Slide 19: Beam position monitors
BPMs measure
Position
Intensity
Phase

## Slide 20: Intensity BCT-BPM

## Slide 21: Bunch Shape Measurements
Has been used to adjust the bunchers in the test stand
Brought back into operation in the L4 tunnel
No bunch shape measurements yet because of missing buncher control

## Slide 22: Concept of Laser Emittance Meter
Diamond
Detector
H- beam
H0 beamlet
beamlet
after drift
Principle:
H-
H0
Bending Magnet
x-scan
y-scan
H- beam
 | Diamond Detector
Type | Polycrystalline
Size | 18 x 18 mm
Resolution | 5 channels
Thickness | 500 µm
Readout | 5 GS/s Scope
 | Pulsed fiber-laser
Wavelength | 1080 nm
Energy | 0.9 mJ
tpulse | 110 ns
Beam quality – M2 | 1.7
Repetition rate | 30kHz
Beam transport | Fiber
Particlesstripped / bunch | < 4e5
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 23: Setup at 3 MeV Diagnostics Testbench
Laser
Focusing
Laser Delivery optics
Diamond Detector
Bending Magnet
H- beam
H0
Fiber delivery