## Slide 1: Beam Loss Monitors:
Overview of BLM Technology
Kay Wittenburg
3rd oPAC Topical Workshop on Beam Diagnostics
Vienna, 8-9 May 2014
Introduction
Dynamic range, sensitivity
Limitations in
  time and/or
  spatial resolution
Challenges associated to measurements of losses (in different machine types)
Radiation hardness
A comprehensive summary of the current state-of-the-art methods
Needs for further development

## Slide 2: 1. Introduction
Beam loss monitor systems are designed for measuring beam losses around an accelerator or storage ring. A detailed understanding of the loss mechanism, together with an appropriate design of the BLM-System and an appropriate location of the monitors enable a wide field of very useful beam diagnostics and machine protection possibilities.
Regular (controlled, slow) loss
Those losses are typically not avoidable and are localized on the collimator system or on other (hopefully known) aperture limits. They might occur continuously during operational running and correspond to the lifetime/transport efficiency of the beam in the accelerator. The lowest possible loss rate is defined by the theoretical beam lifetime limitation due to various effects.
Irregular (uncontrolled, fast) losses
The irregular losses may be distributed around the machine and not obviously on the collimation system. Can be avoided and should be kept to low levels but may reach very high levels in case of an accident.
It is clearly advantageous to have a BLM System which is able to deal with both loss modes. But this means -> High Dynamic Range System!

## Slide 3
Detector Material | energy to create one electron [eV/e] | number of [e / (cm MIP)] (depends on dE/dx, resp. density) | Sensitivity S (for MIPs) [nC/rad]
Plastic Scintillator: | 250 – 2500 | 103 - 104 | ≈17·103 (∙ PMTgain) (1 ltr.)
Inorganic Scint. | 50 - 250 | 104 - 105 | ≈ 100·103(∙ PMTgain) (1 ltr.)
Gas Ionization: | 22 – 95 | 100 (Ar,1 atm., 200C) | ≈ 500 (∙ Elecgain) (1ltr)
Semiconductor (Si): | 3.6 | 106 | ≈ 50 (∙ Elecgain) (1 cm2 PIN-Diode)
Secondary emission: | 2-5%/MIP (surface only) | 0.02-0.05 e/MIP | ≈ 2·10-3 (∙ PMTgain) (8cm2)
Cherenkov light | 105 - 106 | 10 (H2O) -200 (fused silica) | ≈ 270 (∙ PMTgain) (1 ltr.)
Including Gain
and
1 Gy = 100 rad
2. Dynamic Range, Sensitivity
K. Wittenburg, Beam loss monitors, CAS2008
Specialised Beam Diagnostics School in Dourdan, France, CERN-2009-005
Lars Fröhlich, Beam Loss Monitors
ERL Instrumentation Workshop, Cornell University, 2-3 June 2008
http://tesla.desy.de/~lfroehli/download/ERL_instrumentation_ws_2008_BLMs.pdf
≈ 0.2·10-3 (∙ PMTgain) for 1 m Cherenkov fiber
>108 Difference in Sensitivity between different types
In these two references one can find a lot of details on various BLM types!

## Slide 4: 2. High dynamic loss monitoring by
Different BLM types/materials
FLASH FLASH LHC
SNS

## Slide 5: 2. High dynamic loss monitoring by
Different BLM types/materials
Diamond and Sapphire
Alexandr Ignatenko
Thesis, 2014
Brandenburgische Technische
Universität Cottbus-Senftenberg

## Slide 6: 2. High dynamic loss monitoring by
Small and Large
FLASH HERA
7.56 mm2
150 mm2
Note that the flux density of photons into the light guide is
“incompressible” !
=> The cross section of the scintillator should not be larger
than the cross section of the light guide -> I did not proof this rule, any experience with that?

## Slide 7: 2. High dynamic loss monitoring by
Detector:
PMTs:
Noise at max Gain ≈1 mV
Saturation ≈ 1 V
+Active gain variation ≈ 10³
Dynamic range ≈ 106
Ionization chamber (LHC)
Leakage current < 1 pA
Saturation ≈1 mA
Dynamic range ≈ 109
Electronic:
RF Amplifiers
Dynamic range ≈ 104
Log Amp. ≈ 105
ADC
12 bit ≈ 4 x 103
16 bit ≈ 6 x 104
24 bit ≈ 2 x 107 (SNS: VME ADC but 10 bits noise)
Counting
Dark count rate ≈1 Hz
Signal: Bunch rep. rate ≈10 MHz
Dynamic range > 109 (averaging!)
Noise to saturation

## Slide 8: 3a. Limitations in time resolution

## Slide 9: 3a. Limitations in time resolution
Typical reaction time of
  Rings: 1 - few turns -> > ≈ 10 ms -> Defines the detector time response
  Linac: Bunch distance (≈ 100 ns at bunch train or ≈ ms at single bunch) but important:
    Bunch by Bunch resolution -> Defines the detector time response
    Integration over bunch train
    Integration over some bunch trains
Alexander Kaukher
DESY
7th DITANET Topical Workshop on
Beam Loss Monitoring
December 7, 2011

## Slide 10: 3a. Limitations in time resolution
≥ turn by turn:
  Ionization chambers
  + Low bandwidth ADC allows high dynamic range
  Counting (many bunches)
  Allows super high dynamic range
T. Toyama et al., HB2008
Counting circuit for LHC Ion Chamber
Pin Diodes at HERA

## Slide 11: 3a. Limitations in time resolution
< 100 ns:
  PMT (or APD) + Cherenkov or Scintillator
  SEM + GHz Amplifier (or SEM-PMT)
  Solid State Detectors + GHz Amplifier
  + GHz ADC (limited dynamic range)!!!
The signal width from the R5900 PMT is as short as 20 ns, even after 50m twisted pair cable.

## Slide 12: 3a. Limitations in time resolution
Loss signal
Time
Diamond signal
Ionisation chamber
(40 us integration time)
B.Dehning
Beam Loss: New Developments, Detectors and Electronics ;
7th DITANET Topical Workshop on
Beam Loss Monitoring
December 7, 2011

## Slide 13: 3b. Limitations in spatial resolution (where to put BLMs)

## Slide 14: 3b. Limitations in spatial resolution (where to put BLMs)
Beam direction
Electron beam = small shower

## Slide 15: 3b. Limitations in spatial resolution (where to put BLMs)
Trajectory of Electrons after Energy Loss
A New Theoretical Design of BLM System for HLS II
Yukai Chen, Lijuan He, Juexin Li, Wei-min Li, Yuxiong Li
IPAC 2013
Electron Beam Loss Monitors for HERA
F. Ridoutt, W. Bialowons, K. Wittenburg; EPAC 1994
Important:
Beam Optics!
Tracking Codes
Beam direction

## Slide 16: 3b. Limitations in spatial resolution (where to put BLMs)
A local orbit distortion creates losses at high beta (in general at aperture limitations)
Secondaries at surface
Important:
Particle Shower
Monte Carlo Codes
Beam direction
High Energy Proton Beam = Large Shower

## Slide 17: 3b. Limitations in spatial resolution (where to put BLMs)
Loss location = middle of first quadrupole
Loss angle = 1.5 mrad
Loss intensity = 10^12 protons/sec
Technical Note ESS/AD/0032
Important:
Loss signal has to be calibrated by energy
Monte Carlo Codes
Beam direction
Low Energy Proton Beam = Shielded Shower

## Slide 18: 3b. Limitations in spatial resolution (where to put BLMs)
Beam direction
Local measurement
Vs.
Do losses appear somewhere?
Neutrons are everywhere in the accelerator tunnel in case of losses:
Few neutron detector BLMs are sufficient

## Slide 19: 3b. Limitations in spatial resolution (where to put BLMs)
detects the alphas

## Slide 20: 4) Challenges associated to measurements of losses
Just a small selection:
Very low energy machines
High background

## Slide 21: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    LINAC/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators

## Slide 22: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators
IBIC 2013 poster
-> Also: Cryogenic BLMs (not at low energy)

## Slide 23: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators
IBIC 2013 poster

## Slide 24: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators

## Slide 25: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators

## Slide 26: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators
  Known loss location
  Aperture limit
  Highest loss rate (hopefully)
  Machine + Collimator Protection
“Tails” in distributions are from the beam.
BLM signal is linear with proton intensity.
Left-right asymmetry of the shower depends on the collimator gap size and gap position.
Slight top-bottom asymmetry?
BLM signal depends on the impact position on the jaw.
Compares ~ OK with simulations (TT40).

## Slide 27: 4a) Challenges associated to measurements of losses at low energies
The problem: No or very few secondaries outside the vacuum chamber
Solutions:
Differential current measurement
    Limited position resolution
    Linac/transport only
BLMs very close to beam pipe
  Risk of wrong position
BLMs sensitive to neutrons
  Limited position resolution
Very sensitive BLMs
Use of BLMs at collimators
  Known loss location
  Aperture limit
  Highest loss rate (hopefully)
  Machine + Collimator Protection
0.9 and 72 MEV Proton collimator at PSI Cyclotron (Pic. From R. Dölling-PSI)
SNS PAC07 LINAC10
Plot of measured beam loss along the HEBT and IDump (Injection Dump) with the MEBT collimator in and out. Data shows significant reduction in beam loss in the HEBT and IDump with the MEBT collimation.

## Slide 28: 4b) Challenges associated to measurements of losses at high background
The problem: Limits the dynamic range
EM Noise
Reasons:
Shielding
Ground loops
RF
PS ripple (from magnets, from HV)
Kickers, septum
Magic
Ghosts
Sabotage
…
Solutions:
Blame the others!
(not very useful, I know…)

## Slide 29: 4b) Challenges associated to measurements of losses at high background
The problem: Limits the dynamic range
X-ray from cavities
Reasons:
Released electrons from cavity
Magic
Ghosts
Sabotage
…
Solutions:
Subtraction by software
Use of a x-ray insensitive detector
1 mS
X-ray
Beam loss
SNS
XFEL Beam Loss Monitor System
A. Kaukher, I. Krouptchenkov, D. Noelle (D. Nölle), H. Tiessen, K. Wittenburg
IPAC12

## Slide 30: 4b) Challenges associated to measurements of losses at high background
The problem: Limits the dynamic range
X-ray from cavities
Reasons:
Released electrons from cavity
Magic
Ghosts
Sabotage
…
Solutions:
Subtraction by software
Use of a x-ray insensitive detector
Beam Loss Detected by Scintillation Monitor Akihiko Miura, et al.
IPAC´11

## Slide 31: 4b) Challenges associated to measurements of losses at high background
The problem: Limits the dynamic range
X-ray from Synchrotron radiation
Reasons:
SR is unavoidable
Solutions:
Subtraction by software
Use of a x-ray insensitive detector: Cherenkov material: Quartz
Coincidence
K. Scheid, ESRF, 7th DITANET Topical Workshop on Beam Loss Monitoring; 2011
L. Fröhlich, FERMI, 7th DITANET Topical Workshop on Beam Loss Monitoring; 2011
S.L. Krameer NSLS II, 7th DITANET Topical Workshop on Beam Loss Monitoring; 2011
J. Perry, Jlab, PAC93

## Slide 32: 4b) Challenges associated to measurements of losses at high background
The problem: Limits the dynamic range
X-ray from Synchrotron radiation
Reasons:
SR is unavoidable
Solutions:
Subtraction by software
Use of a x-ray insensitive detector: Cherenkov material: Quartz
Coincidence: Counting
Coincidence technique: SR-Photons stop in one or the other PIN diode and are not counted!
The Beam Loss Monitoring System at ELSA
Dennis Proft , IPAC12
Installation and Test of a Beam Loss Monitor System for the S-DALINAC
Robert Stegmann, IPAC12
Beam Loss Monitors for the HERA Proton Ring
DESY HERA 90-11

## Slide 33
5) Radiation hardness

## Slide 34
S. Goulding, R.H. Pohl 1972
5) Radiation hardness

## Slide 35: 6) A comprehensive summary of the current state-of-the-art methods

## Slide 36: 6) A comprehensive summary of the current state-of-the-art methods
Cannot be answered since all kind of methods are in use (as seen from previous slides)
There is no “best method” since a useful method depends on various accelerator parameters
Therefore: Don’t trust on “state of the art”, often a well established method can be the best (Ion chambers and Scintillators+PMT are the most common BLMs)
However, new problems need new solutions: e.g. x-ray background-> Cherenkov rods and fibers, PIN-Coincidence
Still searching for a fast and sensitive detector with high dynamic range and high radiation damage threshold -> Diamonds?
Simulations are important to understand losses and the BLM response

## Slide 37: 6) A comprehensive summary of the current state-of-the-art methods
Simulation:
IBIC2013
Beam Delivery Simulation (BDSIM): A Geant4 Based Toolkit for Diagnostics and Loss Simulation
Monte Carlo Simulations of Beam Losses in the Test Beam Line of CTF3 Simulation for Radiation Field Caused by Beam Loss of C-ADS Injector II
Beam Loss Monitoring at the European Spallation Source
IPAC11
Comparative Studies into 3D Beam Loss Simulations
Monte Carlo Simulation of the Total Dose Distribution around the 12 MeV UPC Race-track Microtron and Radiation Shielding Calculations
Beamloss Study at J-PARC Linac by using Geant4 Simulation
b) Fiber based BLMs
IBIC13
Update on Beam Loss Monitoring at CTF3 for CLIC
Optical Fiber Based Beam Loss Monitor for Electron Storage Ring
Cherenkov Radiation for Beam Loss Monitor Systems
BIW12IBIC12
Development of Optical Fiber Beam Loss Monitor System for the KEK Photon Factory
Simulation and Measurement of Beam Loss in the Narrow-Gap Undulator Straight Section of the Advanced Photon Source Storage Ring
c) Diamond BLMs
IBIC12/13
Operation of Silicon, Diamond and Liquid Helium Detectors in the Range of Room Temperature to 1.9 Kelvin and After an Irradiation Dose of Several Mega Gray
A Prototype Readout System for the Diamond Beam Loss Monitors at LHC
Performance of Detectors using Diamond Sensors at the LHC and CMS
IPAC12
Advances in CVD Diamond for Accelerator Applications
BEAM HALO MONITOR FOR FLASH AND THE EUROPEAN XFEL
Investigation of the Use of Silicon, Diamond and Liquid Helium Detectors for Beam Loss Measurements at 2 Kelvin

## Slide 38
7) Needs for further development:

## Slide 39
7) Needs for further development:
Calibration of BLM signal in terms of lost particles
Dealing with saturation, avoiding, detecting
Extending the useful dynamic range and speed of loss measurements
Quite often, BLMs are used just to minimize losses. Mainly in superconducting hadron accelerators a calibration of the loss signal was done to define thresholds for quenches.
There is a need to calibrate the losses in terms of dose at high intense hadron accelerators to avoid activation, checking the 1W/m rule.
-> Reliable integration of BLMs into MPS
Beam lifetime measured by current and loss monitors agreed by factor 2 in HERAp. I’ve heard about the same at LHC.
BEAM LOSS LIMITS IN HIGH POWER PROTON LINEAR
ACCELERATORS
L. Tchelidze, IPAC2013

## Slide 40
7) Needs for further development:
Calibration of BLM signal in terms of lost particles
Dealing with saturation, avoiding, detecting
Extending the useful dynamic range and speed of loss measurements
It is not always obvious if your detector, amplifier, ADC circuit is saturating. PMTs behave crazy at saturation.
It’s known, but how to deal with it in operation?
Not much in literature available.

## Slide 41
7) Needs for further development:
Calibration of BLM signal in terms of lost particles
Dealing with saturation, avoiding, detecting
Extending the useful dynamic range and speed of loss measurements

## Slide 42
7) Needs for further development.
One BLM System is not enough
My final message:
Thanks for attention, questions?
your
accelerator!!!
for