## Slide 1
Least invasive beam profile measurements:
Ionization Profile Monitors and Beam Induced Fluorescence
P. Forck, C. Andre, F. Becker, T. Giacomini, Y. Shutko, B. Walasek-Höhne
GSI Helmholtz-Zentrum für Schwerionenforschung, Darmstadt, Germany
In collaboration with: T. Dandl, T. Heindl, A. Ulrich, Technical University München
J. Egberts, J. Marroncle, T. Papaevangelou et al., CEA/Saclay
OPAC Workshop Vienna, May 8th, 2014
Outline of the talk:
Ionization Profile Monitor IPM technical realization
 Beam based measurements at GSI synchrotron and storage ring
Beam Induced Fluorescence BIF monitor realization
 Energy scaling of signal and background 60MeV/u < Ekin< 750MeV/u
 Spectroscopic investigations for rare gases and N2
 Profiles & spectroscopy for pressure range 10-3 mbar < p < 30 mbar
Comparison IPM  BIF

## Slide 2
Expected Signal Strength for IPM and BIF-Monitor
Target electron density:
Proportional to vacuum pressure
 Adaptation of signal strength
 1/Ekin (for Ekin> 1GeV nearly constant)
Energy loss in 10-7 mbar N2 by SRIM
ion source
LINAC, cyclotron
synchrotron
Ionization probability proportional to
dE/dx by Bethe-Bloch formula:
Physics:
Energy loss of ions in gas dE/dx
Profile determination from residual gas
Ionization: roughly  100 eV/ionization
Excitation + optical photon emission:
roughly  3 keV/photon

## Slide 3
Expected Signal Strength for IPM and BIF-Monitor
Target electron density:
Proportional to vacuum pressure
 Adaptation of signal strength
 1/Ekin (for Ekin> 1GeV nearly constant)
Strong dependence on projectile charge for ions Zp2
Modification proton  ions: Zp(Ekin) . Charge equilibrium is assumed for dE/dx
Physics:
Energy loss of ions in gas dE/dx
Profile determination from residual gas
Ionization: roughly  100 eV/ionization
Excitation + optical photon emission:
roughly  3 keV/photon
Energy loss for l  1m: dE/dx l << E kin
acceptable for single pass beams
Care: synchr.multi pass; cryogenic envir.
1H
12C
40Ar
238U
Energy loss in 10-7 mbar N2 by SRIM
Ionization probability proportional to
dE/dx by Bethe-Bloch formula:

## Slide 4
Ionization Profile Monitor: Principle
Advantage: ‘4-detection scheme’ for ionization products
Detection scheme:
Secondary e- or ions accelerated by E-field
electrodes & side strips E 50… 300 kV/m
MCP (Micro Channel Plate)
electron converter & 106-fold amplifier
either Phosphor screen & CCD
 high spatial resolution of 100 m
or wire array down to 250 m pitch
 high time resolution
IPMs are installed in nearly all synchrotrons
However, no ‘standard’ realization exists!
CCD
Phosphor
MCP 2
MCP 1
Light
Electrons
Channels 10 m
Residual gas ion
Ion beam

## Slide 5
Ionization Profile Monitor Realization at GSI Storage Ring
The realization for the heavy ion storage ring ESR at GSI:
Horizontal camera
Horizontal IPM:
E-field box
MCP
IPM support
& UV lamp
Ø250 mm
beam
E-field separation disks
View port Ø150 mm
Electrodes
Insertion 650 mm
175mm
Vertical IPM
Vertical camera

## Slide 6
IPM: Multi Channel Plate MCP for Synchrotron Installation
MCP are used as particle detectors with secondary electron amplification.
A MCP is:
1 mm glass plate with 10 μm holes
thin Cr-Ni layer on surface
voltage 1 kV/plate across
 e− amplification of  103 per plate.
 resolution  0.1 mm (2 MCPs)
Anode technologies:
SEM-grid,  0.5 mm spacing
 limited resolution
 fast electronics readout
phosphor screen + CCD
 high resolution, but slow timing
 fast readout by photo-multipliers
20 m
Electron microscope image:
Challenges:
Fast readout with < 100 ns resolution
Proper MCP holder design
Calibration for sensitivity correction
HV switching of MCP to prevent for destruction

## Slide 7
Example:
U73+ beam at GSI for intensity increase
stacking by electron cooling
and acc. 11.4  400 MeV/u
IPM: Observation of Cooling and Stacking
Task for IPM:
Observation of cooling
Emittance evaluation during cycle
P. Forck (GSI) et al., DIPAC’05
| 5 injections + cooling | | acc. |
horizontal
V. Kamerdzhiev (FZJ) et al., IPAC’11
IPM: Profile recording every 10 ms
measurement within one cycle .

## Slide 8
Important application:
Injection matching
to prevent for emittance enlargement
Observation during ‘bunch gymnastics’
turn–by-turn measurement
Required time resolution  100 ns
Example: Injection to J-PARC RCS at 0.4 GeV Anode: wire array with 1mm pitch
IPM: Turn-by-Turn Measurement
H. Hotchi (J-PARC), HB’08, A Satou (J-PARC) et al., EPAC’08
Turn-by-turn IPMs at
BNL, CERN, FNAL etc.
Not realized at GSI yet!
-40 -20 0 20 40
-40 -20 0 20 40
un-matched
matched
1st turn
9th turn

## Slide 9
IPM: Space Charge Influence for Intense Beams
Ion detection: For intense beams
 broadening due to space charge
Electron detection:
B-field required for e- guidance toward MCP.
Effects: 3-dim start velocity of electrons
Ekin(90%) < 50 eV, max  900
 rcyl < 100 m for B  0.1 T
Monte-Carlo simulation:
Ion versus e- detection
1012 charges
 Only e- scheme
gives correct image
B-field & electron detection
ion detection

## Slide 10
IPM: Magnet Design
Design by G. de Villiers (iThemba Lab), T. Giacomini (GSI)
Further types of magnets e.g. K.Satou (J-PARC) et al., EPAC’08, J.Zagel (FNAL) et al., PAC’01,
R.Connolly (RHIC) et al., PAC’01, C. Fischer (CERN) et al. BIW’04
Maximum image distortion:
5% of beam width  B/B < 1 %
Challenges:
High B-field homogeneity of 1%
Clearance up to 500 mm
Corrector magnets required
to compensate beam steering
Insertion length 2.5 m incl. correctors
For MCP wire-array readout
lower clearance required
Magnetic field for electron guidance:
Corrector
480mm
Corrector
Horizontal IPM
Vertical IPM
Insertion length
2.5 m
300mm
At transfer line:
Vacuum pressure up to 10-5 mbar
IPM without MCP realized
much less mechanical efforts

## Slide 11
IPM: Magnet Design
Design by G. de Villiers (iThemba Lab), T. Giacomini (GSI)
Further types of magnets e.g. K.Satou (J-PARC) et al., EPAC’08, J.Zagel (FNAL) et al., PAC’01,
R.Connolly (RHIC) et al., PAC’01, C. Fischer (CERN) et al. BIW’04
Maximum image distortion:
5% of beam width  B/B < 1 %
Magnetic field for electron guidance:
Corrector
480mm
Corrector
Horizontal IPM
Vertical IPM
Insertion length
2.5 m
300mm

## Slide 12
Status:
Non-destructive method in operation in nearly all hadron synchrotrons
Proposed or operated in some hadron LINACs (often without MCP)
Physics well understood
For high beam current i.e. high space charge field magnet B  0.1 T required
 long insertion length
MCP efficiency drops significantly during high current operation
 efficiency calibration & HV switching required
Challenges (no standard realization exists) :
High voltage (up to 60 kV) realization for intense beams
Stable operation for MCP incl. efficiency calibration
Design and tests for correction algorithm for space charge broadening
Remark:
Gas curtain monitor with well localized gas volume realized
Comparable device used for synchrotron light monitor realized
Summary Ionization Profile Monitor

## Slide 13
Beam Induced Fluorescence Monitor: Principle
Ion beam
Blackened walls
Vacuum gauge
Valve
Viewport
150mm flange
Lens, Image-Intensifier
and CCD FireWire-Camera
N2-fluorescent gas
equally distributed
Detecting photons from residual gas molecules, e.g. Nitrogen
N2 + Ion  (N2+)* + Ion  N2+ +  + Ion
390 nm< < 470 nm
emitted into solid angle  to camera
single photon detection scheme
Features:
Single pulse observation possible
down to 1 s time resolution
High resolution (here 0.2 mm/pixel)
can be easily matched to application
Commercial Image Intensifier
Less installations inside vacuum as for IPM
 compact installation e.g. 20 cm for both panes
Beam: 4x1010 Xe48+ at 200MeV/u, p=10-3 mbar

## Slide 14
BIF-Monitor: Technical Realization at GSI LINAC
Beam
Vertical BIF
Image Int. CCD
Horizontal BIF
Photocathode
Phosphor
double
MCP

many 
e-
Image intensifier
Six BIF stations at GSI-LINAC (length 200m):
2 x image intensified CCD cameras each
double MCP (‘Chevron geometry’)
Optics with reproduction scale 0.2 mm/pixel
Gas inlet + vacuum gauge
Pneumatic actuator for calibration
Insertion length 25 cm for both directions only
Advantage: single macro-pulse observation
F. Becker (GSI) et al., Proc. DIPAC’07, C. Andre (GSI) et al., Proc. DIPAC’11

## Slide 15
Image from 1·109 U
p= 2·10-3 mbar,
mounted ≈ 2 m
before beam-dump:
Ekin dependence for signal
& background close to beam-dump:
60 MeV/u
350 MeV/u
750 MeV/u
viewport
Background prop. Ekin2  shielding required
Background suppression by 1 m fiber bundle
Signal proportional to energy loss
Suited for FAIR-HEBT with ≥ 1010 ions/pulse
Energy Scaling behind SIS18 at GSI
F. Becker (GSI) et al., Proc. DIPAC’07

## Slide 16
Results of detailed investigations:
Rare gases and N2: green to near-UV
Compact wavelength interval for N2
Fluorescence yield: N2  4x higher as rare gases
 N2 and Xe are well suited !
gas | Y for p | Y for p/ne
Xe | 86 % | 22 %
Kr | 63 % | 25 %
Ar | 38 % | 30 %
He | 4 % | 26 %
N2 | 100 % | 100 %
Relative fluorescence yield Y (all wavelength):
ne: gas electron density  energy loss  beam influence
BIF-Monitor: Spectroscopy – Fluorescence Yield
F. Becker (GSI) et al., Proc. DIPAC’09, Collaboration with TU-München
Beam: S6+ at 5.16 MeV/u, pN2 =10-3 mbar

## Slide 17
Normalized profile reading for all :
Profile reading equal
for all gases except He
BIF-Monitor: Spectroscopy – Profile Reading
Results of detailed investigations:
Rare gases and N2: green to near-UV
Compact wavelength interval for N2
Fluorescence yield: N2  4x higher as rare gases
Same profile reading for all gas except He
 N2 and Xe are well suited !
F. Becker (GSI) et al., Proc. DIPAC’09, Collaboration with TU-München
Beam: S6+ at 5.16 MeV/u, pN2 =10-3 mbar

## Slide 18
For N2 working gas the spectra for different ion impact is measured:
Spectroscopy – Excitation by different Ions
Results:
Comparable spectra for all ions
Small modification due to N2+
dissociation by heavy ion impact
Results fits to measurements
for proton up to 100 GeV at CERN
Stable operation possible for N2
Care: Different physics for Ekin < 100 keV/u  vcoll < v Bohr
 Different spectra measured
M. Plum et al., NIM A (2002) & A. Variola, R. Jung, G. Ferioli, Phys. Rev. Acc. Beams (2007),

## Slide 19
Observation: Trans. of ionic states e.g. N2+  profile width independent on pressure
Trans. of neutral states e.g. N2  width strongly dependent on pressure!
Ionic transitions =391 nm:
N2 + ion(N2+)* +e-+ ionN2++ +e- + ion
N2+ @391nm: B2+u(v=0)  X2+g(v=0)
large σ for ion-excitation, low for e-
N2+ trans.@391 nm
N2 trans.
@337 nm
p = 0.1 mbar
N2
p = 30 mbar
N2
F. Becker et al., IPAC’12 &HB’12
Neutral transitions =337 nm:
N2 + e-  (N2)* + e-  N2 +  + e-
N2 @337nm: C3u(v=0)  B3g(v=0)
large σ of e- excitation., low for ions
at p  0.1 mbar free mean path 1 cm!
N2
p = 0.003 mbar
Image Spectroscopy – Different Gas Pressures and Profile Width

## Slide 20
Beam: S at 3 MeV/u at TU-München TANDEM
100mm
10-2 mbar
rmfp~30 mm
10-1 mbar
rmfp~ 3 mm
30mm
10+1 mbar
rmfp~ 30 m
Image Spectroscopy – Different Gas Pressures and total Profile Width
F. Becker et al., IPAC’12 and HB’12
Entire spectral range  effect is smaller
but significant disturbance for He and Ne
Task: To which pressure the methods delivers
a correct profile reproduction?
Results:
avoid 10-2 mbar < p < 10 mbar 
chose either rmfp >> rbeam or rmfp<< rbeam
use transition of the charged specious
all transitions

## Slide 21
Alternative Single Photon Camera: emCCD
Principle of electron multiplication CCD:
Multiplication by avalanche diodes:
Parameter of Hamamatsu C9100-13
Pixel: 512x512, size16x16m2 , -80 OC
Maximum amplification: x1200
Readout noise: about 1 e- per pixel
Results: Suited for single photon detection
x5 higher spatial resolution as ICCD
less beam-induced background
more noise due to electrical amplification
 Acts as an alternative
F. Becker et al., BIW’08
I= 60 A Ni13+:
tpulse = 1.2 ms
p=10-4 mbar

## Slide 22
Non-destructive profile method in operation for E < 11 MeV/u for typ. p < 10-5 mbar
Considered for higher beam energies E > 100 MeV/u ongoing
Independence of profile reading for pressures up to 10-2 mbar for N2, Xe, Kr, Ar
N2 is well suited: blue wavelength, high light yield, good vacuum properties
Xe is an alternative due to 10-fold shorter lifetime: less influence in beam’s E-field
He is excluded as working gas due to wrong profile reproduction
Modern emCCD might be an alternative
Topics under development:
Investigation of shielding and radiation hardness of components
Modeling of atomics physics processes for different pressure ranges
Generally: Method proposed or used for:
High current hadron LINAC (e.g. LIPAc, FRANZ, IPHI.....)
Proton synchtrotrons (e.g. CERN...)
Electron sources, LINACs and e-coolers (e.g. Uni-Mainz...)
Summary Beam Induced Fluorescence Monitor

## Slide 23
Beam: 1.1 mA Xe21+, 4.7 MeV/u
Comparison BIF  IPM at GSI LINAC with 4.7 MeV/u Xe21+
Collaboration with J. Egberts, J. Marroncle, T. Papaevangelou CEA/Saclay
J. Egberts (CEA) et al., DIPAC’11, F. Becker (GSI) et al , DIPAC’11
Test with LIPAc design and various beams
Comparison IPM without MCP and BIF
Advantage IPM: 10 x lower threshold as BIF
Disadvantage IPM: Complex vacuum installation,
image broadening by beam’s space charge
Design by CEA
for LIPAc

## Slide 24
Collaboration with J. Egberts, J. Marroncle, T. Papaevangelou CEA/Saclay
J. Egberts (CEA) et al., DIPAC’11, F. Becker (GSI) et al , DIPAC’11
Variation of Helium gas pressure:
Profile broadening for both detectors
Large effect for BIF (emission of photons)
Comparison to SEM-Grid and BIF
Helium is not suited as working gas for BIF & IPM
Beam: 1.1 mA Xe21+, 4.7 MeV/u
Comparison BIF  IPM for He Gas
Design by CEA
for LIPAc

## Slide 25
Simplified Comparison of BIF and IPM Method
 | BIF | IPM
Signal source | γ from residual gas  Low solid angle Ω 10-4 | e- from residual gas Large Ω = 4π due to E-field
Detector principle | γ  e-  108 γ by MCP & Phosphor & CCD | γ  e-  108 γ by MCP & Phosphor & CCD or MCP & I/U converter & ADC
Advantage | Non-destructive Nearly no mechanics | Non-destructive Medium signal strength
Disadvantage | Low signal strength Might need gas inlet Smaller space charge influence for Xe | Complex device Expensive For high currents: Magnet required
Main Application | High current at LINAC No well suited for super-cond. LINAC Target diagnostics | Synchrotons
Thank you for your attention !
Comparison for application at high current hadron LINAC, transport lines & synchrotrons