## Slide 1
III ARDENT
Annual Workshop
29/09-02/10/2014
IBA Dosimetry
Francesca Bisello
ESR 10

## Slide 2
Outline
Research Activities
Conferences
Training and Secondments
I. Research Activities
II. Conferences
III. Trainings and Secondments
Outline

## Slide 3
Research
Activities

## Slide 4: Dosepix Characterization General overview
Outline
Research Activities
Conferences
Training and Secondments
Dosepix detector as Quality Assurance tool in diagnostics

## Slide 5: Dosepix Characterization kVp-meter in mammography
Outline
Research Activities
Conferences
Training and Secondments

## Slide 6: Dosepix Characterization Sum up of activities
Outline
Research Activities
Conferences
Training and Secondments
October 2013 – October 2014:
Dosepix as kVp-meter in mammography
Radiation Tolerance studies
Temperature characterization

## Slide 7: Dosepix Characterization kVp-meter in mammography
Outline
Research Activities
Conferences
Training and Secondments
Image quality and dose depend on the xRay tube voltage
Tube setting:
30 kV, 20 mAs
Mo-Mo
Big Pixels
Tube setting:
30 kV, 20 mAs
Mo-Mo

## Slide 8: Dosepix Characterization kVp-meter in mammography
Outline
Research Activities
Conferences
Training and Secondments
kVp: fitting the end point of the spectrum
Chart: kVp MagicMax (kV)
| 22 | 22.9 | 23.8 | 24.8 | 25.7 | 26.6 | 27.5 | 28.4 | 29.3 | 30.09 | 30.9 | 31.8 | 32.9 | 33.799999999999997
 | 23.5521 | 23.5723 | 24.6581 | 25.7655 | 26.524 | 27.7469 | 28.644 | 29.263 | 29.954 | 30.7359 | 31.6135 | 32.2765 | 33.0071 | 33.4796
Chart: kVp (MagicMAx) (KV)
| 22 | 22.9 | 23.8 | 24.8 | 25.7 | 26.6 | 27.5 | 28.4 | 29.3 | 30.09 | 30.9 | 31.8 | 32.9 | 33.799999999999997
 | 1.5521 | 0.67229 | 0.85808 | 0.96547 | 0.82405 | 1.1469 | 1.144 | 0.86299 | 0.65401 | 0.64587 | 0.71349 | 0.47646 | 0.10711 | 0.32036
Additional filtration: 0.3mm Cu
Counts/big pixel/s: 1.5 · 104 Hz [ count rate without filtration : 5· 104 Hz ]

## Slide 9: Dosepix Characterization RadiationTolerance studies
Outline
Research Activities
Conferences
Training and Secondments
10-28/02/2014
Secondment at CERN
Radiation Tolerance Studies
Measurement Session:
Accumulated dose for the test:
5 kGy in step of 200 Gy
Tube settings:
40 kV, 2 mA dose rate 4.83 Gy/min

## Slide 10: Dosepix Characterization Radiation Tolerance studies
Outline
Research Activities
Conferences
Training and Secondments

## Slide 11: Dosepix Characterization Radiation Tolerance studies
Outline
Research Activities
Conferences
Training and Secondments
Poster Presention IEEE 2014, 8-15 November 2014
“Investigation on the radiation tolerance of Dosepix detector”
F. Bisello, W.S. Wong , E.Frojdh , J.Alozy I. Ritter, A.Zang, M. Campbell, J.C.Celi, G.Anton and T.Michel
  Deeper analysis of effect around mammographic energies
  Investigations in damages in the sensor or in the chip

## Slide 12: Dosepix Characterization Temperature Characterization
Outline
Research Activities
Conferences
Training and Secondments
First Measurement Session: CERN (secondment)
Addidtiona Investigations: IBA Dosimetry
Measurement Session:
Range required by the regulation: 15°- 35 °C
Instability < 1%

## Slide 13: Dosepix Characterization Temperature Characterization
Outline
Research Activities
Conferences
Training and Secondments

## Slide 14: Dosepix Characterization Temperature Characterization
Outline
Research Activities
Conferences
Training and Secondments
Source of instability (I):
PLL circuit
Source of instability (II):
Bias current of initial gain stage of the discriminator

## Slide 15: Dosepix Characterization Temperature Characterization
Outline
Research Activities
Conferences
Training and Secondments
Next steps:
Evaluate the possibility to bypass the PLL circuit
Find other more stable configurations for the analog front end

## Slide 16: Other Activities SFD diodes
Outline
Research Activities
Conferences
Training and Secondments
Machine and Plan QA
Small Field
external Radiotherapy

## Slide 17: Other Activities SFD diodes
Outline
Research Activities
Conferences
Training and Secondments
Studies on sensitivity dependence from:
Dose
Dose per pulse
Dose rate

## Slide 18: Other Activities SFD diodes
Outline
Research Activities
Conferences
Training and Secondments
Oral Presentation, RESMDD14, 8-10 October 2014
“Development of silicon monolithic arrays for dosimetry in external
beam radiotherapy”
F. Bisello, M. Bucciolini, M. Bruzzi, D. Menichelli, M. Scaringella, C. Talamonti, M. Zani1,2

## Slide 19
Conference
Trainings
Secondments

## Slide 20: Conferences 2014
Outline
Research Activities
Conferences
Training and Secondments
DPG Mainz, 25th March 2014
Oral Presentation
“Characterization of the Dosepix detector under clinical conditions ”
RESMDD Florence, 8-10 October 2014
Oral Presentation
“Development of silicon monolithic arrays for dosimetry in external beam radiotherapy”
IEEE 2014 Seattle, 8-15 November 2014
Poster Presentation
“Investigation on the radiation tolerance of Dosepix detector”

## Slide 21: Planned Training and Secondments
Outline
Research Activities
Conferences
Training and Secondments
Business and Administration training, ST Microelectronics
Secondment at the UPENN University Hospital, Phyladelphia
(tentative: 16-28 November 2014):
  Measurements on plan QA with silicon monolithic array
  Trainings on clinical quality assurance routine (to be defined)

## Slide 22
Thank you
very much
for your attention