## Slide 1: Results (and Problems) of PHIN Run Summer 2013
Christoph Hessler, Eric Chevallay, Steffen Doebert, Valentin Fedosseev, Irene Martini, Mikhail Martyanov
CLIC/CTF3 experimental verification meeting 04.03.2014

## Slide 2: Problems (and Results) of PHIN Run Summer 2013
Christoph Hessler, Eric Chevallay, Steffen Doebert, Valentin Fedosseev, Irene Martini, Mikhail Martyanov
CLIC/CTF3 experimental verification meeting 04.03.2014

## Slide 3: Original Plans
Lifetime studies with Cs3Sb cathode and green light
  Under improved vacuum conditions (new NEG pump)
  With 5 µs (2.5 µs) long pulse trains
  With 5 Hz operation
Feedback stability tests
Dark current studies
  Measurements with Cs3Sb and Cu cathodes
  Try to see field emission spots with camera looking into the gun

## Slide 4: Original Plans
Lifetime studies with Cs3Sb cathode and green light
  Under improved vacuum conditions (new NEG pump)
  With 5 µs (2.5 µs) long pulse trains
  With 5 Hz operation
Feedback stability tests
Dark current studies
  Measurements with Cs3Sb and Cu cathodes
  Try to see field emission spots with camera looking into the gun

## Slide 5: Layout of PHIN
FCT: Fast current transformer VM: Vacuum mirror SM: Steering magnet BPM: Beam position monitor MSM: Multi-slit Mask OTR: Optical transition radiation screen
MTV: Gated cameras SD: Segmented dump FC: Faraday cup

## Slide 6: Available Beam Instrumentation (1)
For the long-pulse tests, the Faraday cup (FC) was separated from the vacuum system and operated in air. A vacuum window was installed.
As a consequence the FC was no longer 50  terminated.
During the preparation works a leak at the insulator glass tube, needed for the fast current transformer (FCT), was detected.
As a consequence the FCT has been replaced by an old wall current monitor (WCM) with different bandwidth.
 Only “unknown” instruments for current / charge measurement were available!

## Slide 7: Available Beam Instrumentation (2)
BTV screen CT.MTV0265 was “blind” during PHIN run 2012 due to residual green laser light, back reflected from the cathode.
To solve this issue a notch filter for 532 nm was installed in front of the Proxitronic camera:
Some residual green light was also visible on screen CT.MTV0305 for high intensities, but this was less critical.
Installation of notch filter for 2nd screen for next PHIN run.
with notch filter
without notch filter

## Slide 8: Virtual Cathode Camera
Camera has been replaced at the beginning by one, which was directly placed in the image plane.
However, the field of view was to small for the range in which the beam position can move.
It has then be replaced again by the old camera in the old configuration (focused to the image plane).
However, the old camera shows strange horizontal lines and should be replaced for the next run:

## Slide 9: Vacuum Window and Protection (1)
In 2012 strong pressure increase due to Faraday cup heating was observed for relatively high beam intensity.
Maximum achievable beam intensity: 5 µs long pulse trains with 2.7 nC/bunch (limited by klystron power) → 1.3E14 electrons/train → 500 W maximum beam power
Intensity too high for uncooled Faraday cup!
Operation of FC in air, separated by exit window from beam line.
Fluka studies showed, that only Beryllium can withstand maximum beam intensity at minimum achievable beam size. (Thanks to M. Brugger and M. Delonca)
Cathode #189 (PHIN run 2012)
Courtesy:
M. Delonca
Pressure increase to 2e-7 mbar

## Slide 10: Vacuum Window and Protection (2)
Disadvantages of Beryllium:
  Toxicity
  Long lead times
  Carbon-composite support plate required with unknown and potentially harmful effect (?) on Cs3Sb cathodes. To be investigated more in detail. ↔ In conflict with actions to improve vacuum level.
Therefore, a stainless steel window was chosen for this PHIN run, which requires a reduced beam intensity.
Supervision of window temperature with infrared camera.
Bunch charge 2.3 nC Train length 1000 ns

## Slide 11: FC Measurements
Problems:
  Ringing on the FC signal: Due to separation from the vacuum system FC is not 50 Ω terminated anymore.
  Poor transmission WCM -> FC, especially for higher bunch charges
Trials to improve the signal:
  Add flange with a hole at the entrance of FC
  FC moved as close as possible to window (loss of temperature control with IR camera)
  Connection of FC to beam line with cable and Cu foil
  Replacement of FC with Cu plate
Conclusion: No solution without ringing and good transmission found and simultaneous measurement of the window temperature.
Initial situation 24.7.
Cu plate 8.8.2013
FC moved 25.7.
FC signal
FC signal
FC signal

## Slide 12: FC: Possible Solutions?
Possible explanations for the poor transmission:
  Original FC in air: Scattering of electrons in the window might increase the size of the beam, which then does not fully hit the FC.
  Cu plate: Due to the larger size of the plate more electrons hit the plate, but back scattered and secondary electrons can leave the FC easier.
Possible solutions:
  Built new FC with same geometry as original FC but much bigger size. But will be still difficult (and maybe not possible) to have it in a sufficient distance to be able to observe the window with the IR camera.
  FC with repeller electrode.
  Installation of a second FCT right before or after the window to replace the FC as current measurement device. Effect of back-scattered electrons from the window on FCT?
Solutions to avoid ringing? → Maybe we can live with the ringing?

## Slide 13: WCM Measurements
In the past an FCT with 1.6 kHz – 1 GHz bandwidth was used and integral of signal was measured.
The WCM used in this PHIN run has a bandwidth of 1 MHz – 10 GHz and is therefore not suitable for measuring long pulses. As a work-around the amplitude was measured instead of the integral.
FCT signal
WCM signal

## Slide 14: Improvement of Vacuum in PHIN
1/e lifetime 185 h
1 nC, 800 ns, l=524 nm, Cs3Sb
1/e lifetime 26 h
1 nC, 800 ns, l=262 nm, Cs3Sb
7e-10 mbar
1.3e-10 mbar
Dynamic vacuum level:
4e-9 mbar
Static vacuum level:
2.2e-10 mbar
March 2012
March 2011
Activation of NEG chamber around gun
Installation of additional NEG pump
<2e-10 mbar
2.4e-11 mbar
July 2013
1/e lifetime n/a
due to different
problems during PHIN run

## Slide 15: Improvement of Vacuum
New RGA installed
RGA does not work during beam operation due to disturbance caused by the magnetic fields of the solenoids
RGA spectrum

## Slide 16: Problems with Lifetime Measurements
Despite good QE in DC gun, photocathodes had low QE on arrival in PHIN gun
Despite good dynamic base vacuum level there were a lot of breakdowns causing desorption spikes:
Rapidly decreasing QE, (partially) recovering during “QE jumps”
24h oscillations on beam current
Frequent loss of synchronization of laser system
Cathode #189 (PHIN run 2012)
Cathode #191 (PHIN run 2013)

## Slide 17: Photocathodes Used during PHIN Run
Number | Material | Age | QE in DC gun | QE in PHIN
#188 | Cs3Sb | Old cathode, previously used for studies in photoemission lab | 0.24% 21.03.2013 | 0.33%
#193 | Cs3Sb | New cathode, produced 28.02.2013 | 5.8% after preparation | 1%
#191 | Cs3Sb | New cathode, produced 12.02.2013 | 7.5% after preparation | 1%
#194 | Cs3Sb | Fresh, new cathode, produced 22.08.2013 on Cu plug previously used for RF conditioning at PHIN. | 4.5% after preparation | 1.3%
#186 | Cs2Te | Old Cs2Te cathode, but with good QE | 9.0% 03.05.2013 (QE not uniform) | 6.5%
Number | Material | Age | QE in DC gun
#188 | Cs3Sb | Old cathode, previously used for studies in photoemission lab | 0.24% 21.03.2013
#193 | Cs3Sb | New cathode, produced 28.02.2013 | 5.8% after preparation
#191 | Cs3Sb | New cathode, produced 12.02.2013 | 7.5% after preparation
#194 | Cs3Sb | Fresh, new cathode, produced 22.08.2013 on Cu plug previously used for RF conditioning at PHIN. | 4.5% after preparation
#186 | Cs2Te | Old Cs2Te cathode, but with good QE | 9.0% 03.05.2013 (QE not uniform)

## Slide 18: Cathode #188
1st Cathode used during PHIN run 2013.
Surface treatment: Diamond powder polishing.

## Slide 19: Cathode #193
2nd Cathode used during PHIN run 2013.
Surface treatment: Diamond turning.

## Slide 20: Cathode #189 (PHIN run 2012)
Cathode used during PHIN run 2012 with long lifetime.
Surface treatment: Diamond powder polishing.

## Slide 21: Fresh Copper Plugs
Surface treatment: Diamond powder polishing.
Surface treatment:
Diamond turned.

## Slide 22: Photocathodes Used during PHIN Run
Number | Material | Age | QE in DC gun | QE in PHIN
#188 | Cs3Sb | Old cathode, previously used for studies in photoemission lab | 0.24% 21.03.2013 | 0.33%
#193 | Cs3Sb | New cathode, produced 28.02.2013 | 5.8% after preparation | 1%
#191 | Cs3Sb | New cathode, produced 12.02.2013 | 7.5% after preparation | 1%
#194 | Cs3Sb | Fresh, new cathode, produced 22.08.2013 on Cu plug previously used for RF conditioning at PHIN. | 4.5% after preparation | 1.3%
#186 | Cs2Te | Old Cs2Te cathode, but with good QE | 9.0% 03.05.2013 (QE not uniform) | 6.5%
Number | Material | Age | QE in DC gun | QE in PHIN | Surface treatment
#188 | Cs3Sb | Old cathode, previously used for studies in photoemission lab | 0.24% 21.03.2013 | 0.33% | Diamond powder polishing
#193 | Cs3Sb | New cathode, produced 28.02.2013 | 5.8% after preparation | 1% | Diamond turning
#191 | Cs3Sb | New cathode, produced 12.02.2013 | 7.5% after preparation | 1% | Diamond turning
#194 | Cs3Sb | Fresh, new cathode, produced 22.08.2013 on Cu plug previously used for RF conditioning at PHIN. | 4.5% after preparation | 1.3% | Diamond powder polishing + RF conditioning
#186 | Cs2Te | Old Cs2Te cathode, but with good QE | 9.0% 03.05.2013 (QE not uniform) | 6.5% | Diamond powder polishing
Number | Material | Age | QE in DC gun | QE in PHIN | Surface treatment | Layer thickness
#188 | Cs3Sb | Old cathode, previously used for studies in photoemission lab | 0.24% 21.03.2013 | 0.33% | Diamond powder polishing | 170 nm
#193 | Cs3Sb | New cathode, produced 28.02.2013 | 5.8% after preparation | 1% | Diamond turning | 19 nm
#191 | Cs3Sb | New cathode, produced 12.02.2013 | 7.5% after preparation | 1% | Diamond turning | 171 nm
#194 | Cs3Sb | Fresh, new cathode, produced 22.08.2013 on Cu plug previously used for RF conditioning at PHIN. | 4.5% after preparation | 1.3% | Diamond powder polishing + RF conditioning | 39 nm
#186 | Cs2Te | Old Cs2Te cathode, but with good QE | 9.0% 03.05.2013 (QE not uniform) | 6.5% | Diamond powder polishing | 215 nm
Could be explained by the fact that pumping time of SAS of transport carrier was not sufficient after its reconnection to PHIN

## Slide 23: RF Break-downs
#188
#193
#191
#194
Diamond
turning
Diamond
turning
Diamond
powder
polishing
Diamond
powder
polishing
Significantly less break-downs
2.1 break-downs/h
10.6 break-downs/h
5.7 break-downs/h
5.2 break-downs/h

## Slide 24: Quantum Efficiency Jumps
Cathode #191, Cs3Sb, train length 800 ns, bunch charge 1 nC
Sudden improvement of quantum efficiency
(“QE jumps”)
→ Not yet understood!
Laser energy was regulated to keep the charge constant at 1 nC

## Slide 25: 24h Oscillations on Beam Current
Klystron phase was kept constant
24h oscillations probably caused by non-working air-conditioning in klystron gallery:
→ Temperature variation of RF reference signal cable for laser
→ Phase shift of laser with respect to klystron
Laser energy was at its limit and could not be increased further
Iris opened, more laser energy available

## Slide 26: Loss of Synchronisation of the Laser System
Frequent loss of synchronisation of the laser system at the end of the PHIN run
Incoming RF power for the laser synchronisation unit was outside the recommended range (-20 to 0 dBm). An increase of the RF power to -16 dBm did not help.
Investigations showed that there was a back reflection from a lens after Amp 1, which caused the oscillator to lose the synchronisation

## Slide 27: Magnet powering / Beam Focusing
Even with a small laser beam, it was difficult to get a reasonable electron beam size on the OTR screen.
The second solenoid had to be powered almost at its limit.
12.08.2013
250 A
200 A

## Slide 28: Magnet powering / Beam Focusing
For comparison PHIN run September 2011:
Parameters: Phase: 60 deg CF.SNH0110: 190 CF.SNJ0130: 195 Klystron voltage: 38.50kV
07.09.2011

## Slide 29: Original Plans
Lifetime studies with Cs3Sb cathode and green light
  Under improved vacuum conditions (new NEG pump)
  With 5 µs (2.5 µs) long pulse trains
  With 5 Hz operation
Feedback stability tests
Dark current studies
  Measurements with Cs3Sb and Cu cathodes
  Try to see field emission spots with camera looking into the gun

## Slide 30: Feedback Stabilisation Scheme

## Slide 31: Feedback Stabilisation Tests
Cathode #193
Feedback OFF
Cathode #188
Feedback ON
Feedback OFF
1.34 % rms
factor 3
improvement !
3.03% rms
0.97% rms
Feedback ON
0.43 % rms
Laser beam
Electron beam

## Slide 32: Original Plans
Lifetime studies with Cs3Sb cathode and green light
  Under improved vacuum conditions (new NEG pump)
  With 5 µs (2.5 µs) long pulse trains
  With 5 Hz operation
Feedback stability tests
Dark current studies
  Measurements with Cs3Sb, Cs2Te and Cu cathodes
  Try to see field emission spots with camera looking into the gun

## Slide 33: Dark Current Measurements
Measurement performed with a 1.3 µs long klystron pulse
Dark current by a factor ~5 higher in Cs3Sb than in Cs2Te
b by factor ~2 smaller for Cs3Sb than for Cs2Te or copper

## Slide 34: Observation of Field Emission Spots?
RF off

## Slide 35: Observation of Field Emission Spots?
RF on
No field emission spots visible, but…

## Slide 36: Observation of Field Emission Spots?
Break-downs

## Slide 37: Observation of Field Emission Spots?
Break-downs

## Slide 38: Observation of Field Emission Spots?
Break-downs

## Slide 39: Observation of Field Emission Spots?
Break-downs

## Slide 40: Observation of Field Emission Spots?
Break-downs

## Slide 41: Observation of Field Emission Spots?
Break-downs

## Slide 42: Observation of Field Emission Spots?
Break-downs

## Slide 43: Observation of Field Emission Spots?
Break-downs

## Slide 44: Observation of Field Emission Spots?
Break-downs

## Slide 45: Observation of Field Emission Spots?
Break-downs

## Slide 46: Observation of Field Emission Spots?
Break-downs

## Slide 47: Observation of Field Emission Spots?
Break-downs

## Slide 48: Plans for PHIN, Photoemission Lab and Laser System
PHIN:
Excellent vacuum conditions have been achieved → Lifetime studies under these conditions with Cs3Sb and Cs2Te cathodes
Study the effect of surface roughness on cathode lifetime → Electro-polished cathode plugs
Continue studies of Cs3Sb cathodes using a green laser beam
Studies with 5 µs long pulse trains and 5 Hz repetition rate → Be window needed → separate PHIN run required
Photoemission lab:
Continue high-charge studies (with different cathode layer thicknesses)
XPS studies
Upgrade of cathode preparation system with load-lock system
Laser:
Study of performance of laser system with CLIC specs (new 500 MHz front end, 50 Hz tests)
Continue stability studies with new front end
PHIN run
2014
PHIN run
2015

## Slide 49: Prerequisites for next PHIN Run
Possibility to easily switch to 5 Hz for RF conditioning.
Stable air-conditioning klystron gallery and laser lab
Check of solenoid resistance
Replacement insulator piece of FCT
Installation additional FCT at the end of the beam line
Re-design FC? Or re-connection to beam line for short pulse studies? Add additional ion pump?
Setup for 2D QE scan at PHIN

## Slide 50: Acknowledgement
Controls: Sergio Batuca, Mark Butcher, Mathieu Donze, Alessandro Masi, Christophe Mitifiot
Load-lock system: Szymon Sroka
FLUKA simulations: Markus Brugger, Melanie Delonca
Beam instrumentation: Benoit Bolzon, Thibaut Lefevre
Vacuum: Paolo Chiggiato, Berthold Jenninger, Esa Paju
RF: Stephane Curt, Luca Timeo
Wilfrid Farabolini
XPS studies: Holger Neupert, Elise Usureau
Collaborators at LAL and IAP-RAS
… and many others
… and thank you for your attention!

## Slide 52: Failure of Air Conditioning in Laser Lab
Air conditioning in laser lab failed on 15.08.2013
No laser operation possible for a couple of days
Temperature evolution:

## Slide 53: PHIN and CLIC Parameters
Parameter | PHIN | CLIC
Charge / bunch (nC) | 2.3 | 8.4
Macro pulse length (μs) | 1.2 | 140
Bunch spacing (ns) | 0.66 | 2.0
Bunch rep. rate (GHz) | 1.5 | 0.5
Number of bunches / macro pulse | 1800 | 70000
Macro pulse rep. rate (Hz) | 5 | 50
Charge / macro pulse (μC) | 4.1 | 590
Beam current / macro pulse (A) | 3.4 | 4.2
Bunch length (ps) | 10 | 10
Charge stability | <0.25% | <0.1%
Cathode lifetime (h) at QE > 3% (Cs2Te) | >50 | >150
Norm. emittance (μm) | <25 | <100