## Slide 1: A new timing calibration method for switched capacitor array chips to achieve sub-picosecond resolutions
Stefan Ritt, Paul Scherrer Institute, Switzerland

## Slide 2: Overview
Limitations of timing resolutions in Switched Capacitor Arrays (SCA)
New method for SCA calibration
Results

## Slide 3: Schematic Overview (DRS4)

## Slide 4: Noise limited time accuracy
1
2
3

## Slide 5: A few examples
U [mV] | DU[mV] | tr [ns] | f3dB [MHz] | fS [GSPS] | Dt [ps]
10 | 1 | 1 | 333 | 3 | 60
100 | 1.8 | 1 | 333 | 2 | 13
100 | 0.65 | 1 | 333 | 3 | 3.8
100 | 0.7 | 0.23 | 1500 | 10 | 1.9
100 | 0.35 | 0.36 | 950 | 5 | 1
SAM
DRS4
TARGET
PSEC4
Small (<20 mV) signals give poor timing
Theoretical limit for TARGET/SAM/PSEC4 is reached
Limit for DRS4 is lower than current resolution ( O(20ps) )
What is the reason for this?

## Slide 6: Time definition
single inverter delay: Dti → differential nonlinearity
Sum of single inverter delay: Dta,b = SDti → integral nonlinearity

## Slide 7: Local Calibration
Credits: D. Stricker-Shaver, Univ. Tübingen, NSS/MIC 2012 proceedings

## Slide 8: Local Calibration

## Slide 9: How to calibrate?
External function generator
  Not convenient for re-calibration
  Arbitrary waveform generation has ~20 ps jitter
FGPA: LVDS output with slew rate limitation
  Anything coming from an FPGA has 50-100 ps jitter
Dedicated quartz oscillator
  Can be on-board
  Generate sine wave via passive low-pass filter
  Only use region around zero as linear slope

## Slide 10: Oscillator
OSC.
BUF
Low
Pass
to all
channels
100 MHz

## Slide 11: Effect of local calibration
before
after

## Slide 12: Distribution of Dti
nominal value = 0.2 ns

## Slide 13: How to quantify accuracy of calibration
sine with random phase
10 ns
cells

## Slide 14: Global Calibration
Dta,b = Sti = 10 ns
cells
!
Measure period of sine wave (linearly interpolate zero crossings)
Calculate correct factor
Distribute correction equally to cells between zero crossings
Iterate with random phase

## Slide 15: Effect of global calibration

## Slide 16: Measure resolution with real pulses
Pulse
Generator
DRS4
Evaluation
Board
CH1
CH2

## Slide 17: Result of time difference @ 5 GSPS
2.5 ps
Rise time: 500 mV / 2 ns = 100 mV / 0.36 ns
Delay = 11.6 ns
Value in RMS
Linear interpolation
Single measurement 2.5ps/√2 = 1.8 ps

## Slide 18: Calibration of all channels
Calibrating one channel of DRS4 chip → 12 ps
Calibrating each channel individually → 2.5 ps
Large system needs clock distribution with ~ps accuracy

## Slide 19: Advanced Analysis
Accuracy can be increased by taking more points into account (unlike TDC)
threshold
line distance
Simple Threshold
Line Fit
Cross-correlation
with interpolation
Win √n !

## Slide 20: Results in dependence of delay
Dt [ps RMS]
Improvement of ~40 x to old calibration
No degradation for delay > 0
10x better than Acquiris ADC 2 GSPS/10 bit
Cross correlation gives best result
Result < 1ps (single measurement) for delay < 30 ns
Dt [ps RMS]
1.42 ps

## Slide 21: Temperature Dependence
Calibration seems stable on the 1-2 ps level over wide temperature range

## Slide 22: Limitation of time measurement
SNR of detector signal
Variation of Vth → calibration
Noise inside inverter chain
DLL phase noise
100 ps
2.5 V
Vth
→ Dt=0.03 ps

## Slide 23: DLL Phase Noise
REFCLK
CLK
PLLOUT

## Slide 24: Application of precise timing
Readout of straw tubes / drift chambers
Position along the wire with charge division
Time measurement with cluster technique should give 3 ps → 0.5 mm
HV

## Slide 25: Conclusions
New calibration for DRS4 chip give sub-ps resolution for delay < 30 ns
Patent PCT/EP2013/070892 <Bernd.Pichler@med.uni-tuebingen.de>
SNR of detector signal limits timing
To improve timing even further:
  work hard on SNR
  carefully calibrate voltage response of EACH cell
  reduce DLL phase noise

## Slide 26
FEMTO
2017