## Slide 1: Outstanding Issues of HTS for High Quality Magnets
H. Maeda and Y. Yanagisawa
RIKEN, Yokohama, Japan
May 21-23, 2014
1st workshop on Accelerator magnets in HTS at DESY
1
NMR Facility

## Slide 2
Introduction
Magnetic field distortion and temporal magnetic field drift due to the screening current
Effect of screening current on the characteristic of LTS/Bi2223 NMR and LTS/REBCO NMR
CONTENTS

## Slide 3: Introduction

## Slide 4
4
HTS conductors
HTS
Magnet systems
Technological Gap
Gap between HTS conductors and HTS magnets
NMR
Accelerator
REBCO
It is necessary to develop HTS magnet technology
Notes: 4

## Slide 5
LTS magnets
HTS magnets
Magnet technology
Stability
AC loss
Quench protection
Structure
Magnetic field
Cryogenic engineering
Degradation in the coil performance due to delamination(REBCO)
Screening current-induced magnetic field (HTS)
Protection against thermal runaway(HTS)
Structure
Magnetic field
Cryogenic engineering

## Slide 6: Magnetic field distortion and temporal magnetic field drift due to the screening current

## Slide 7
Basic phenomena of the screening current-induced magnetic field

## Slide 8: Screening current-induced magnetic field for LTS magnets
LTS Coil
LTS Filaments
Screening current-induced magnetic field, Bs
Coil current magnetic field, Bc
For early LTS magnets, the screening current was enormous.
However, for modern LTS magnets, effect of screening current becomes negligible as LTS filaments are fine.
Net field
B0 = Bc + Bs
Filament
Bz
LTS mf conductor
Axial magnetic field
Screening current

## Slide 9
REBCO
coil
Screening current
Screening current-induced
field, -Bs
Radial magnetic
field
REBCO
layer
Coil current magnetic field, Bc
Net field
B0 = Bc - Bs
Screening current-induced magnetic field for HTS magnet
Br
Notes: In this slide, I want to explain the mechanism of the screening field. This is a schematic of a YBCO circular coil. During charging of the coil, radial component of the magnetic field is applied to the YBCO conductor surface and this field induces the screening current in the YBCO layer as excluding magnetic flux from superconductor. The screening current generates a screening field in the coil center in the opposite direction to the coil field. Therefore, the apparent field is composed of the coil field and screening field. Furthermore, the screening current decreases with time resulting in the long term field drift. So, the major problems for YBCO coil are reduction in the central field and field drift with time. In the following slides, I will show experimental data of screening field.

## Slide 10
Major problems:
Hysteresis effect
Reduction in the central field
Field drift with time
ID 18mm
OD 45mm
L 25mm
Solenoid
Y. Yanagisawa et al., IEEE Trans. Appl. Supercond. 20, 744-747(2013)
@4.2K
Hysteresis effect of the screening current-induced magnetic field for a REBCO magnet
Notes: Here you can see the central magnetic field for a small YBCO solenoid as a function of coil current. During charging and discharging process the values of the field dose not coincide. This nonlinear effect is due to screening field. You can extract the screening field from the measured data and this is the screening field as a function of coil current. For the virgin ramp the screening field decreases with coil current. Then, the field saturates, and starts to decrease in its absolute value. After that the field shows the general type of hysteresis for magnetization. The important point here is that the screening field reduces the central field. For example, when the coil current is 100 A, the central field is 1.1 T while the screening field is minus 0.17T. So the field reduction rate is as large as 15%, which is very large rate. Dashed line in both graph shows a numerical simulation result. The detail of the numerical method is described in our previous paper. So I will skip the explanation of the method and move on to the simulation results for various coils.

## Slide 11
Bi2223 coil
REBCO coil
Comparison between Bi2223 and REBCO
-29mT
Saturation
Screening current-induced magnet field for a REBCO coil is 5-fold larger than that for a Bi2223 coil
Id 81mm
Od 119mm
L 400mm
Unsaturation
@4.2K
-155mT
Id 81mm
Od 121mm
L 375mm

## Slide 12: Dependence of hysteresis on the coil shape
: Shape factor of the solenoid
@4.2K
Thin cylinder
Disk
Min. volume design
Numerical simulation for REBCO
The central magnetic field is the same, while the shape is different
Notes: Here we have hysteresis loops of screening field for the coils. Vertical axis shows screening field while horizontal axis shows coil current normalized by the critical current of the coil. This is a long and thin coil while this is a disk shape coil and this is minimum volume coil. What is interesting here is that the minimum volume designed coil shows the most apparent screening field. Let me compare the field reduction rates for these coils. For the long and thin coil, the coil field is 1.58T while the screening field is -0.0363T at this point. The field reduction rate is as small as 2.3%. For the disk shape coil, the rate is also as small as 3.1%. But, for the minimum volume coil, the field reduction rate is as large as 16%. The most important point here is that the field reduction rate shows maximum when the coil is minimum volume configuration. In other words, the minimum volume coil is effective for generating not only central field but also screening field. This rate is large problem for the practical superconducting coil. However, in the present case, the inner diameter of the coils are as small as 18mm. So we have to investigate the screening field for much larger coils.

## Slide 13
Remedies for the screening current-induced magnetic field

## Slide 14: Current sweep reversal method
Y. Yanagisawa et al., Physica C 469, 1996-1999(2009)
z
@4.2K
z
Flux
Flux
@77K
Several % of current sweep reversal is necessary for HTS magnets, while 0.2-0.5% for LTS magnets.
The magnet should endure a current of 105% of the operation current
REBCO coil

## Slide 15: Temperature change
Y. Yanagisawa et al., Adv. Cryo. Eng. 1434, 1373-1380 (2012)
Lowered
If the coil temperature is lowered, the critical current density is enhanced, thus flux creep is suppressed
REBCO coil

## Slide 16: Scribed multi-filamentary REBCO conductor
Mechanically scribed 3-filasment REBCO conductor
@77K
Usual REBCO conductor
The screening current-induced field is remarkably reduced, if a scribed REBCO conductor is used
4mm
4mm

## Slide 17: Bi2212 multi-filamentary round conductor
If multi-filamentary/ twisted Bi2212 round conductor is used, the screening current-induced magnetic field will be negligible as for LTS coils.
U. P. Trociewitz, et al. Appl. Phys. Lett. 99, 202506(2011)
D. C. Larbalestier et al., “ A transformative superconducting magnet technology, 2013.
33.8T coil, developed in NHMFL
Bi2212 round conductor

## Slide 18: Effect of screening current on the characteristic of LTS/Bi2223 NMR and LTS/REBCO NMR

## Slide 19
Necessity of the LTS/ HTS NMR magnet

## Slide 20
History of the NMR magnetic field
High Temperature Superconducting (HTS) Magnets.
Hc2>50 T
Low Temperature Superconducting (LTS) Magnets.
Hc2< 25 T
1 GHz, 23.5 T
1 GHz NMR
920 MHz NMR
RIKEN started to develop super-high field NMRs
Notes: 20

## Slide 21
4 mm
Superconducting multi-filaments
Two available HTS conductors
Silver matrix
0.3-0.5 mm
4-12 mm
Hastelloy
Superconducting thin layer(1-2 μm）
Copper
Buffer layer
0.1-0.15 mm
Bi2223 conductors
REBCO conductors
A: Long conductor 1 km
D: Low mechanical strength Lower current density  larger magnet
A: High mechanical strength
High current density
 compact size magnet
D: Shorter length 100-200m

## Slide 22: Technical problems for an NMR magnet
Strong magnetic field intensity
Spatial field homogeneity
1ppb／5mm dia. cylinder
Temporal field stability
0.1ppb
NMR sample
NMR magnet
22

## Slide 23
Axial error component
Radial error component
Spatial homogeneity of the magnetic field
Superconducting shim coils(z,z2,x,y, zx, zy, xy,x2-y2) < 10 ppm
Field correction coils
（z2 and z4 are canceled）
< 500ppm
Room temperature shim coils & iron shims <1ppb
Legendre Polynomials
Target: 1ppb

## Slide 24
Comparison between LTS/ Bi2223 NMR magnet and LTS/REBCO NMR magnet

## Slide 25
REBCO inner coil
NMR probe
(1H, 13C, 15N )
World’s first LTS/Bi2223 NMR magnet and LTS/REBCO NMR magnet(400-500MHz)
Ultra-stabilized
DC-power supply
< 1ppm
Bi2223 inner coil
)
Yanagisawa et al, JMR (2010)
Kiyoshi et al, IEEE TAS (2010)
Pulse-tube cryocooler
NbTi/ Nb3Sn outer coils
(2009)
(2013)
Id 81mm
Od 121mm
L 375mm
Notes: We installed the REBCO coil in the conventional LTS NMR at the inner most part. The configuration of the REBCO coil is nearly equal to the Bi2223 innermost coil developed and tested in 2008. The magnet was operated in driven mode with a highly stabilized DC power supply. We have a 4K pulse tube cryocooler on the top flange of the magnet. The evaporated liquid helium due to heat leak was recondensed using pulse-tube cryocooler for helium refill free continuous operation.

## Slide 26
Temporal field drift after the magnet charge
Temporal field drift due to screening current for an LTS/REBCO NMR is >20-times steeper than that for an LTS/Bi2223 NMR
Notes: I am going to show the magnetic field drift with time after magnet charging due to the relaxation of the screening current. For Bi2223 NMR, the drift is not so large and it took 20 days to obtain 0.01ppm/h which is the required stability of NMR magnets. So we just waited for 20 days. On the other hand, for the REBCO NMR, the drift was substantial. It is estimated that we require 600 days to get the rate of 0.01ppm/h. This is impractical time interval. We used current sweep reversal operation with about 10% overshooting to suppress the drift and obtained the stable magnetic field.

## Slide 27
Spatial distribution of the magnetic field
Z1: -23ppm at ±1cm
Z2: -83ppm at ±1cm
Z1: -2.1ppm at ±1cm
Z2: -0.41ppm at ±1cm
LTS/Bi2223 NMR
(500MHz, 11.7T)
LTS/REBCO NMR
(500MHz, 11.7T)
The field generated by the field correction coil is reduced by 27% due to the LTS/REBCO coil, resulting in excessive residual z2 component. Therefore, additional iron shims were required.
Axial position (cm)
>> superconducting shim coils(2ppm)
<< superconducting shim coils(2ppm)
Notes: Next, I am going to talk about the magnetic field distribution. This is magnetic field versus coil axial position for Bi2223 NMR. There is not notable field inhomogeneity and we corrected the field error harmonics using SC shim and RT shim. On the other hand, for REBCO NMR, we found the remarkable inhomogeneity. Here you can see the notable Z2 error harmonics. This is

## Slide 28
Degradation in the superconducting(SC) shim coil performance
Axial SC shim coils
Radial SC shim coils
LTS/Bi2223
40% of the design
LTS/REBCO
 20% of the design
LTS/Bi2223
20-40% of the design
LTS/REBCO
 < 5% of the
design
The SC shim coil performance is remarkably degraded due to the screening current induced in the HTS coil, especially for REBCO
Difficult to make a field correction!
Notes: Radial shims coils generate transverse magnetic field like this. So the shim coil field is effectively shielded by the screening current. That’s why we could not correct the residual radial field error harmonics.

## Slide 29
 | LTS NMR | LTS/Bi2223 NMR | LTS/REBCO NMR
NMR resolution | <1Hz(2ppb) | 0.7Hz(1.4ppb) | 15Hz(30ppb)
NMR sensitivity S/N | >600 | 512 | 28
Comparison of NMR resolution and sensitivity
The same quality
20-fold lower resolution and sensitivity
Resolution
Noise
Signal
Sensitivity(S/N)
NMR spectrum

## Slide 30
NMR spectra for a protein
HSQC
LTS/Bi2223 NMR magnet
LTS/REBCO NMR magnet
Nearly the same quality as conventional NMRs
1H (ppm)
15N (ppm)
The number of the available NMR measurement is limited
2D-HNCO
Ubiquitin (protein)
30

## Slide 31
Collaboration between NHMFL and RIKEN
on a 400 MHz LTS/Bi2212 NMR magnet
Develop and evaluate an LTS/Bi2212 NMR magnet in collaboration between NHMFL, coil fabrication, and RIKEN, NMR measurement.
The experiments in RIKEN is planned to be conducted in 2015.
Bi2212 inner coil
NMR probe
RIKEN
NHMFL

## Slide 32
 | LTS NMR | LTS/Bi2223 NMR | LTS/Bi2212 NMR | LTS/REBCO NMR
NMR resolution | <1Hz(2ppb) | 0.7Hz(1.4ppb) | XXX | 15Hz(30ppb)
NMR sensitivity S/N | >600 | 512 | XXX | 28
Comparison of NMR resolution and sensitivity
Resolution
Noise
Signal
Sensitivity(S/N)
NMR spectrum

## Slide 33
Summary
Effect of screening current induced magnetic field on the performance of high quality magnet such as NMR, MRI and accelerator is enormous.
It can be effectively suppressed by current sweep cycling and temperature change cycle, in addition to using adequate conductor.
The quality of the NMR spectra achieved by LTS/Bi2223 NMR magnet is nearly the same as that by LTS NMR magnet.
NMR spectra achieved by the LTS/REBCO NMR magnet is 20-fold broader than those by LTS NMR. Therefore, available NMR measurements are rather limited.

## Slide 34
Acknowledgement
This work is partially supported by JST
RIKEN
Dr. T. Yamazaki
Dr. M. Takahashi
Dr. X. Jin
JASTEC
Dr. M. Hamada
Dr. K. Kominato
NIMS
Late Dr. T. Kiyoshi
Dr. S. Matsumoto
JEOL RESONANCE
Mr. H. Suematsu
Chiba University
Prof. Hideki Nakagome
Sophia University
Prof. T. Takao
Mr. S. Iguchi

## Slide 35
22
35
Thank you very much for your attention!

## Slide 37
Why higher magnetic field is required in NMR ?
0.94T (40MHz)
11.75T(500MHz)
Protein sample
Frequency
Frequency
NMR resolution and sensitivity (S/N) are remarkably enhanced with the increase of the magnetic field
NMR magnet
NMR sample

## Slide 38: Numerical simulation of the screening current induced magnetic field
■Current diffusion equations for each current element
■Relation between transport current and current density
■Power law E-J characteristics
Convergence
■Jc-B characteristics
Current element
Notes: Here, we have numerical simulation method for calculating the screening field. The detail of the method is described in our previous paper, so I will skip the detail and show simulation result in the next slide.

## Slide 39
1200
600MHz
(14.1T)
700 MHz
(16.5T)
800 MHz
(18.8T)
930MHz
(21.9T)
Bottleneck of high-field NMR magnets
1.2GHz
(21.8T)
LTS/REBCO NMR magnet
LTS/Bi2223 NMR magnet
LTS NMR magnet
600
700
800
930
Notes: By using this kind of technology, we want to make this kind of paradigm shift for NMR magnets. We are thinking about the 1.1GHz NMR magnet with the size of 700MHz magnet.

## Slide 40
Ultra-stabilized DC current supply <1ppm
High temporal stability of the local magnetic field is achieved by using a 2H field-frequency lock system <1ppb
Current
Temporal stabilization of the magnetic field
Target: 0.1ppb