## Slide 1: Fully depleted MAPS: Pegasus and MIMOSA 33
Maciej Kachel, Wojciech Duliński
PICSEL group, IPHC Strasbourg
For low energy X-ray applications

## Slide 2: Motivation
Goal (success oriented):
  Low energy X-rays (< 10 keV)
  Single photon counting (with > 106 ph / 100x100 µm2)
  Good spatial resolution (few µm)
maciej.kachel@iphc.cnrs.fr
Chip requirements:
          Depletion depth ~ 50-100 µm
          Low noise (sensitivity to few 100 e- signal)
          Fast frame rate >> kHz
          Pixel pitch ~ 20 µm
Intermediate step ASIC:
  Counting > 103 ph/pixel·s
  Energy range 1 keV – 10 keV
  Noise below 25 e-
  Pixel size 25 x 25 µm2
  Chip size ~1 cm2
How to get depletion??

## Slide 3: MAPS
maciej.kachel@iphc.cnrs.fr
Notes: Thin sensitive area – only few um of depletion… - limited in Xray appications, but work with indirect detection

## Slide 4: Detection with MAPS
Standard low resistivity epi (few Ω cm)
High resistivity epi (k Ω cm)
With the same electronics you gain the depletion zone depth
But this is not enough!!
Diode
Provided by A. Dorokhov
maciej.kachel@iphc.cnrs.fr

## Slide 5: Absorption efficiency of Si
Chart: Photon energy [keV]
Si - 300 µm
Chart: Photon energy [keV]
Si - 10 µm
Chart: Photon energy [keV]
Si - 50 µm
maciej.kachel@iphc.cnrs.fr

## Slide 6: Technology choice
Test structures (collecting diodes and amplifiers) were made in both technologies
TOWER (0.18 um)
ESPROS (0.15 um)
Quadruple well technology
P substrate
Used by CERN
Fully depleted 50 μm
Post processing included
N substrate (detector grade)
„Out of the box solution” ?
maciej.kachel@iphc.cnrs.fr

## Slide 7: Mimosa 33 submission ESPROS technology
Study different biasing approches
  - Bias the diode from bottom
  - Bias from the top – DC coupling
maciej.kachel@iphc.cnrs.fr
Study the charge collection with different pixel pitch
Prototype of a simple X-ray counter

## Slide 8: M33: Standard SF pixel
Isolated NMOS transistors (but used with std voltage)
Diode biased from the anode side with negative voltage
SF bias - self biased solution
maciej.kachel@iphc.cnrs.fr

## Slide 9: M33: Standard SF mesurements
Increasing the Vphanode improves the charge collection
High noise ~ 50 e- rms ?
Low gain because of the large capacitance at the input node ??
Fe55 measurements with different biasing(Vback, Vphanode)
where the extracted capacitance at the gate of input transistor ~ 1-2fF
maciej.kachel@iphc.cnrs.fr
Seed pixel

## Slide 10: M33: Standard SF mesurements (I) depletion depth
Mimosa 33 was tested with Sr beta source
Chart: DAC units
Apart from that – good charge collection -> most of the charge is collected by the single pixel
It looks like there is only 15µm of depleted detector (out of 50)
Beta particle leave ~80 e- / µm of detector
1200 e- -> 15um of depleted substrate ??
maciej.kachel@iphc.cnrs.fr
Seed pixel

## Slide 11: M33: bias the diode from the top
Chart: ADC units
First stage of the pixel can be supplied from higher voltages (i.e. 3.2 V – 5.0 V)
~4V at the diode
With vback=-3V (or even 0V) response from beta particles is similar to the previous case
maciej.kachel@iphc.cnrs.fr
Seed pixel

## Slide 12: M33: X-ray counter
SR90 beta source
Fe55 X-ray source
SR90 beta source
Fe55 X-ray source
Chip is responsive
Due to the low gain problem the chip was not tested extensively
Developed as a proof of concept
16 x 18 pixels – 50 um pitch
In pixel:
CSA, shaper, comparator , 4bit counter and 4bit DC correction
maciej.kachel@iphc.cnrs.fr

## Slide 13: M33: problem & second submission
The reason for the low gain/high noise was the wrong sensing element placed by the foundry
  Corrected version of chip was resubmitted and is expected in November 2014.
Foreseen tests
  Study the charge collection with different pixel pitch (25 um, 50um)
  Custom made sensing elements inserted
  Pixel designs modified, with AC coupled diode
maciej.kachel@iphc.cnrs.fr

## Slide 14: Use standard CMOS technology
Courtesy of T. Hemperek
Bonn University
maciej.kachel@iphc.cnrs.fr

## Slide 15: Pegasus
180nm CMOS (Tower)
Various types of epi in submission
25 x 25µm2 pixels
Matrix - 56x32 with 4 versions of pixels
  Two source followers
  Two amplifiers
Amplifier
Input diode AC coupled with amplifier
Based on inverter
No bias needed
maciej.kachel@iphc.cnrs.fr

## Slide 16: Pegasus – measurements results
Fe55 spectrum with different diode voltage (seed pixel)
ENC ~25 e-
Tests performed on the chips with a 18 µm thick epitaxial layer
maciej.kachel@iphc.cnrs.fr

## Slide 17: Pegasus depletion depth estimation
Pegasus chip (in-pixel amplifier)
Chart: DAC units
1400 e- -> substrate with 18μm epi
maciej.kachel@iphc.cnrs.fr
Seed pixel

## Slide 18: Pegasus 2
Similar to Pegasus 1 (also 4 versions of pixels)
Problems with high frame rate in first submission
pixels optimized
  Rolling shutter readout -> 150ns per row
Tests just started
maciej.kachel@iphc.cnrs.fr

## Slide 19: Pegasus 2 Cluster shape
maciej.kachel@iphc.cnrs.fr

## Slide 20: Pegasus 2 preliminary measurements
Source follower:
Gain 17 uV/e- => Cin = 9.38 fF
ENC 17 e- (base)
28 e- (Fe peak)
Amplifier:
Gain 70 uV/e-
ENC 16 e- (base)
30e- (Fe peak)
maciej.kachel@iphc.cnrs.fr
Seed pixel
Seed pixel

## Slide 21: Conclusions
With those „new” technologies MAPS looks promising for wide variety of low energy X-ray applications
Both approaches are promissing
  Charge collection – TOWER – full depletion of epi - most of the events in 1-2 pixels
  Small noise – 100 eV resolution at 6 keV
  Frame rate – to achieve the high frame rate goal -> 3D approach like professionals do
Study of the charge collection and depletion capabilities will be performed
  Charge collection with different pixel pitch (4µ, 25µ, 50µ)
  Response linearity with energy
  Sensitivity to 1 keV photons
  Depletion uniformity– Sr source, Edge TCT studies
  Neutron irradiated samples
maciej.kachel@iphc.cnrs.fr

## Slide 22
Thank you for your attention

## Slide 24: Motivation
Advantages :
Smaller pixels
Standard CMOS technology ($)
No bump bonding needed
Noise lower than in hybrid det. Due to small capacitance at the input
Disadvantages:
Low energy X-rays only ( < 20 keV )
Radiation tolerance smaller than in hybrids
Monolithic sensor
Hybrids
Advantages :
Different materials of detectors
Thick detectors for higher energies available
Radiation hardness
Disadvantages:
Pixel pitch
Hard to set the threshold low
Costs: Detectors, Bonding
Maciej Kachel

## Slide 25: MAPS
Standard MAPS
Introduced at the end of the XXth century for digital cameras
Only NMOS transistors in pixels (PMOS would take part in charge collection)
Charge collection mainly through diffusion (low efficiency, no depletion)
Thin sensitive area – limited in Xray applications, but work with indirect detection
Used in various charge particle detection applications
Maciej Kachel
Notes: Thin sensitive area – only few um of depletion… - limited in Xray appications, but work with indirect detection