## Slide 1: Beam instrumentation for machine protection
Evian workshop
2-4 June 2014
E. Bravin – CERN
For the BE-BI group

## Slide 2: Outline
Interlocked BPMs
  Changes
  Status at startup
  Scrubbing with doublets
FBCCM (dI/dt)
  Changes
  Status at startup
Abort gap monitor
  Changes
  Status at startup

## Slide 3: Interlocked BPMs
Strip line pick-ups installed in IR6 just after Q4 (BPMSX was BPMSA) and just before the TCDQ (BPMSI was BPMSB)
Prevent beam on TCDS
Acquisition is based on the LHC BPM design with dedicated firmware
Two operational ranges used (high and low sensitivity modes)

## Slide 4: Issues in run 1
Interlock fires for bunches near the detection threshold
Reflections from intense bunches cause false triggers (seen as weak bunches)
Detection threshold adjusted (by changing attenuators) between signal and reflections
Insufficient diagnostics for the PM analysis

## Slide 5: Actions during LS1
Replaced shorted strip-lines with 50 Ω terminated s.l.
Added absorptive filter at strip-line output (100 MHz)
Separated orbit and interlock functions (2 DABs)
Improved HW & firmware: bunch-by-bunch interlock post-mortem data (3564 slots x 294 turns)
Installed thermal controlled racks (stability)
FESA adapted to new firmware (and new CPUs)
New GUI for the interlock post-mortem t.b.d. (BI/OP)

## Slide 6: Old BPM setup

## Slide 7: New BPM setup
The BPMs have been renamed
(LHC-BP-EC-0002)
The Interlock and orbit systems have been separated

## Slide 8: BPMs Reflections
LPF BW
> 20dB

## Slide 9: Reflections in time domain
Shorted strip-lines reflections
Measurement: -27 dB
Simulation: -34 dB
Terminated strip-lines with LPF:
Simulations: <-46 dB

## Slide 10: Interlock thresholds
Chart: BPMSA.A4L6.V1
| 3.79099839844087E10 | 3.37873087907395E10 | 3.01129706567611E10 | 2.68382133478323E10 | 2.39195828240892E10 | 2.13183506317373E10 | 1.9E10 | 1.69337678245412E10 | 1.50922364597614E10 | 1.34509699032986E10 | 1.19881895451237E10 | 1.06844851786166E10 | 9.52255743891817E9 | 8.4869882508683E9 | 7.56403624051645E9 | 6.74145439543793E9 | 6.00832755431992E9 | 5.35492756940246E9 | 4.7725842198682E9 | 4.25357016327985E9 | 3.79099839844087E9 | 3.37873087907395E9 | 3.01129706567612E9 | 2.68382133478323E9 | 2.39195828240892E9 | 2.13183506317373E9 | 1.9E9 | 1.69337678245412E9 | 1.50922364597614E9 | 1.34509699032986E9 | 1.19881895451237E9 | 1.06844851786166E9 | 9.52255743891817E8 | 8.4869882508683E8 | 7.56403624051644E8 | 6.74145439543793E8
Lin H sens | 0.0019443 | 0.0012962 | 0.0019443 | 0.0012962 | 0.0019443 | 0.0019443 | 0.0019443 | 0.0019443 | 0.00064809 | 0 | 0.0012962 | 0 | -0.00064809 | 0 | 0.00064809 | 0.0025924 | 0.0019443 | -0.0012962 | -0.00064809 | 0 | -0.00064809 | -0.0045366 | -0.0012962 | 0.00064809 | -0.0019443 | 0.0038885 | -0.0032404 | -0.0019443 | 0.0090732 | 0.012962 | 0.012314 | -0.011666 | 0.0012962 | 0.0084251 | -0.018146 | 6.4809e+46
Lin L sens | 0 | -0.0012962 | 0.0025924 | 0.0025924 | 0.0019443 | 0.0019443 | 0.0019443 | 0.0019443 | 0.0038885 | 0.0038885 | 0.0019443 | 0.0038885 | 0.0025924 | 0.0025924 | 0.0019443 | 0.0019443 | 0.00064809 | 0.0012962 | 0.00064809 | 0.0012962 | 0.0019443 | 0.0045366 | 0.0032404 | -0.0012962 | 0.0012962 | 0.0025924 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46 | 6.4809e+46
Pilot | -2 | 2
Nominal | -2 | 2
Ultimate | -2 | 2
Nominal
Ultimate
Noise floor
Thresholds to be defined
Verifications with beam required

## Slide 11: Scrubbing doublets
Chart: Effect of the bunch spacing at the BPMSB (interlock)
| 25.0 | 20.0 | 15.0 | 10.0 | 7.5 | 6.0 | 5.0 | 4.0 | 3.0 | 2.85
"Low sensitivity" | 0 | -0.016142 | -0.026902 | -0.050752 | 0.1377 | 0.086539 | -0.50921 | -0.40894 | 0.01142 | 0.0071748
"High sensitivity" | 0 | 0.0078303 | 0.031263 | -0.062746 | 0.055687 | -0.17201 | -0.93458 | 0.21198 | 0.16158 | 0.060992
Beam “simulator” tests (beam signal replaced by pulse generator)
May be possible to test on SPS with beam
Chart: Effect of the bunch spacing at the arc BPM
| 25.0 | 20.0 | 15.0 | 10.0 | 7.5 | 6.0 | 5.0 | 4.0 | 3.0 | 2.85
"Low sensitivity" | 0 | -0.0063976 | -0.010662 | -0.020115 | 0.054574 | 0.034298 | -0.20182 | -0.16208 | 0.0045262 | 0.0028436
"High sensitivity" | 0 | 0.0031034 | 0.01239 | -0.024869 | 0.022071 | -0.068174 | -0.37041 | 0.084014 | 0.064042 | 0.024173

## Slide 12: FBCCM (dI/dt)
Interlock system based on the reading of the fast beam current transformers
Detects rapid changes of beam currents i.e. losses or debunching (of one or more bunches)
Six different time integration windows
  1, 4, 16, 64, 256 and 1024 turns
32 energy thresholds
Prototypes tested in 2012

## Slide 13: FBCCM Signal Processing
LP Filter
Band Pass Filter
40 MHz (FIR)
IQ demodulation
Moving average
1
ADC
Threshold
Comparator
Energy from telegram
To CIBU
Loss Calculation
From FBCT
Moving average
4
Threshold
Comparator
Loss Calculation
Moving average
1024
Threshold
Comparator
Loss Calculation
.
.
.
.
.
.
.
.
.
(16, 64, 256)

## Slide 14: FBCCM units
6 Units produced based on prototype
  Reduced noise
  One beam per box
4 Will be installed in July (2 oper. B1/B2 + 2 dev. B1/B2)

## Slide 15: FBCCM for 2015
Two systems per beam running in parallel
  One will be operational (System A) (stable HW and FW)
  One will be used for debugging/development (system B)
New fast current transformers (to address position sensitivity)
  System A remains old FBCT (position and σz sensitivity)
  System B will be CERN/BERGOZ BCTI on Beam 1 and CERN BCTW on Beam 2
  Switch to BCTIs or BCTWs on both beams (without breaking vacuum) later on depending on results with beam
FESA class 90% ready, expert GUI almost ready, Post mortem analysis tool to be defined (BI/OP)
Connected to the BIS but initially masked (allowing to collect statistics)

## Slide 16: FBCCM commissioning
Need dedicated beam time for repeating the lab measurement (controlled losses, scraping)
Lot of learning, debugging and setting-up can be done in parallel with the normal operation of the machine
Possibility to carry out realistic beam simulations in the lab?

## Slide 17: Abort Gap Monitor
Monitors the particles population inside the 3μs abort gap (needed by the dump kickers)
It is based on the detection of synchrotron light by a gated photomultiplier and is integrated in the BSRT telescope

## Slide 18: Issues during run 1
Interference between BSRT and AGM
Extraction mirror heating
Software issue affecting reliability
Lots of manual actions for calibration and verification (expert only)

## Slide 19: Changes during LS1
New extraction mirror [LHC-BSRTM-EC-0002]
  No more heating problem (TBC with beam)
New layout of optical table (optimized)
New readout electronic (not deployed at start up)
New automated software procedures for calibration and diagnostics [EDMS 1337184]
  Actions to be included in the sequencer

## Slide 20: AGM and SIS
AGM will eventually be connected to SIS
SIS also used for triggering AG cleaning
AGM-FESA to handle cleaning and dump logic
Thresholds need to be finalized (based on new quench limits)

## Slide 21: Conclusions
The impact of the interlocked BPMs on the length of fills should be minimized (surely for ions)
  Thresholds levels to be defined
4 FBCCM will be available at startup
  Prototype and lab test give encouraging results, some debugging and fine tuning required when beam comes back
Reliability of AGM improved
  No significant changes in performances (better if any)
  System should become much more reliable and self-diagnosing

## Slide 22: Thanks!
The End

## Slide 23
Support slides

## Slide 24: Dump Channel
The main aim of these BPMs is to avoid large orbit offsets leading to high losses on the septum protection during a dump
BPMSX (.A/.B)
BPMSI (.A/.B)

## Slide 25: Interlock Mechanism
Two separate trigger logics
  70 reading in the last 100 turns out of limits (a single bunch can trigger the dump)
  250 reading in the last 10 turns out of limits (response to fast orbit changes)
Limits set to ~±3mm
The whole chain from readout to beam dump trigger is in hardware (and firmware)
Interlock signal connected to a maskable channel of the BIS

## Slide 26: BPMs Signals (Protons)
Chart: Intensity dependence of WBTN cards
| 5.0E9 | 5.0E9
Pilot | -2 | 2
Nominal | -2 | 2
Ultimate | -2 | 2
Sens Switching | -2 | 2
Pilot
Nominal
Ultimate
Sensitivity
switch
Reflections
-27 dB
Hi sensitivity
Low sensitivity
Note: 1% ~220um for BPMSA and ~330um for BPMSB
Moved limits beginning of 2012 to allow longer physics fills
(reduced attenuators)
Useful range
New useful range

## Slide 27: BPMs Signals (Ions)
Chart: Intensity dependence of WBTN cards
| 5.0E9 | 5.0E9
Pilot | -2 | 2
Nominal | -2 | 2
Ultimate | -2 | 2
Sens Switching | -2 | 2
Pilot
Reflections
-27 dB
Hi sensitivity
Low sensitivity
Nominal
With p-Pb only use Hi Sensitivity mode (only 1.5 1010 charges / bunch)
P-Pb MD showed that the satellites of the nominal beam triggered the interlock
Moved limits beginning of 2013 to cover ~1.5E10 bunches
New end of fill

## Slide 28: Doublets simulations 1
Chart: Normalized Position (half aperture=1)
| -0.1 | -0.08 | -0.06 | -0.04 | -0.02 | 0.0 | 0.02 | 0.04 | 0.06 | 0.08 | 0.1
0.5 * B1=B2, 2.5ns | -0.045223 | -0.032478 | -0.026624 | -0.020406 | -0.013349 | 3.0947e-05 | 0.0085697 | 0.022262 | 0.027902 | 0.033792 | 0.040324
0.5 * B1=B2, 5ns | -0.085668 | -0.061436 | -0.043516 | -0.030956 | -0.015679 | 0.0055394 | 0.026413 | 0.050782 | 0.070807 | 0.08295 | 0.093873
B1=B2, 2.5ns | -0.077555 | -0.058677 | -0.046121 | -0.033034 | -0.01972 | 0.00026147 | 0.015646 | 0.035646 | 0.047853 | 0.060171 | 0.073214
B1=B2, 5ns | -0.21776 | -0.1722 | -0.12911 | -0.08225 | -0.03761 | 0.01264 | 0.057243 | 0.10917 | 0.14554 | 0.19819 | 0.24286
1.1 * B1=B2, 2.5ns | -0.075154 | -0.056254 | -0.044312 | -0.031816 | -0.01917 | 0.00011429 | 0.014789 | 0.034641 | 0.046332 | 0.058725 | 0.070664
1.1 * B1=B2, 5ns | -0.25094 | -0.19784 | -0.15001 | -0.095167 | -0.04413 | 0.010919 | 0.064859 | 0.12486 | 0.16924 | 0.21963 | 0.26634
1.2 * B1=B2, 2.5ns | -0.07186 | -0.054012 | -0.042238 | -0.030713 | -0.018689 | 0.00023132 | 0.014201 | 0.03303 | 0.044616 | 0.055721 | 0.067479
1.2 * B1=B2, 5ns | -0.28437 | -0.21717 | -0.1647 | -0.1118 | -0.052346 | 0.020378 | 0.077542 | 0.13996 | 0.18724 | 0.24961 | 0.30534
1.5 * B1=B2, 2.5ns | -0.062254 | -0.046596 | -0.037107 | -0.026558 | -0.015426 | 0.0015893 | 0.013242 | 0.03 | 0.039555 | 0.048349 | 0.059746
1.5 * B1=B2, 5ns | -0.44035 | -0.3473 | -0.2596 | -0.16799 | -0.067126 | 0.0036304 | 0.105 | 0.20335 | 0.28957 | 0.37204 | 0.46068
Bunch 1 and bunch 2 with same position
Y scale for arc BPM
Multiply time 2.5 for BPMS

## Slide 29: Doublets simulations 2
Chart: Normalized position (half aperture=1)
| -0.1 | -0.08 | -0.06 | -0.04 | -0.02 | 0.0 | 0.02 | 0.04 | 0.06 | 0.08 | 0.1
0.5 * B1=B2, 2.5ns | -0.20152 | -0.15979 | -0.11976 | -0.080423 | -0.041599 | 0.0003401 | 0.03979 | 0.081616 | 0.11957 | 0.15939 | 0.19875
0.5 * B1=B2, 5ns | -0.82979 | -0.66209 | -0.49621 | -0.33056 | -0.16251 | 0.0079171 | 0.17686 | 0.34889 | 0.51015 | 0.66952 | 0.84331
B1=B2, 2.5ns | 0.46164 | 0.37166 | 0.27812 | 0.18424 | 0.090678 | 0.00026979 | -0.092502 | -0.18277 | -0.27695 | -0.37009 | -0.46352
B1=B2, 5ns | -1.0328 | -0.82145 | -0.61427 | -0.40441 | -0.19704 | 0.01264 | 0.21846 | 0.42869 | 0.63154 | 0.84833 | 1.0521
1.1 * B1 B2, 2.5ns | 0.51688 | 0.41501 | 0.31082 | 0.20622 | 0.10157 | 0.00022794 | -0.10383 | -0.20514 | -0.31038 | -0.41514 | -0.51902
1.1 * B1=B2, 5ns | -1.0829 | -0.85447 | -0.63832 | -0.4262 | -0.21359 | 0.014633 | 0.23518 | 0.45148 | 0.66378 | 0.88357 | 1.0942
1.2 * B1=B2, 2.5ns | 0.55943 | 0.44914 | 0.33646 | 0.22357 | 0.11055 | 0.00083115 | -0.11162 | -0.22135 | -0.33514 | -0.44835 | -0.56065
1.2 *B1=B2, 5ns | -1.14 | -0.90666 | -0.6845 | -0.44995 | -0.22053 | 0.014861 | 0.23802 | 0.45803 | 0.69666 | 0.91925 | 1.1614
1.5 * B1=B2, 2.5ns | 0.6378 | 0.51283 | 0.38399 | 0.25475 | 0.1265 | 0.00096886 | -0.12693 | -0.25277 | -0.38116 | -0.50958 | -0.63874
1.5 * B1=B2, 5ns | -1.4034 | -1.0989 | -0.82574 | -0.54237 | -0.24881 | 0.010247 | 0.28942 | 0.56805 | 0.85992 | 1.1329 | 1.4165
Bunch 2 always centred

## Slide 30: Detail of Implementation

## Slide 31: AGM Performance
A-gap population published and logged at 1Hz
Sensitivity and accuracy depends on the energy and the particle type
  The higher the emission the more precise the measurement
For protons the sensitivity is better than 10% of quench level (fulfilling the specs) for all Energies
For Pb ions the AGM only fulfills the specs above 1.5 TeV
  Would need a new undulator to solve this
The error on the measurement is of the order of 50%

## Slide 32: AGM Error Sources
Alignment and steering
  Light collection efficiency very sensitive to these
Attenuation of light in optical components
  Can change due to dust, radiation etc.
PMT gain vs. voltage stability and HV control
Photocathode ageing
Electromagnetic noise in the signal

## Slide 33: BSR extraction mirror
Old
New

## Slide 34: Frequency domain simulations
In reality the peak is much higher
2012 Design
New Design
Power dissipated reduced ~100 times (~1w now)

## Slide 35: Original BSRT Telescope Layout
Light path shared between BSRT and AGM
Light path only used by AGM
AGM
From extraction mirror
Steering mirror
Splitting mirror
Telescope was based on spherical mirrors (hoping to use IR for ions at inj.)
In order to cope with the different synchrotron light sources an optical delay line (“trombone”) was used. This meant that the light had to go through many mirrors making alignment difficult
Errors in the alignment resulted in large variation of light collection efficiency

## Slide 36: New BSRT Layout
Light path shared between BSRT and AGM
Light path only used by AGM
AGM
From extraction mirror
Steering mirror
Splitting mirror
Spherical mirrors replaced by lenses
Since we can move the lenses, there is no need for the optical delay line anymore (“trombone”)
No moving parts before the AGM
Alignment much simpler and light collection less influenced
Fixed lens
Moving lens