## Slide 1
Ecole de Physique des Houches
Joel Fajans
Winter School on Physics with Trapped Charged Particles
19-30 January 2015
Particle and Plasma Diagnostics
Hollow antiproton plasma imaged on a microchannel plate,
ALPHA 2012

## Slide 2
Surko-Style Positron Accumulator
ALPHA-II Antihydrogen Apparatus
BASE (Anti)Proton Magnetic Moment Apparatus
How Do You Make a Particle/Plasma Experiment Work?
Notes: BEFORE NEXT SLIDE ASK AUDIENCE

## Slide 3
Interesting Parameters
Things we would like to know:
  Particle Number.
      Typical range: 103 to 109.
  What type of particles are present?
    Electrons, positrons, antiprotons, ions…
T. B. Mitchell, J. J. Bollinger, D. H. E. Dubin, X.-P. Huang, W. M. Itano,
R. H. Baughman, Science,282 1290 (1998)

## Slide 4
Interesting Parameters
Things we would like to know:
  Particle Number.
      Typical range: 103 to 109.
  What type of particles are present?
    Electrons, positrons, antiprotons, ions…
  Radial Profile.
      Typical range: 0.2mm to 10cm.
    Can the system be parameterized by a single number, i.e. the radius?

## Slide 5
Interesting Parameters
Things we would like to know:
  Particle Number.
      Typical range: 103 to 109.
  What type of particles are present?
    Electrons, positrons, antiprotons, ions…
  Radial Profile.
      Typical range: 0.2mm to 10cm.
    Can system be parameterized by a single number, i.e. the radius?
  Length.
      Typical range: 0.5cm to 100cm.
  Density
    Typical range: 103 to 109cm-3.
  Temperature
      Typical range 10 to 107K (0.5meV to 1000eV).
    Can system be parameterized by a single number, i.e. the temperature?
A.W. Hyatt, C.F. Driscoll and J.H. Malmberg "Measurement of the Anisotropic
Temperature Relaxation Rate in a Pure Electron Plasma," Phys. Rev. Lett. 59, 2975 (1987)

## Slide 6
Destructive vs. Nondestructive Diagnostics
Nondestructive:
  Advantages
    Can work with the plasmas afterwards.
  Disadvantages
    Can be slow.
    Can leave a wave imposed on the plasmas.
    Can heat the plasmas.
    May be difficult to calibrate.
Destructive:
  Advantages
    Sometimes easier to calibrate.
  Disadvantages
    Destroys the plasmas.
    Because the plasmas are changed radically, may have large systematic errors.
Notes: ASK STUDENTS FIRST

## Slide 7
Non-Laser Based Diagnostics
Concentrate on particles/plasmas for which laser diagnostics don’t work.
  Electrons, positrons, protons, antiprotons.

## Slide 8
Particle Counting-Faraday Cups
Faraday cups are used to count charged particles.
  Have been in use for a very long time.
  Basic design is very simple.
Principal parts:
  Charge collection cup: Self evident.
  Grounded shield: Self evident
  Secondary electron suppressor:
    When the primary charged particle hit the cup, secondary electrons may be emitted
    Secondary electron suppressor is biased negative relative to the charge collection cup.
    Sends secondary electrons back to the cup.
Notes: ASK STUDENTS FIRST

## Slide 9
http://dev.physicslab.org/Document.aspx?doctype=3&filename=Electrostatics_ShellsConductors.xml
http://en.wikipedia.org/wiki/Faraday's_ice_pail_experiment
Why are Faraday Cups Called Faraday Cups?
Faraday’s Ice Pail
(Effect first observed by Benjamin Franklin in 1755)

## Slide 10
Faraday Cup Implementation
http://en.wikipedia.org/wiki/Faraday_cup
Complicated
Simple

## Slide 11
Faraday Cup Circuit
Circuit Model
Typical Circuit Response (ALPHA)
is the charge deposited on the plate.

## Slide 12
Faraday Cup Noise
Circuit Model
Noise sources:
  Thermal (Johnson or Nyquist) noise.
    Johnson Noise , where is the bandwidth in which the signal is measured
    Total noise on plate is independent of the resistor .
    Capacitance comes from cable and cup.
    Typical capacitance is 1nF.
      This yields V of noise, or 12500e.
Notes: Why is the noise on the plate independent of the resistor?

## Slide 13
Faraday Cup Noise
Circuit Model
Noise sources:
  Thermal (Johnson or Nyquist) noise.
      This yields V of noise, or 12500e.
  Amplifier Noise
    Typical amplifier bandwidth might be 100kHz.
    Typical amplifier noise might be 2
      This yields a V of noise, or 3750e.
For these parameters, thermal noise dominates.

## Slide 14
Faraday Cup Noise
Circuit Model
Typical Circuit Response (ALPHA)
Actual Noise: about 200k electrons.
Where does the noise come from?
  50/60Hz coupling.
  Noise pickup from nearby cables.
  Accidental “thermocouple” junctions.
  Bad solder joints.
  Bad connectors.
  Triboelectric effect.
    System acts like a microphone.
Notes: Ask Students

## Slide 15
Triboelectric Noise Test
Diagram of triboelectric noise test set-up specified in ANSI/AAMI 5.5.4 EC53
All images from http://www.affinitymed.com/news/2014/august.aspx
Weight is 40 times the weight of one foot of cable.
Notes: Ask Students

## Slide 16
Radial Profile Diagnostic
Knowing the radius, and more generally the radial profile can be very helpful.
  Can be used to compute the plasma density. This is critical for studies of waves, 3-body recombination, calculating the collision frequency, etc..
  Can be used to check for and study instabilities.
  Can be used to determine basic plasma shape.
  Will tell you if the charge will fit through subsequent apertures.
Increasing Octupole Field: ALPHA 2008
Berkeley

## Slide 17
Radial Profile
Kyhl, R. L.; Webster, H. F., Breakup of hollow cylindrical electron beams, IRE Trans. Electron Dev., 3, 172 (1956).
Hollow beam electron gun
Phosphor Screen

## Slide 18
Radial Profile
A. J. Peurrung and J. Fajans, A pulsed, microchannel plate-based, non-neutral plasma imaging system. Rev Sci. Instrum., 64:52, 1993.
G.B. Andresen, et al Antiproton, Positron, and Electron Imaging with a Microchannel Plate/Phosphor Detector, Rev. Sci. Inst., 80, 123701, 2009.
Hollow antiproton plasma imaged on a microchannel plate,
ALPHA 2012

## Slide 19
Phosphor Screens
http://www.beamimaging.com/pscreen.html
http://www.proxivision.de/products/phosphor-screen.html

## Slide 20
Phosphor Screen
type | composition | emission |  | decay
P43 | Gd2O2S:Tb | 540 nm | green | 1 ms
P46 | Y3Al5O12:Ce | 530 nm | green | 300 ns
P47 | Y2SiO5:Ce,Tb | 400 nm | blue | 100 ns
Type | efficiency (lm/µA) |  |  |  | efficiency (W/mA) |  |  |  | efficiency (ph/el) |  |  | 
 | 6 kV | 10 kV | 12 kV | 15 kV | 6 kV | 10 kV | 12 kV | 15 kV | 6 kV | 10 kV | 12 kV | 15 kV
P43 | 0,24 | 0,43 | 0,54 | 0,71 | 0,43 | 0,77 | 0,97 | 1,28 | 185 | 330 | 420 | 550
P46 | 0,08 | 0,15 | 0,19 | 0,25 | 0,22 | 0,39 | 0,49 | 0,65 | 90 | 160 | 200 | 265
P47 | 0,06 | 0,11 | 0,14 | 0,18 | 0,62 | 1,35 | 1,71 | 2,24 | 212 | 380 | 480 | 
http://www.proxivision.de/products/phosphor-screen.html
Detection Efficiency:
  Phosphor Efficiency.
  Detector (likely a CCD camera) efficiency.
Generally, slow phosphors are more efficient.
http://www.andor.com/learning-academy/
ccd-spectral-response-%28qe%29-defining-the-qe-of-a-ccd

## Slide 21
Phosphor Screen Brightness
With Al Layer
No Al layer
Minimum voltage with a screen can be as low as about 1500V.
Approximately 100 photons per electron is typical.
A. J. Peurrung and J. Fajans, A pulsed, microchannel plate-based, non-neutral plasma imaging system. Rev Sci. Instrum., 64:52, 1993.

## Slide 22
Microchannel Plates
http://www.tectra.de/MCP.htm
http://www.tectra.de/MCP.htm
photonis.com
Typical pore size 2 - 25µm
http://www.pe1hmm.nl/xx1080/Introduction%20to%20Image%20Intensifier%20Tubes.htm
1000V

## Slide 23
J. Wiza, Microchannel Plate Detectors, Nuclear Instruments and Methods, 162, 587 (1979).
http://www.optimacorp.co.jp/BurleEO/MCP.htm
Microchannel Plate Gain

## Slide 24
J. Wiza, Microchannel Plate Detectors, Nuclear Instruments and Methods, 162, 587 (1979).
What limits Microchannel Plate?
MCP starts to arc at high voltages.
Channel get exhausted.
http://www.pe1hmm.nl/xx1080/Introduction%20to
%20Image%20Intensifier%20Tubes.htm

## Slide 25
Microchannel Plate Systems
MCP Front: Typically between -100 to +100V.
MCP Back: Typically between 500V to 1200V.
Phosphor Screen: Typically 3000-10000V.
Photonis APD Detector Brochure
http://www.dmphotonics.com/MCP_MCPImageIntensifiers/microchannel_plates.htm
Don’t do this yourself.

## Slide 26
Microchannel Plate System Performance
J. Wiza, Microchannel Plate Detectors, Nuclear Instruments and Methods, 162, 587 (1979).
Coatings available to enhance detection efficiency
For UV and X-rays.

## Slide 27
Microchannel Plate System Performance
ALPHA, Antiproton, Positron, and Electron Imaging with
Microchannel Plate/Phosphor Detector, Rev. Sci. Inst.,
80, 123701, (2009).
Hollow antiproton plasma imaged on a microchannel plate,
ALPHA 2012
J. Wiza, Microchannel Plate Detectors, Nuclear Instruments and Methods, 162, 587 (1979).

## Slide 28
O. Hadjar, W.K. Fowler, G. Kibelka, W.C. Schnute, Preliminary demonstration of an IonCCD as an alternative pixelated anode for direct MCP readout in a compact MS-based detector, J. Am. Soc. Mass. Spectrom., 23, 418, (2012).
Other Radial Diagnostics
The dream….direct reading of charge onto a CCD-like chip.
  Electronics would have to be UHV compatible.
  No opportunity for service.
  Detector area likely to be small.
  Possibly complex electrical feedthrough issues.
Notes: Back thinned CCD Fragile, uncommon.

## Slide 29
Other Radial Diagnostics
Multiple ring Faraday cups and possibly expansion by magnetic field reduction.
  Very crude.
  Very slow if magnetic expansion is used. Plasma may change during expansion.
  Usually requires many trials.
B. R. Beck, J. Fajans and J. H Malmberg, Temperature and anisotropic-temperature relaxation measurements in cold, pure-electron plasmas. Phys. Plasmas, 3:1250, 1996.
B.R. Beck, Measurement of the magnetic and temperature dependence of the electron-electron anisotropic temperature relaxation rate., Ph.D. Thesis, UCSD (1990).

## Slide 30
Other Radial Diagnostics
Magnetic field lowering and antiproton annihilation.
  Very crude.
    Plasma may change during expansion.
Total Charge and Inner Line Density Faraday Cups
  Requires the assumption of global thermal equilibrium.
Hollow antiproton plasma imaged on a microchannel plate,
ALPHA 2012
G. Gabrielse et al, Centrifugal separation of antiprotons and electrons, Phys. Rev. Lett., 105, 213002, (2010).
P. Oxley et al, Aperture method to determine the density and geometry of anti-particle plasmas, Phys. Lett. B 595, 60 (2004).
Berkeley
Berkeley

## Slide 31
Density Reconstruction
There are methods to measure the density directly, but frequently the density is inferred from:
  Measurements of the total charge.
  Measurements of the radial profile.
  Measurement (or assumption) of the plasma temperature.
  Knowledge of the potentials on the electrodes.
Algorithm:
  Assume local thermal equilibrium along magnetic field lines. Then simultaneously solve:
    Poisson’s equation: , where is the local density, and is the unit charge.
    Boltzmann’s equation at every radius:
    A normalization condition: .
    Note that is itself normalized so that is the total charge.
Notes: Draw out radial problem/
Discuss global and local thermal equilibrium.

## Slide 32
Density Reconstruction
Procedure:
  From an image of the plasma, construct the radial profile (generally, solvers assume azimuthal symmetry).
  Then solve the set of equations iteratively.
    Finite temperature method. Begin by gridding the trap in .
      Solve Poisson’s equation for by Fourier transforming at every . This yields a set of sparsely coupled equations indexed by , which can be inverted.
        Even though one uses an FFT, this is the time consuming step.
      Then use to calculate an new “trial” guess for at every radius.
      Calculate the final new guess . Adjust to optimize convergence.
      Iterate steps 1-2-3 until the desired convergence is obtained.
Notes: Numerical derivative

## Slide 33
Density Reconstruction
Procedure:
  From an image of the plasma, construct the radial profile (generally, solvers assume azimuthal symmetry).
  Then solve the set of equations iteratively.
    Finite temperature method. Begin by gridding the trap in .
    Convergence is slow.
      The grid size must be well less than in for convergence, and again well less than to properly model the radial edge.
      Convergence is particularly slow at low temperatures because becomes small, and the solutions oscillate if the grid is too crude. But making the grid fine makes the algorithm slow.
        The practical lower limit is a few hundred Kelvin.
      If the plasma is not well confined, i.e. if it is near to leaking out over the potential barrier, convergence is also difficult.
      Typical solve times range from 30s to a few hours…to never.

## Slide 34
Density Reconstruction
Procedure:
  From an image of the plasma, construct the radial profile (generally, solvers assume azimuthal symmetry).
  Then solve the set of equations iteratively.
    Finite temperature method. Begin by gridding the trap in .
    Zero temperatures method.
      By assuming a “waterbag” model, it is possible to solve this system at
        A waterbag is like a plasma filled balloon, with constant density along field lines inside, and zero density outside.
      With a waterbag, Poisson’s equation need only be solved at the waterbag boundary. No need to grid.
      Convergence is much faster…a few seconds to minutes.

## Slide 35
Density Reconstruction
Waterbag example: a barely confined plasma.

## Slide 36
Temperature Diagnostic
Consider a non-interacting, finite temperature, charged gas held in a potential well.
Notes: Contradiction between charged and non-interacting.

## Slide 37
Temperature Diagnostic: Noninteracting Gas
Consider a noninteracting charged gas held in a potential well.
How many gas particles escape?

## Slide 38
Temperature Diagnostic: Noninteracting Gas
Assume the particles in the gas are Maxwellian distributed.
  Distribution function: .
Lower the barrier height to some level .
  Then the charge that can escape over the barrier is .
    This is a complementary error function.
What is the slope on the log graph?

## Slide 39
Temperature Diagnostic: Noninteracting Gas
Starting with: .
We want to calculate .
Taking the derivative, we get:
This can be rewritten as:

## Slide 40
Temperature Diagnostic: Noninteracting Gas
This can be further simplified as:
When we are reasonably far out on the tail, , so we can approximate the above expression to:
Therefore, the slope of the log of the escaped charge is inversely proportional to the temperature.
A better approximation yields:

## Slide 41
Temperature Diagnostic: Simple Complication
The wall potential is not the well potential:
  The on-axis potential maximum is typically about 50% of the wall potential.
    The exact ratio depends on the electrode radius to length radius.
  The potential in the middle of the trap is not zero.
Both these effect need to be accounted for in the voltage derivative.
These issues are increasingly important in short traps.

## Slide 42
Temperature Diagnostic: Modest Complication

## Slide 43
Temperature Diagnostic: Modest Complication
Plasma expands during release.
  1D adiabatic expansion: temperature declines to 41% of its original value in this example.
  Plasma may undergo 3D expansion: temperature declines to 74% of its original value in this example.
  Actual expansion may be between 1D and 3D.

## Slide 44
Typical alpha parameters:
Temperature Diagnostic: Collective Effects
The self field of a plasma severely complicates the theoretical description of this diagnostic.
  The potential of a solid infinite cylinder is inside the plasma, and outside the plasma, where is the density, is the plasma radius and is the electrode wall radius.
The self potential is largest on the axis.

## Slide 45
Temperature Diagnostic: Radial Effects
Well is deeper off-axis.

## Slide 46
Temperature Diagnostic: Collective Effects
Well is deeper off-axis.
The self potential is largest on the axis.
Both these effects lead to the plasma escaping on the axis first.

## Slide 47
Temperature Diagnostic: From Where Does the Plasma Escape?
Plasma End View
Initially, the plasma escapes from a central region on the order of several Debye lengths:
Eventually, the center erodes significantly.
Once this happens, many things go wrong.
  Nontail particle begin to escape. The diagnostic was predicated on only tail particles escaping.
  The plasma hollows out, and the plasma potential changes significantly.

## Slide 48
Temperature Diagnostic: Useful Charge
Plasma End View
Initially, the plasma escapes from a central region on the order of several Debye lengths:
Eventually, the center erodes significantly.
Once the center erodes, the subsequent charge contains little temperature information.
  The rate at which the charge comes out is determined largely by the plasma self space charge.
Because of the scaling of the Debye length with temperature, ever fewer particles contain temperature information as the temperature increases.
D. L. Eggleston, C. F. Driscoll, B. R. Beck, A. W. Hyatt, and J. H. Malmberg, Parallel energy analyzer for pure electron plasma devices, Phys. Fluids B 4, 3432 (1992).
B. R. Beck, Measurement of the magnetic and temperature dependence of the electron-electron anisotropic temperature relaxation rate, Ph.D Thesis, UCSD, 1990.

## Slide 49
Andresen et al (ALPHA Collaboration), Evaporative Cooling of Antiprotons to Cryogenic Temperatures, Phys. Rev. Lett., 105 013003, 2010.
Temperature Diagnostic: Antiproton Temperatures

## Slide 50
Temperature Diagnostic: Hot Lepton Temperatures

## Slide 51
Temperature Diagnostic: Cold Lepton Temperatures

## Slide 52
Temperature Diagnostic: Cold Temperatures and Noise
The straight line region is short, and somewhat arbitrary at low temperatures.
  Pickup and amplifier noise limits the bottom end.
  Roll over from self potential effects limits the top end.
To get into the tens of Kelvin region, one typically needs sensitivities on the order of hundreds to thousands of charges.

## Slide 53
Temperature Diagnostic: Noise
The base temperature that one can measure is limited by amplifier and pickup noise.
  Acoustic noise from triboelectric effects is in the same band as the electrical pickup.
    This is particularly a problem in cyrogenic experiments.
  Consider:
    Mounting amps very close to the pickup. This may mean that the amps are:
      In vacuum.
      At cryogenic temperatures.
  Alpha uses a MCP preamp very successfully. This is the same system as we use for our radial profiles.

## Slide 54
Temperature Diagnostic: Cold Temperatures and Noise
Fitting a straight line seems arbitrary, but isn’t.
Fits are consistent to about 20%.
Systematic effects are on the order of 20% to 100%.
Notes: Takes persuading.

## Slide 55
Multi Species Temperature Diagnostic
ALPHA, Centrifugal Separation And Equilibration Dynamics In An Electron-Antiproton Plasma, Phys. Rev. Lett. 106, 145001 (2011)
As the plasma cools, it separates. Good, qualitative diagnostic.

## Slide 56
Beam Temperature Diagnostic: Pepperpot
http://www.dreebit.com
http://www.dreebit.com
http://www.davidshure.com/
item/d3111b-silver-pepper-pots
Beam emittance detector.
http://www.ischebeck.net/media/Beam%20
Profile%20Images/slides/PepperPot.html

## Slide 57
Nondestructive Diagnostics
If you can’t use lasers, how can you talk to a plasma?
Use waves or modes.
  Theory pioneered by Dan Dubin, and will be discussed by Francois.
  Two modes particularly useful.
    Sloshing mode: Used for calibration.
    Breathing mode: Can be used to determine density, aspect ratio, temperature.
M.D. Tinkle, R.G. Greaves, C. Surko, Lower order longitudinal modes
of single-component plasmas. Phys. Plasmas, 2, 2880, (1995).

## Slide 58
Modes Diagnostics
Put a plasma into a harmonic well.
  Actually, doesn’t need to be too harmonic
Which pickup is used for which mode?

## Slide 59
Modes Diagnostics
Measure the frequencies of both modes.
Measure the total charge.
From these measurements, use some somewhat nasty equations involving Legendre functions of the first and second kind to determine the aspect ratio, length and density of the plasma.
M.D. Tinkle, R.G. Greaves, C. Surko, Lower order longitudinal modes
of single-component plasmas. Phys. Plasmas, 2, 2880, (1995).

## Slide 60
Modes Diagnostics
Closer look:
  The theory predicts the mode frequency well, particularly for the second mode. (The first mode is used for calibration, and is defined to be exact.)
  The inverse prediction, which is what we want, is not particularly good.
M.D. Tinkle, R.G. Greaves, C. Surko, Lower order longitudinal modes
of single-component plasmas. Phys. Plasmas, 2, 2880, (1995).

## Slide 61
Modes Diagnostics
Why is the inverse prediction off?
  The equations are numerically unstable, particularly for high aspect ratio plasmas (high .
  The equations do not take into account the trap wall, which has significant effect.
    Numerical solutions have been attempted, but the inclusion of these effects is variously reported as “essentially impossible”1 and straightforward.2
    In my personal experience, the inversion problem does not satisfy elementary self-consistency checks for high plasmas. Others disagree.
M.D. Tinkle, R.G. Greaves, C. Surko, Lower order longitudinal modes
of single-component plasmas. Phys. Plasmas, 2, 2880, (1995).
1Francis Robicheaux, personal communication.
2Speck, Density and geometry of single component plasmas, Phys. Lett. B, 650, 119 (2007).

## Slide 62
Modes Diagnostics
Why is the inverse prediction off?
  The basic theory assumes the plasmas are cold.
  The theory has been extended for warm plasmas.
  This makes for a useful, relative, temperature diagnostic.
M.D. Tinkle, R.G. Greaves, C. Surko, Lower order longitudinal modes
of single-component plasmas. Phys. Plasmas, 2, 2880, (1995).
Notes: Here, the aspect ratio is already known.

## Slide 63
Modes Temperature Diagnostic
  The modes are a useful, relative, temperature diagnostic.
    Base is unknown.
ALPHA, In situ electromagnetic field diagnostics with an electron plasma in a Penning–Malmberg trap, New J. Phys., 16, 013037 (2014).
Notes: Measured shift of, say, 40kHz here leads to a large uncertainty in the temperature.

## Slide 64
N. Shiga, F. Anderegg, D.H.E. Dubin, C.F. Driscoll and R.W. Gould , Thermally excited fluctuations as a pure electron plasma temperature diagnostic, Phys. Plasmas 13, 022109:, (2006).
Thermal Modes Diagnostics
Using nontrivial apparatus, the thermal spectrum of the mode in the plasma can be measured.

## Slide 65
N. Shiga, F. Anderegg, D.H.E. Dubin, C.F. Driscoll and R.W. Gould , Thermally excited fluctuations as a pure electron plasma temperature diagnostic, Phys. Plasmas 13, 022109:, (2006).
Thermal Modes Diagnostics
Using nontrivial apparatus, the thermal spectrum of the mode in the plasma can be measured.
From the power in the spectra, the temperature can be inferred.
This is the perfect nondestructive temperature diagnostic…
  Except:
    Calibration is difficult.
    Probably requires a very quiet environment.
    May require amplifiers whose noise temperature is lower than the temperature of the plasmas under measurement.