## Slide 1: Deatailed design of a 25 T Cryogen-free Superconducting Magnet (25T-CSM)
High Field Laboratory Superconducting Materials,
IMR Tohoku University
Satoshi Awaji
Nov.13-14, 2014, Kyoto, Japan

## Slide 2: Collaborators
HFLSM, IMR, Tohoku Univ. H. Oguro, Y. Tuchiya, K. Watanabe
Toshiba H. Miyazaki, M. Takahashi, Y. Tosaka,
K. Tasaki, S. Hanai, S. Ioka
Fujikura S. Fujita, M. Daibo, H. Iijima
Furukawa M. Sugimoto, H. Tsubouchi
NIMS G. Nishijima，S. Matsumoto，S. Nimori，
H. Kumakura, T. Shimizu
LNCMI-CNRS Y. Miyoshi, X. Chaud, F. Debray

## Slide 3: Progress of Superconducting Magnets
15T-52
4T-36
20T-52
1GHz-NMR
(23.5T-f54)
25T-CSM

## Slide 4: Content
Development of 25T-CSM
  Overview
  Key technology of LTS coils (CuNb/Nb3Sn Rutherford )
  HTS coils (Gd123 CC )
    Load line and mechanical stress
    Epoxy impregnation (Deamination)
    AC-loss
    Quench protection

## Slide 5: Overview of the 25T-CSM
712mm
96mm
628mm
3 CuNb/Nb3Sn Rutherford solenoid
3 NbTi Rutherford solenoid
68 GdBCO single pancakes
2 x GM/JT cryocooler (8.6W@4.3K)
2 x 4K-GM cryocooler (3W@4.2K)
2 x 1 stg GM cryocooler
LTS part: 14T
HTS part: 11.5T

## Slide 6: Conceptual design of 25T-CSM
L ≈ 97H, Stored Energy ≈ 10.7MJ
（stress is for a whole cross-section of the conductors.）
 |  | REBCO | Nb3Sn | Nb3Sn | Nb3Sn | NbTi | NbTi | NbTi
Current | A | 135 | 851 |  |  |  |  | 
Inner radius | mm | 48.0 | 150.0 | 185.9 | 229.2 | 272.5 | 301.3 | 313.9
Outer radius | mm | 140.0 | 182.9 | 226.2 | 269.5 | 301.3 | 310.9 | 355.8
Height | mm | 394.4 | 540.0 | 627.8 | 627.8 | 627.0 | 628.1 | 628.1
Coil current density | A/mm2 | 110.8 | 68.9 | 68.9 | 68.9 | 71.6 | 90.0 | 90.0
No of turns/layer | - | 68 | 80 | 93 | 93 | 95 | 107 | 107
No of layer | - | 438 | 18 | 22 | 22 | 16 | 6 | 26
Total No of turns | - | 29784 | 1440 | 2046 | 2046 | 1520 | 642 | 2782
Bmax | T | 25.66 | 13.77 | 11.35 | 8.37 | 6.83 | 6.22 | 5.84
Br | T | 4.80 | 4.65 | 5.58 | 5.71 | 5.71 | 5.71 | 5.52
B0 | T | 11.56 | 2.43 | 2.91 | 2.73 | 1.91 | 0.78 | 3.24
Width of conductor | mm | 5.00 | 6.45 | 6.45 | 6.45 | 6.30 | 5.57 | 5.57
Thickness of conductor | mm | 0.13 | 1.83 | 1.83 | 1.83 | 1.80 | 1.61 | 1.61
Thickness of layer insulation | mm | 0.080 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
Jcon | A/mm2 | 207.7 | 105.8 | 105.8 | 105.8 | 105.8 | 138.1 | 138.1
Tcs | K |  | 5.87 | 7.28 | 8.58 | 5.92 | 6.12 | 6.32
Averaged compressive stress | MPa | -35 | -39 | -51 | -50 | -49 | -59 | -53
Hoop Stress BJR | MPa | 376 | 219 | 223 | 203 | 154 | 129 | 92
Hoop stress Wilson | MPa | 461 | 251 | 243 | 200 | 138 | 112 | 52

## Slide 7: Mechanical stress effects – strand in Rutherford cable–
Axial stress
Transverse stress
CuNb/Nb3Sn Rutherford cable
251MPa
5０MPa
(b) Strand and pre-bent cable for 25T-CSM

## Slide 8
Performance of Gd123 CC
455MPa

## Slide 9: Jc (B, q) at 4.2 K
Measured by Fujikura
The c-axis component of magnetic field determines Jc (B, q).
Intrinsic pinning is dominant.

## Slide 10: Field and Jc distribution（@25.5T）
Bmax = 25.7 T
B (T)
Angle of B (deg)
Ic (A)
Icmin=403A
T = 4.2K

## Slide 11: Epoxy Impregnation (Delamination)
Al cooling plate
Pancake coil
Cooling plate
F coated polyimide tape
REBCO tape
Epoxy resin
All turns are un-bonded but the edge part is connected to the cooling plate.
3 sections : riout/riin ≈ 1.26
5 sections : riout/riin ≈ 1.15
2nd stage
Notes: Bond, stick

## Slide 12: R&D test @LNCMI, Grenoble
Gd123 coil | 
Tape | Fujikura Cu-plated Gd123
Tape Ic | 297 A
Tape w | 5 mm
Tape tick | 0.12 mm
Turn tick | 0.18 mm
Coil Din | 96 mm
Coil Dout | 119.7 mm
Coil h | 10.7 mm
Turn No. | 58 x 2
Tape L | 20 x 2 m
Ic@77K, sf | 109 A (upper) 108 A (lower)
n-values | 28 (upper) 29 (lower)
M10@LNCMI,
Grenoble, France
19 T-170mm RT bore
section 2
Dout/Din ≈ 1.10

## Slide 13: 6 stacked coil
Cu-plated
Cu laminated
Cooling plate
coils
 | #1, #3, #5 | #2, #4, #6
Tape | Cu-plated Gd123 | Cu lamination on Cu plated Gd123
Tape width | 5 mm | 
Tape thickness | 0.12 mm | 0.17 mm
Inner diameter | 96 mm | 
Outer diameter | 172 mm | 
No of turns | 208 | 160
Tape length | 87 m | 68 m
t0.05-Cu
Joints

## Slide 14: Test results -6 stacked coil-
Magnetic field @462 A
0.51T
7.6T
20K
E (V/cm)
I (A)
 | 60 K |  | 50 K |  | 40 K |  | 30 K |  | 20 K | 
 | IC(A) | N | IC(A) | N | IC(A) | N | IC(A) | N | IC(A) | N
#1-#2 | - | - | 260 | 25 | 337 | 24 | 416 | 32 | - | -
#3-#4 | 182 | 21 | - | - | - | - | - | - | - | -
#5-#6 | - | - | - | - | - | - | 420 | 32 | 521 | 33

## Slide 15: AC loss of Gd123 tape
Experimental data were provided by Prof. Iwakuma using the saddle-shaped pick-up coil.
Very good agreement can be seen, but we have to take the effects of stacking into account!
Br
Stacked strip
Single strip

## Slide 16: Models of AC loss
W0 = m0JcwHc, c = pw/D, h = Hm/Hc andHm is an amplitude of applied magnetic field.
Brandt model for a strip (H. Brandt et al. , PRB, 48(1993)12893.)
Mawatari model for stacked strips (Y. Mawatari, PRB, 54 (1996) 13215.)
Kajikawa model for a slab (Kajikawa et al., 4LPo2B-06, http://arxiv.org/abs/1405.7765.)
B
B
B
thickness << width
thickness >>width

## Slide 17: Effect of tape stacking on AC-loss
Kajikawa et al., Physica C 469 (2009) 1436.
The AC-loss approaches from the strip model to the slab model, when the number of stacks increases.
Due to an increase of Bp by tape stacking, AC losses in low field region decreases.
N
strip
slab

## Slide 18: Calculation of AC loss power (time derivative)
Slab model
Strip model (Brandt) w/o transport current (only external field)
Br
Bz
Jc

## Slide 19: Perpendicular-field loss
The perpendicular-field ac loss of about 5 W estimated, if the effect of tape stacking, although it is about the twice of loss without the stacking effects (stripe model). In particular, the field dependency are much different each other.

## Slide 20: Cooling the HTS coil in the 25T-CSM
If the temperature rise to about 10 K is allowed, the heat load of 20 W
(10 W/cooler) can be permitted.

## Slide 21: Cooling the LTS coil in the 25T-CSM
Heat load to the GM/JT cryocoolers at 4 K
mode 1 (300–50 K)
mode 2 (50–20 K)
mode 3 (20–4 K)
Cooling modes

## Slide 22: Design of a 25T-CSM (modified)
L ≈ 97H, Stored Energy ≈ 10.7MJ
（stress is for a whole cross-section of the conductors.）
 |  | YBCO | Nb3Sn | Nb3Sn | Nb3Sn | NbTi | NbTi | NbTi
Current | A | 140 | 851 |  |  |  |  | 
Inner radius | mm | 51 | 150.0 | 185.9 | 229.2 | 272.5 | 301.3 | 313.9
Outer radius | mm | 138 | 182.9 | 226.2 | 269.5 | 301.3 | 310.9 | 355.8
Height | mm | 408 | 540.0 | 627.8 | 627.8 | 627.0 | 628.1 | 628.1
Space current density | A/mm2 | 110.8 | 68.9 | 68.9 | 68.9 | 71.6 | 90.0 | 90.0
No of turns/layer | - | 68 | 80 | 93 | 93 | 95 | 107 | 107
No of layer | - | 435 | 18 | 22 | 22 | 16 | 6 | 26
Total No of turns | - | 29580 | 1440 | 2046 | 2046 | 1520 | 642 | 2782
Bmax | T | 25.68 | 13.77 | 11.35 | 8.37 | 6.83 | 6.22 | 5.84
Br | T | 4.80 | 4.65 | 5.58 | 5.71 | 5.71 | 5.71 | 5.52
B0 | T | 11.5 | 2.43 | 2.91 | 2.73 | 1.91 | 0.78 | 3.24
Width of conductor | mm | 5.00 | 6.45 | 6.45 | 6.45 | 6.30 | 5.57 | 5.57
Thickness of conductor | mm | 0.13 | 1.83 | 1.83 | 1.83 | 1.80 | 1.61 | 1.61
Thickness of layer insulation | mm | 0.070 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075 | 0.075
Jcon | A/mm2 | 217 | 105.8 | 105.8 | 105.8 | 105.8 | 138.1 | 138.1
Tcs | K |  | 5.87 | 7.28 | 8.58 | 5.92 | 6.12 | 6.32
Averaged compressive stress | MPa | -35 | -39 | -51 | -50 | -49 | -59 | -53
Hoop Stress BJR | MPa | 417 | 219 | 223 | 203 | 154 | 129 | 92
Hoop stress Wilson | MPa | 455 | 251 | 243 | 200 | 138 | 112 | 52

## Slide 23: Protection circuit
LTS part
HTS part
RA=RA=1W
The quench is detected by the balance voltage of LTS.
The quench of HTS is not cared because of its high stability.
140A
HTS
IHTS(A)
ILTS(A)
RH=11W
LTS-A
LTS-B
Vmax ≈ 1.87kV

## Slide 24: Quench test of small coils
Outer LTS Magnet
REBCO
Single pancake coil
 | LTS | REBCO
D out(mm) | 380 | 156
D in (mm) | 317 | 96
Height(mm) | 377 | 4.1
L (mH) | 24 | 1.73
M (mH) | 34 | 
SC thickness(μm) |  | 1.5
Cu thickness(μm) |  | 100
Resistive V
Dump Resistance

## Slide 25: Quench behavior
Resistive Vres (V)
Dump resistance dependence
Initial Temp. dependence
Maximum Tmax
w temperature rise
Higher T for Lower Initial-T
w/o temperature rise

## Slide 26: Temperature rise
Protection by
temp. margin
Protection by Quench heater
Imax = Ic
Tmax = Tc
Temperature Increase Tinc(Tini, R)
Negligible T-rise

## Slide 27: Quench analysis
Hot-spot at High-T is more safty!
Low temp.
High temp.
Temperature rise
低温ほどホットスポットによる温度上昇は大きい
←コイル内径
コイル外径→
低温ほどホットスポット領域が狭い
Initial-T dependence
Initial-T | low | high
Temperature rise | large | small
Hot-spot area | narrow | wide
Risk of burn-out | large | small
Inner
Outer
Smaller hot-spot at lower-T
Higher Temp rise at lower-T

## Slide 28: Performance of HTS tapes
淡路 智，まぐね， vol.9, No4 (2014)

## Slide 29: HTS insert for 25T-CSM (modified)
 |  | YBCO | Bi2223
Current | A | 140 | 224
Inner radius | mm | 51.1 | 48.0
Outer radius | mm | 140.0 | 139.3
Height | mm | 394.4 | 391.4
Space current density | A/mm2 | 117 | 112
No of turns/layer | - | 68 | 76
No of layer | - | 435 | 234
Total No of turns | - | 29580 | 17784
Bmax | T | 25.66 | 25.62
Br | T | 5.11 | 4.95
B0 | T | 11.5 | 11.50
Width of conductor | mm | 5.00 | 5.15
Thickness of conductor | mm | 0.13 | 0.32
Thickness of layer insulation | mm | 0.080 | 0.07
Jcon | A/mm2 | 215 | 157
Averaged compressive stress | MPa | -35 | -32
Hoop stress Wilson | MPa | 455 | 330

## Slide 30: HFLSM annex building and 25T-CSM
Experimental
preparation
room
Machinery
Measuring-
equipment
storage
WC
WC
Hall
Front
door
25T-CSM
Sink
Closet
HFLSM Annex １F
1390WX2840LX1190H

## Slide 31: Summary
The 25T-CSM project is going on now at the HFLSM, Sendai, Japan.
We adopted a high stress design for both of LTS and HTS coils.
The issues of impregnation, AC losses, quench protection are being overcome.
However, we just start to consider the possibility of high strength Bi2223 tape (HT-XX) for the HTS insert.
The construction must be finished by March, 2015.