## Slide 1: ISOLDE Technical Report
70th Meeting of the ISCC
24th June 2014
Richard Catherall, EN-STI

## Slide 2: Outline
Target Area
  Robots and auxiliary equipment
    services
  Modifications for Medicis
    Shielding, Montrac, shroud and FC door modifications
  Hot Cell
  PAD/MAD installation
Targets
RILIS
HRS20
Update and Critical Path

## Slide 3: Robots
Robots are in place and commissioning has started this week
Controls and safety interlocks are in place
Need to test, test and test again

## Slide 4: Shelf Modifications
Courtesy of M. Owen

## Slide 5: Services
The services to be displaced included:
Electrical cables and cable tray
Flexible tubes to ventilate the HRS FC
Water tubes for the BTY line magnet cooling
Welded for the 4th time last Friday
X-rays to be done this evening
May be a bottle neck for the start up!

## Slide 6: Collective dose for robot installation
Chart: 
| Phase 5 : Tests des robots | Phase 4: Installation des robots | Phase 3: Modification des services | Phase 2: Démantèlement des robots | Phase 1: Installation des blindages
Real dose (man.µSv) | 100 | 12 | 60 | 950 | 19000
Estimated (man.µSv) | 1609 | 0 | 0 | 1547 | 1648
Dose collective totale estimée = 1547 µSv
Dose collective totale estimée = 2313 µSv
Dose collective totale estimée = 2905 µSv
PHASE EN COURS – 39 % de la dose estimée
PHASE EN COURS – 197 % de la dose estimée
PHASE EN COURS – 152 % de la dose estimée
PHASE EN COURS – 6 % de la dose estimée
Dose collective totale estimée = 1609 µSv
Dose collective totale estimée = 1648 µSv
Dose collective cumulée = 610 µSv
Dose collective cumulée = 4557 µSv
Dose collective cumulée = 4410 µSv
Dose collective cumulée = 1505 µSv
Dose collective cumulée = 101 µSv
PHASE EN COURS – 91 % de la dose estimée
Phase 1: Installation des blindages
Phase 2: Démantèlement des robots
Phase 3: Modification des services
Phase 4 : Installation des robots
Phase 5 : Tests des robots
Courtesy of C. Lemesre

## Slide 7: MEDICIS shielding
Displacement of services
Opening of wall for future Montrac installation
Installation of shielding has started and should be finished this week

## Slide 8: MEDICIS Irradiation Point
Installation of irradiation point and Montrac system including HT protection (shroud)
Faraday cage pistons modified

## Slide 9: ISOLDE Hot Cell

## Slide 10: PAD/MAD Access to Target Area
Installed and now operational. A few access delays during implementation
SIR, ISOPRIM access, dosimeter and registered on IMPACT

## Slide 11: 11C Beams for combined PET/Hadrontherapy
1+ ions
HV platform
TRAP
EBIS
60 kV
Analyzing magnet,
Bunches and cools
ions
Charge
breeding
Triplet
Multiple charged
ions
Switched
HV platform
60 « 20 kV
RFQ
IH
Accelerator
from
5 -> 300 keV/u
Accelerator
from
0.3 -> 1.2 MeV/u
Doublets
Triplet
7-gap
resonator
Accelerator
from
1.2 -> max 2.2 MeV/u
Target and
detectors
ISOLDE
11C Post accelerated radioactive ion beam possible
for hadrontherapy
Unexpected developments from Molten salt target tests
T. Mendonca et al., CERN-ACC-NOTE-2014-0028

## Slide 12: LaC2–C Spallation Target Material for n-def Ba beams
Adopting procedure as developed for nano-UCx in 2012 towards rare-earth carbides
La(OH)3 particle size reduction by suspension grinding
Theoretical Production Yields
after comparing all possible spallation target nuclei
Hydroxide instead of oxide as precursor
nanometric carbon for higher homogeneity
MWCNT to increase thermal stability
Developing wet-milling procedure for the production of nanopowders
Ultrasound-assisted nano-incorporation of La(OH)3 into MWCNT fibers
1 µm
micrometric graphite
micrometric La(OH)3 (as ISOLDE standard)
MWCNT
micrometric La(OH)3 (dry mixing)
MWCNT
micrometric La(OH)3 (new suspension mixing)
Result of new nano-incorporation:
La(OH)3 particle
agglomerate of
MWCNT fibres
MWCNT fibres
Alexander Gottberg and Julien Guillot

## Slide 13: RILIS laser alignment target unit
Mounted as first target on HRS after LS1. Lasers successfully aligned. Unit Functioning as expected.
 facilitates laser alignment - normally only possible with ion beam
 helps to clarify laser beam transport issues (thermal lensing)
 enables remote check of status of magnet windows when in doubt, reducing interventions to separator areas
Thanks to: Bernard, Ermanno, Tim & our 2013 summer student Sasha
Tungsten ionizer acts as light source for pre-alignment after shutdown
Laser power meter measures power delivered through ion source
readout via existing thermocouple infrastructure
RILIS lasers
Magnet window
HRS/GPS
Target Unit
Courtesy of S. Rothe

## Slide 14: Status HRS20
Courtesy of Jose A. Ferreira on behalf of TE/VSC

## Slide 15: Layout
P ~ 10 mbar
Modification laser window
Sector valve
Sector valve
Laser window

## Slide 16: HRS90

## Slide 17: Actions
* Sniffer test  Helium pressurization and external detection (non quantitative)
24th April: First leak detection using a sniffer. Leak detected close to the window (1e-4 mbar·L/s)
28th of April: Exchange laser window (old window) and o-ring, pump-down
29th of April: new o-ring installed (pump-down)
30th of April: New sniffer test. All the sector tested, signal from the window (1e-5 mbar·L/s)
7th of May: Sniffer test and shutter mechanism. Again the leak is found on the window. We checked all the sector without any other result.
23th of May: Sniffer test: Clear signal from window (1e-3 mbar·L/s). New plug instead of window.
30th of May, sealing agent, unfortunately pressure was not reduced significantly (always mbar range)

## Slide 18: Actions (sector splitted)

## Slide 19: Results
Leak between the magnet and the frontend.
Faraday cage (visual inspection)
The clamp is bigger than the flange!!
O-ring not accessible
Not documented!
Same layout in GPS?

## Slide 20: Results

## Slide 21: Results

## Slide 22: HRS20 Conclusions
Corrosion of non-stainless steel pieces (brass)
Almost all the clamps replaced to prevent further failures:
  BTY3 and HRS20
  1xDN250 in GPS not accessible (NOT DONE)
  2(?)xViton o-ring not accessible to replace.
  Standard clamps on non-standard conical flanges
Leak test of the sector (ongoing). Still leaking  Test (today), installation of Viton seal on the external part of the flange
LS2 consolidation required

## Slide 23: Update
GPS extraction electrode repaired
  Ba beams being produced since 3 weeks for RILIS development
Modification of Frontend coupling tables done
Radhard pistons installed on both FC doors
Change of clamps for preventive maintenance
Improved leak tightness of the target area
Change of lighting in part of the target area
New laser window system in HRS.MAG90
RFQ Cooler installed and aligned externally
  Need beam to further check alignment
  New RF amplifier also installed
Change of deflector motors and controls implemented
New HRS magnet controls and FESA class implemented
Most of the software now running

## Slide 24: Critical Path
Final commissioning of the robots
Stable beam through the RFQ Cooler
Testing of the repaired tape station
Protons still planned for the 10th July
  SEM grid tests, HT on HRS with protons, RIB from GPS for tape station tests
Request further stable beam on HRS to commission RFQ Cooler…~2 weeks